## GTR1 G315 z0 x
GTR1 Spot size in x axis at iso for G315º every 10 MeV
Lynx PT
notification | 5%
warrning | 8%
action level | 10%
validation of new set point for deflector
|  | E [MeV] | REF | [16 empty] | After deflector adjustment due to IDDC measurements showing deviations |  |  | SMPS replacement | [6 empty] | Lynx SN 63057 | Lynx SN 63057 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 |  | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021
REF | time axis 2019-01-10 to 2023-06-20
|  |  |  | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm]
|  | 60 | 7.085 | 6.92 | 6.74 | 6.94 | 6.9 | 6.926 | 6.51 | 6.95 | 6.92 | 6.95 | 6.95 | 6.91 | 6.92 | 6.94 | 6.92 | 6.92 | 6.98 | 6.94 | 6.91 | 6.94 | 6.92 | 6.94 | 6.92 | 6.94 | 6.97 | 6.95 | 6.91 | 6.95 | 6.95 | 7.02 | 7 | 7.01 | 7 | 6.99 | 7 | 6.99 | 7 | 6.99 | 7.01 | 7.01 | 7 | 7 | 6.99 | 6.98 | 6.97 | 6.95 | 6.96 | 6.96 | 6.97
|  | 70 | 6.23 | 6.08 | 6.29 | 6.11 | 6.05 | 6.109 | 6.29 | 6.11 | 6.09 | 6.1 | 6.1 | 6.07 | 6.08 | 6.09 | 6.07 | 6.07 | 6.13 | 6.08 | 6.07 | 6.08 | 6.08 | 6.08 | 6.07 | 6.09 | 6.12 | 6.09 | 6.06 | 6.11 | 6.1 | 6.15 | 6.14 | 6.14 | 6.14 | 6.15 | 6.16 | 6.14 | 6.17 | 6.15 | 6.17 | 6.16 | 6.14 | 6.13 | 6.14 | 6.13 | 6.13 | 6.12 | 6.13 | 6.12 | 6.12
|  | 80 | 5.93 | 5.7 | 5.58 | 5.8 | 5.69 | 5.765 | 5.53 | 5.78 | 5.75 | 5.73 | 5.74 | 5.76 | 5.75 | 5.74 | 5.73 | 5.72 | 5.75 | 5.72 | 5.73 | 5.7 | 5.75 | 5.76 | 5.72 | 5.77 | 5.75 | 5.76 | 5.76 | 5.77 | 5.79 | 5.83 | 5.78 | 5.81 | 5.85 | 5.84 | 5.88 | 5.82 | 5.83 | 5.83 | 5.89 | 5.87 | 5.84 | 5.83 | 5.85 | 5.84 | 5.85 | 5.82 | 5.83 | 5.8 | 5.8
|  | 90 | 5.519 | 5.28 | 5.12 | 5.32 | 5.33 | 5.391 | 5.18 | 5.34 | 5.36 | 5.32 | 5.35 | 5.31 | 5.33 | 5.32 | 5.3 | 5.31 | 5.39 | 5.33 | 5.32 | 5.34 | 5.32 | 5.32 | 5.31 | 5.32 | 5.35 | 5.35 | 5.33 | 5.35 | 5.36 | 5.37 | 5.4 | 5.39 | 5.39 | 5.43 | 5.42 | 5.42 | 5.42 | 5.43 | 5.46 | 5.46 | 5.4 | 5.42 | 5.42 | 5.43 | 5.4 | 5.39 | 5.41 | 5.41 | 5.38
|  | 100 | 5.11 | 4.93 | 4.7 | 4.98 | 4.93 | 4.966 | 4.51 | 4.96 | 4.95 | 4.95 | 4.97 | 4.93 | 4.93 | 4.92 | 4.91 | 4.92 | 4.98 | 4.93 | 4.89 | 4.91 | 4.9 | 4.92 | 4.89 | 4.93 | 4.95 | 4.95 | 4.94 | 4.96 | 4.96 | 5.03 | 5.03 | 5.02 | 5.02 | 5.04 | 5.06 | 5.04 | 5.05 | 5.03 | 5.1 | 5.08 | 5.03 | 5.01 | 5.03 | 5.05 | 5.02 | 5.04 | 5.04 | 5 | 5.01
|  | 110 | 4.754 | 4.59 | 4.37 | 4.62 | 4.58 | 4.607 | 4.19 | 4.59 | 4.61 | 4.58 | 4.62 | 4.57 | 4.58 | 4.59 | 4.57 | 4.57 | 4.63 | 4.59 | 4.57 | 4.56 | 4.56 | 4.57 | 4.57 | 4.59 | 4.6 | 4.62 | 4.6 | 4.6 | 4.61 | 4.68 | 4.66 | 4.68 | 4.69 | 4.7 | 4.69 | 4.7 | 4.69 | 4.69 | 4.72 | 4.71 | 4.69 | 4.68 | 4.68 | 4.7 | 4.67 | 4.67 | 4.69 | 4.69 | 4.68
|  | 120 | 4.457 | 4.32 | 4.08 | 4.35 | 4.3 | 4.335 | 3.91 | 4.35 | 4.32 | 4.33 | 4.33 | 4.31 | 4.3 | 4.31 | 4.29 | 4.29 | 4.35 | 4.32 | 4.3 | 4.31 | 4.3 | 4.29 | 4.3 | 4.3 | 4.32 | 4.33 | 4.32 | 4.35 | 4.34 | 4.39 | 4.39 | 4.43 | 4.42 | 4.44 | 4.44 | 4.42 | 4.41 | 4.41 | 4.47 | 4.45 | 4.4 | 4.39 | 4.41 | 4.42 | 4.4 | 4.41 | 4.4 | 4.4 | 4.38
|  | 130 | 4.199 | 4.02 | 3.9 | 4.06 | 4.03 | 4.062 | 3.84 | 4.07 | 4.05 | 4.05 | 4.04 | 4.03 | 4.03 | 4.03 | 4 | 4.02 | 4.08 | 4.04 | 4.01 | 4.03 | 4.01 | 4.02 | 4.01 | 4.02 | 4.04 | 4.06 | 4.04 | 4.04 | 4.05 | 4.12 | 4.09 | 4.14 | 4.13 | 4.14 | 4.12 | 4.12 | 4.13 | 4.12 | 4.16 | 4.17 | 4.11 | 4.12 | 4.11 | 4.12 | 4.11 | 4.11 | 4.11 | 4.12 | 4.11
|  | 140 | 4.011 | 3.82 | 3.83 | 3.9 | 3.84 | 3.843 | 3.87 | 3.86 | 3.86 | 3.84 | 3.85 | 3.83 | 3.83 | 3.83 | 3.8 | 3.81 | 3.88 | 3.82 | 3.82 | 3.82 | 3.81 | 3.82 | 3.82 | 3.83 | 3.85 | 3.84 | 3.85 | 3.86 | 3.85 | 3.9 | 3.9 | 3.92 | 3.92 | 3.93 | 3.96 | 3.92 | 3.93 | 3.92 | 3.99 | 3.98 | 3.89 | 3.88 | 3.89 | 3.91 | 3.89 | 3.9 | 3.91 | 3.9 | 3.88
|  | 150 | 3.833 | 3.68 | 3.69 | 3.73 | 3.67 | 3.698 | 3.76 | 3.71 | 3.68 | 3.68 | 3.68 | 3.67 | 3.66 | 3.65 | 3.62 | 3.64 | 3.71 | 3.67 | 3.65 | 3.65 | 3.66 | 3.66 | 3.66 | 3.67 | 3.69 | 3.69 | 3.67 | 3.7 | 3.69 | 3.75 | 3.73 | 3.77 | 3.77 | 3.79 | 3.79 | 3.75 | 3.78 | 3.75 | 3.83 | 3.82 | 3.71 | 3.71 | 3.72 | 3.75 | 3.72 | 3.75 | 3.75 | 3.73 | 3.72
|  | 160 | 3.614 | 3.5 | 3.59 | 3.53 | 3.49 | 3.485 | 3.71 | 3.51 | 3.49 | 3.47 | 3.49 | 3.47 | 3.45 | 3.44 | 3.41 | 3.43 | 3.5 | 3.45 | 3.44 | 3.45 | 3.45 | 3.45 | 3.45 | 3.47 | 3.5 | 3.48 | 3.47 | 3.5 | 3.48 | 3.53 | 3.53 | 3.58 | 3.57 | 3.59 | 3.6 | 3.55 | 3.57 | 3.55 | 3.65 | 3.63 | 3.49 | 3.48 | 3.5 | 3.53 | 3.49 | 3.52 | 3.52 | 3.51 | 3.49
|  | 170 | 3.45 | 3.35 | 3.45 | 3.38 | 3.31 | 3.331 | 3.51 | 3.34 | 3.35 | 3.32 | 3.32 | 3.32 | 3.3 | 3.27 | 3.24 | 3.26 | 3.34 | 3.29 | 3.26 | 3.3 | 3.29 | 3.29 | 3.3 | 3.31 | 3.33 | 3.33 | 3.31 | 3.34 | 3.32 | 3.38 | 3.37 | 3.42 | 3.42 | 3.43 | 3.45 | 3.39 | 3.41 | 3.4 | 3.51 | 3.48 | 3.32 | 3.31 | 3.32 | 3.37 | 3.32 | 3.36 | 3.37 | 3.34 | 3.32
|  | 180 | 3.303 | 3.21 | 3.45 | 3.23 | 3.18 | 3.183 | 3.57 | 3.2 | 3.2 | 3.17 | 3.19 | 3.18 | 3.15 | 3.12 | 3.08 | 3.11 | 3.2 | 3.14 | 3.12 | 3.16 | 3.14 | 3.15 | 3.14 | 3.17 | 3.19 | 3.19 | 3.15 | 3.21 | 3.19 | 3.25 | 3.24 | 3.29 | 3.29 | 3.3 | 3.31 | 3.25 | 3.28 | 3.26 | 3.38 | 3.37 | 3.15 | 3.14 | 3.17 | 3.21 | 3.16 | 3.21 | 3.22 | 3.18 | 3.18
|  | 190 | 3.184 | 3.15 | 3.37 | 3.14 | 3.11 | 3.091 | 3.42 | 3.1 | 3.09 | 3.07 | 3.08 | 3.09 | 3.06 | 3 | 2.97 | 2.99 | 3.11 | 3.03 | 3.02 | 3.06 | 3.06 | 3.05 | 3.05 | 3.08 | 3.09 | 3.09 | 3.06 | 3.12 | 3.08 | 3.13 | 3.13 | 3.19 | 3.19 | 3.22 | 3.23 | 3.16 | 3.19 | 3.15 | 3.31 | 3.27 | 3.04 | 3.03 | 3.07 | 3.12 | 3.06 | 3.15 | 3.13 | 3.09 | 3.05
|  | 200 | 3.087 | 3.08 | 3.28 | 3.08 | 3.02 | 3.012 | 3.26 | 3 | 3.03 | 2.97 | 2.98 | 3 | 2.99 | 2.91 | 2.87 | 2.89 | 3.02 | 2.95 | 2.94 | 2.98 | 2.96 | 2.97 | 2.95 | 2.99 | 3.02 | 3.02 | 2.98 | 3.07 | 2.99 | 3.05 | 3.05 | 3.11 | 3.12 | 3.16 | 3.17 | 3.09 | 3.11 | 3.07 | 3.26 | 3.22 | 2.96 | 2.93 | 2.98 | 3.05 | 2.98 | 3.08 | 3.06 | 3.01 | 2.97
|  | 210 | 2.965 | 3.04 | 3.09 | 3.1 | 2.97 | 2.95 | 3.05 | 2.93 | 2.93 | 2.91 | 2.93 | 2.97 | 2.91 | 2.83 | 2.79 | 2.81 | 2.95 | 2.9 | 2.88 | 2.92 | 2.9 | 2.9 | 2.91 | 2.93 | 2.96 | 2.96 | 2.91 | 3.01 | 2.91 | 2.99 | 3 | 3.09 | 3.1 | 3.11 | 3.12 | 3.04 | 3.06 | 3.01 | 3.23 | 3.17 | 2.9 | 2.89 | 2.93 | 3 | 2.91 | 3.05 | 3.01 | 2.95 | 2.92
|  | 220 | 2.888 | 3.01 | 2.99 | 2.99 | 2.91 | 2.875 | 2.89 | 2.92 | 2.86 | 2.86 | 2.91 | 2.93 | 2.9 | 2.8 | 2.76 | 2.78 | 2.92 | 2.87 | 2.85 | 2.92 | 2.88 | 2.88 | 2.89 | 2.92 | 2.95 | 2.94 | 2.9 | 3.01 | 2.88 | 2.98 | 2.99 | 3.07 | 3.09 | 3.12 | 3.15 | 3.07 | 3.04 | 2.97 | 3.26 | 3.21 | 2.93 | 2.91 | 2.95 | 3.01 | 2.93 | 3.08 | 3.04 | 2.96 | 2.95
|  | 226 | 2.9 | 3.01 | 2.98 | 2.98 | 2.92 | 2.895 | 2.92 | 2.89 | 2.9 | 2.87 | 2.93 | 3 | 2.89 | 2.82 | 2.75 | 2.79 | 2.94 | 2.89 | 2.86 | 2.94 | 2.9 | 2.9 | 2.91 | 2.94 | 2.98 | 2.98 | 2.91 | 3.05 | 2.91 | 2.99 | 3.01 | 3.11 | 3.12 | 3.15 | 3.19 | 3.11 | 3.09 | 2.98 | 3.26 | 3.21 | 2.97 | 2.95 | 3.01 | 3.06 | 2.97 | 3.12 | 3.06 | 3.01 | 2.99
Deviation from reference [%]
|  | E [MeV] | OOS | 2020-11-18: After deflector adjustment due to IDDC measurements showing deviations; 2021-01-11: SMPS replacement
|  |  | 2018-05-29 | [5 empty] | 2019-09-11 | 2019-11-12 |  | 2020-02-02 | 2020-03-25 | 2020-05-05 | 2020-06-02 | 2020-07-28 | 2020-09-01 | 2020-10-06 | 2020-10-11 | 2020-11-18 | 2020-12-10 | 2021-01-05 | 2021-01-11 | 2021-02-04 | 2021-03-09 | 2021-04-06 | 2021-06-22 | 2021-07-20 | 2021-08-17 | 2021-10-12 | 2021-10-17 | 2021-11-08 | 2021-12-09 | 2022-01-21 | 2022-02-08 | 2022-03-08 | 2022-03-13 | 2022-04-04 | 2022-05-16 | 2022-07-08 | 2022-07-12 | 2022-08-30 | 2022-10-09 | 2022-11-08 | 2022-12-01 | 2023-02-04 | 2023-03-10 | 2023-03-20 | 2023-04-22 | 2023-05-27 | 2023-06-20 | 0 | 0 | 0 | 0 | 0
|  |  | sigma x [mm] | [5 empty] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm]
|  | 60 | #REF! | 2.384% | 5.119% | 2.089% | 2.681% | 2.296% | -1.905% | -2.329% |  | -1.905% | -1.905% | -2.47% | -2.329% | -2.047% | -2.329% | -2.329% | -1.482% | -2.047% | -2.47% | -2.047% | -2.329% | -2.047% | -2.329% | -2.047% | -1.623% | -1.905% | -2.47% | -1.905% | -1.905% | -0.9174% | -1.2% | -1.059% | -1.2% | -1.341% | -1.2% | -1.341% | -1.2% | -1.341% | -1.059% | -1.059% | -1.2% | -1.2% | -1.341% | -1.482% | -1.623% | -1.905% | -1.764% | -1.764% | -1.623% | -100% | -100% | -100% | -100% | -100%
|  | 70 | #REF! | 2.467% | -0.9539% | 1.964% | 2.975% | 1.981% | -1.926% | -2.247% |  | -2.087% | -2.087% | -2.568% | -2.408% | -2.247% | -2.568% | -2.568% | -1.605% | -2.408% | -2.568% | -2.408% | -2.408% | -2.408% | -2.568% | -2.247% | -1.766% | -2.247% | -2.729% | -1.926% | -2.087% | -1.284% | -1.445% | -1.445% | -1.445% | -1.284% | -1.124% | -1.445% | -0.9631% | -1.284% | -0.9631% | -1.124% | -1.445% | -1.605% | -1.445% | -1.605% | -1.605% | -1.766% | -1.605% | -1.766% | -1.766% | -100% | -100% | -100% | -100% | -100%
|  | 80 | #REF! | 4.035% | 6.272% | 2.241% | 4.218% | 2.862% | -2.53% | -3.035% |  | -3.373% | -3.204% | -2.867% | -3.035% | -3.204% | -3.373% | -3.541% | -3.035% | -3.541% | -3.373% | -3.879% | -3.035% | -2.867% | -3.541% | -2.698% | -3.035% | -2.867% | -2.867% | -2.698% | -2.361% | -1.686% | -2.53% | -2.024% | -1.349% | -1.518% | -0.8432% | -1.855% | -1.686% | -1.686% | -0.6745% | -1.012% | -1.518% | -1.686% | -1.349% | -1.518% | -1.349% | -1.855% | -1.686% | -2.192% | -2.192% | -100% | -100% | -100% | -100% | -100%
|  | 90 | #REF! | 4.527% | 7.793% | 3.741% | 3.546% | 2.374% | -3.243% | -2.881% |  | -3.606% | -3.062% | -3.787% | -3.425% | -3.606% | -3.968% | -3.787% | -2.337% | -3.425% | -3.606% | -3.243% | -3.606% | -3.606% | -3.787% | -3.606% | -3.062% | -3.062% | -3.425% | -3.062% | -2.881% | -2.7% | -2.156% | -2.337% | -2.337% | -1.613% | -1.794% | -1.794% | -1.794% | -1.613% | -1.069% | -1.069% | -2.156% | -1.794% | -1.794% | -1.613% | -2.156% | -2.337% | -1.975% | -1.975% | -2.519% | -100% | -100% | -100% | -100% | -100%
|  | 100 | #REF! | 3.651% | 8.723% | 2.61% | 3.651% | 2.9% | -2.935% | -3.131% |  | -3.131% | -2.74% | -3.523% | -3.523% | -3.718% | -3.914% | -3.718% | -2.544% | -3.523% | -4.305% | -3.914% | -4.11% | -3.718% | -4.305% | -3.523% | -3.131% | -3.131% | -3.327% | -2.935% | -2.935% | -1.566% | -1.566% | -1.761% | -1.761% | -1.37% | -0.9785% | -1.37% | -1.174% | -1.566% | -0.1957% | -0.5871% | -1.566% | -1.957% | -1.566% | -1.174% | -1.761% | -1.37% | -1.37% | -2.153% | -1.957% | -100% | -100% | -100% | -100% | -100%
|  | 110 | #REF! | 3.573% | 8.787% | 2.9% | 3.799% | 3.191% | -3.45% | -3.029% |  | -3.66% | -2.819% | -3.87% | -3.66% | -3.45% | -3.87% | -3.87% | -2.608% | -3.45% | -3.87% | -4.081% | -4.081% | -3.87% | -3.87% | -3.45% | -3.239% | -2.819% | -3.239% | -3.239% | -3.029% | -1.557% | -1.977% | -1.557% | -1.346% | -1.136% | -1.346% | -1.136% | -1.346% | -1.346% | -0.7152% | -0.9255% | -1.346% | -1.557% | -1.557% | -1.136% | -1.767% | -1.767% | -1.346% | -1.346% | -1.557% | -100% | -100% | -100% | -100% | -100%
|  | 120 | #REF! | 3.171% | 9.24% | 2.46% | 3.651% | 2.814% | -2.401% | -3.074% |  | -2.849% | -2.849% | -3.298% | -3.523% | -3.298% | -3.747% | -3.747% | -2.401% | -3.074% | -3.523% | -3.298% | -3.523% | -3.747% | -3.523% | -3.523% | -3.074% | -2.849% | -3.074% | -2.401% | -2.625% | -1.503% | -1.503% | -0.6058% | -0.8302% | -0.3814% | -0.3814% | -0.8302% | -1.055% | -1.055% | 0.2917% | -0.1571% | -1.279% | -1.503% | -1.055% | -0.8302% | -1.279% | -1.055% | -1.279% | -1.279% | -1.728% | -100% | -100% | -100% | -100% | -100%
|  | 130 | #REF! | 4.453% | 7.667% | 3.424% | 4.194% | 3.373% | -3.072% | -3.548% |  | -3.548% | -3.787% | -4.025% | -4.025% | -4.025% | -4.739% | -4.263% | -2.834% | -3.787% | -4.501% | -4.025% | -4.501% | -4.263% | -4.501% | -4.263% | -3.787% | -3.31% | -3.787% | -3.787% | -3.548% | -1.881% | -2.596% | -1.405% | -1.643% | -1.405% | -1.881% | -1.881% | -1.643% | -1.881% | -0.9288% | -0.6906% | -2.12% | -1.881% | -2.12% | -1.881% | -2.12% | -2.12% | -2.12% | -1.881% | -2.12% | -100% | -100% | -100% | -100% | -100%
|  | 140 | #REF! | 5% | 4.726% | 2.846% | 4.453% | 4.372% | -3.765% | -3.765% |  | -4.263% | -4.014% | -4.513% | -4.513% | -4.513% | -5.261% | -5.011% | -3.266% | -4.762% | -4.762% | -4.762% | -5.011% | -4.762% | -4.762% | -4.513% | -4.014% | -4.263% | -4.014% | -3.765% | -4.014% | -2.767% | -2.767% | -2.269% | -2.269% | -2.019% | -1.272% | -2.269% | -2.019% | -2.269% | -0.5236% | -0.7729% | -3.017% | -3.266% | -3.017% | -2.518% | -3.017% | -2.767% | -2.518% | -2.767% | -3.266% | -100% | -100% | -100% | -100% | -100%
|  | 150 | #REF! | 4.158% | 3.875% | 2.761% | 4.441% | 3.651% | -3.209% | -3.992% |  | -3.992% | -3.992% | -4.253% | -4.513% | -4.774% | -5.557% | -5.035% | -3.209% | -4.253% | -4.774% | -4.774% | -4.513% | -4.513% | -4.513% | -4.253% | -3.731% | -3.731% | -4.253% | -3.47% | -3.731% | -2.165% | -2.687% | -1.644% | -1.644% | -1.122% | -1.122% | -2.165% | -1.383% | -2.165% | -0.07827% | -0.3392% | -3.209% | -3.209% | -2.948% | -2.165% | -2.948% | -2.165% | -2.165% | -2.687% | -2.948% | -100% | -100% | -100% | -100% | -100%
|  | 160 | #REF! | 3.257% | 0.6685% | 2.38% | 3.553% | 3.702% | -2.878% | -3.431% |  | -3.985% | -3.431% | -3.985% | -4.538% | -4.815% | -5.645% | -5.091% | -3.154% | -4.538% | -4.815% | -4.538% | -4.538% | -4.538% | -4.538% | -3.985% | -3.154% | -3.708% | -3.985% | -3.154% | -3.708% | -2.324% | -2.324% | -0.9408% | -1.217% | -0.6641% | -0.3874% | -1.771% | -1.217% | -1.771% | 0.9961% | 0.4427% | -3.431% | -3.708% | -3.154% | -2.324% | -3.431% | -2.601% | -2.601% | -2.878% | -3.431% | -100% | -100% | -100% | -100% | -100%
|  | 170 | #REF! | 2.985% | 0% | 2.071% | 4.23% | 3.573% | -3.188% | -2.899% |  | -3.768% | -3.768% | -3.768% | -4.348% | -5.217% | -6.087% | -5.507% | -3.188% | -4.638% | -5.507% | -4.348% | -4.638% | -4.638% | -4.348% | -4.058% | -3.478% | -3.478% | -4.058% | -3.188% | -3.768% | -2.029% | -2.319% | -0.8696% | -0.8696% | -0.5797% | 0% | -1.739% | -1.159% | -1.449% | 1.739% | 0.8696% | -3.768% | -4.058% | -3.768% | -2.319% | -3.768% | -2.609% | -2.319% | -3.188% | -3.768% | -100% | -100% | -100% | -100% | -100%
|  | 180 | #REF! | 2.897% | -4.261% | 2.26% | 3.868% | 3.77% | -3.118% | -3.118% |  | -4.027% | -3.421% | -3.724% | -4.632% | -5.54% | -6.751% | -5.843% | -3.118% | -4.935% | -5.54% | -4.329% | -4.935% | -4.632% | -4.935% | -4.027% | -3.421% | -3.421% | -4.632% | -2.816% | -3.421% | -1.605% | -1.907% | -0.3936% | -0.3936% | -0.09083% | 0.2119% | -1.605% | -0.6963% | -1.302% | 2.331% | 2.028% | -4.632% | -4.935% | -4.027% | -2.816% | -4.329% | -2.816% | -2.513% | -3.724% | -3.724% | -100% | -100% | -100% | -100% | -100%
|  | 190 | #REF! | 1.079% | -5.519% | 1.401% | 2.379% | 3.009% | -2.638% | -2.952% |  | -3.58% | -3.266% | -2.952% | -3.894% | -5.779% | -6.721% | -6.093% | -2.324% | -4.837% | -5.151% | -3.894% | -3.894% | -4.209% | -4.209% | -3.266% | -2.952% | -2.952% | -3.894% | -2.01% | -3.266% | -1.696% | -1.696% | 0.1884% | 0.1884% | 1.131% | 1.445% | -0.7538% | 0.1884% | -1.068% | 3.957% | 2.701% | -4.523% | -4.837% | -3.58% | -2.01% | -3.894% | -1.068% | -1.696% | -2.952% | -4.209% | -100% | -100% | -100% | -100% | -100%
|  | 200 | #REF! | 0.2273% | -5.884% | 0.2273% | 2.219% | 2.49% | -2.818% | -1.846% |  | -3.79% | -3.466% | -2.818% | -3.142% | -5.734% | -7.029% | -6.382% | -2.17% | -4.438% | -4.762% | -3.466% | -4.114% | -3.79% | -4.438% | -3.142% | -2.17% | -2.17% | -3.466% | -0.5507% | -3.142% | -1.199% | -1.199% | 0.7451% | 1.069% | 2.365% | 2.689% | 0.09718% | 0.7451% | -0.5507% | 5.604% | 4.308% | -4.114% | -5.086% | -3.466% | -1.199% | -3.466% | -0.2268% | -0.8746% | -2.494% | -3.79% | -100% | -100% | -100% | -100% | -100%
|  | 210 | #REF! | -2.467% | -4.045% | -4.355% | -0.1684% | 0.5085% | -1.18% | -1.18% |  | -1.855% | -1.18% | 0.1686% | -1.855% | -4.553% | -5.902% | -5.228% | -0.5059% | -2.192% | -2.867% | -1.518% | -2.192% | -2.192% | -1.855% | -1.18% | -0.1686% | -0.1686% | -1.855% | 1.518% | -1.855% | 0.8432% | 1.18% | 4.216% | 4.553% | 4.89% | 5.228% | 2.53% | 3.204% | 1.518% | 8.938% | 6.914% | -2.192% | -2.53% | -1.18% | 1.18% | -1.855% | 2.867% | 1.518% | -0.5059% | -1.518% | -100% | -100% | -100% | -100% | -100%
|  | 220 | #REF! | -4.053% | -3.411% | -3.411% | -0.756% | 0.4522% | 1.108% | -0.9695% |  | -0.9695% | 0.7618% | 1.454% | 0.4155% | -3.047% | -4.432% | -3.74% | 1.108% | -0.6233% | -1.316% | 1.108% | -0.277% | -0.277% | 0.06925% | 1.108% | 2.147% | 1.801% | 0.4155% | 4.224% | -0.277% | 3.186% | 3.532% | 6.302% | 6.994% | 8.033% | 9.072% | 6.302% | 5.263% | 2.839% | 12.88% | 11.15% | 1.454% | 0.7618% | 2.147% | 4.224% | 1.454% | 6.648% | 5.263% | 2.493% | 2.147% | -100% | -100% | -100% | -100% | -100%
|  | 226 | #REF! | -3.654% | -2.685% | -2.685% | -0.6849% | 0.1727% | -0.3448% | 0% |  | -1.034% | 1.034% | 3.448% | -0.3448% | -2.759% | -5.172% | -3.793% | 1.379% | -0.3448% | -1.379% | 1.379% | 0% | 0% | 0.3448% | 1.379% | 2.759% | 2.759% | 0.3448% | 5.172% | 0.3448% | 3.103% | 3.793% | 7.241% | 7.586% | 8.621% | 10% | 7.241% | 6.552% | 2.759% | 12.41% | 10.69% | 2.414% | 1.724% | 3.793% | 5.517% | 2.414% | 7.586% | 5.517% | 3.793% | 3.103% | -100% | -100% | -100% | -100% | -100%

## GTR1 G315 z0 y
GTR1 Spot size in y axis at iso for G315º every 10 MeV
Lynx PT
warrning | 8%
action level | 10%
validation of new set point for deflector
|  | E [MeV] | REF |  |  |  | after reoptim | [12 empty] | After deflector adjustment due to IDDC measurements showing deviations |  |  | SMPS replacement |  | cyclo PM_run2 | [5 empty] | Lynx SN 63057 | Lynx SN 63057 | Lynx SN 63021 | Lynx SN 63021 |  |  |  | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021
 | time axis 2019-01-10 to 2023-06-20
|  |  |  | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm]
|  | 60 | 7.165 | 7.13 | 7.31 | 7.2 | 7.26 | 7.212 | 7.49 | 7.23 | 7.22 | 7.12 | 7.14 | 7.07 | 7.13 | 7.12 | 7.08 | 7.15 | 7.24 | 7.09 | 7.13 | 7.16 | 7.13 | 7.14 | 7.13 | 7.11 | 7.12 | 7.16 | 7.17 | 7.17 | 7.13 | 7.17 | 7.15 | 7.1 | 7.12 | 7.13 | 7.16 | 7.19 | 7.2 | 7.13 | 7.16 | 7.17 | 7.23 | 7.16 | 7.15 | 7.2 | 7.2 | 7.18 | 7.15 | 7.18 | 7.23 | 7.2
|  | 70 | 6.481 | 6.61 | 6.55 | 6.63 | 6.62 | 6.607 | 6.53 | 6.62 | 6.61 | 6.56 | 6.57 | 6.61 | 6.55 | 6.56 | 6.52 | 6.55 | 6.64 | 6.56 | 6.58 | 6.55 | 6.55 | 6.52 | 6.49 | 6.53 | 6.55 | 6.61 | 6.58 | 6.56 | 6.56 | 6.56 | 6.6 | 6.55 | 6.57 | 6.57 | 6.59 | 6.61 | 6.58 | 6.61 | 6.54 | 6.57 | 6.62 | 6.59 | 6.57 | 6.59 | 6.61 | 6.58 | 6.58 | 6.58 | 6.6 | 6.6
|  | 80 | 5.679 | 5.68 | 5.83 | 5.69 | 5.67 | 5.673 | 5.78 | 5.72 | 5.67 | 5.62 | 5.62 | 5.57 | 5.62 | 5.66 | 5.61 | 5.62 | 5.67 | 5.61 | 5.62 | 5.59 | 5.62 | 5.67 | 5.64 | 5.62 | 5.66 | 5.65 | 5.66 | 5.62 | 5.57 | 5.64 | 5.66 | 5.63 | 5.65 | 5.66 | 5.67 | 5.68 | 5.66 | 5.66 | 5.66 | 5.68 | 5.67 | 5.62 | 5.63 | 5.67 | 5.66 | 5.64 | 5.64 | 5.66 | 5.66 | 5.65
|  | 90 | 5.453 | 5.33 | 5.28 | 5.36 | 5.31 | 5.331 | 5.19 | 5.33 | 5.3 | 5.26 | 5.29 | 5.26 | 5.25 | 5.29 | 5.27 | 5.28 | 5.31 | 5.25 | 5.25 | 5.29 | 5.3 | 5.3 | 5.25 | 5.29 | 5.26 | 5.31 | 5.28 | 5.27 | 5.25 | 5.28 | 5.28 | 5.29 | 5.28 | 5.29 | 5.29 | 5.3 | 5.31 | 5.28 | 5.27 | 5.32 | 5.31 | 5.28 | 5.29 | 5.3 | 5.32 | 5.29 | 5.27 | 5.33 | 5.29 | 5.28
|  | 100 | 5.097 | 4.96 | 5.22 | 5.01 | 5.01 | 4.982 | 5.37 | 5.01 | 4.97 | 4.94 | 4.94 | 4.91 | 4.91 | 4.94 | 4.92 | 4.92 | 4.99 | 4.94 | 4.93 | 4.93 | 4.92 | 4.92 | 4.91 | 4.94 | 4.94 | 4.95 | 4.92 | 4.92 | 4.93 | 4.91 | 4.95 | 4.94 | 4.94 | 4.95 | 4.95 | 4.96 | 4.95 | 4.96 | 4.93 | 4.98 | 4.97 | 4.94 | 4.92 | 4.94 | 4.96 | 4.95 | 4.93 | 4.95 | 4.93 | 4.95
|  | 110 | 4.739 | 4.65 | 4.83 | 4.65 | 4.64 | 4.632 | 5.01 | 4.65 | 4.63 | 4.6 | 4.6 | 4.6 | 4.59 | 4.61 | 4.56 | 4.6 | 4.64 | 4.6 | 4.57 | 4.59 | 4.57 | 4.58 | 4.58 | 4.6 | 4.61 | 4.6 | 4.6 | 4.61 | 4.58 | 4.58 | 4.6 | 4.6 | 4.62 | 4.62 | 4.63 | 4.62 | 4.63 | 4.63 | 4.6 | 4.64 | 4.64 | 4.62 | 4.59 | 4.62 | 4.64 | 4.61 | 4.6 | 4.62 | 4.62 | 4.62
|  | 120 | 4.411 | 4.35 | 4.58 | 4.4 | 4.33 | 4.357 | 4.65 | 4.37 | 4.31 | 4.28 | 4.29 | 4.28 | 4.26 | 4.28 | 4.25 | 4.27 | 4.33 | 4.27 | 4.29 | 4.28 | 4.29 | 4.28 | 4.25 | 4.28 | 4.29 | 4.3 | 4.29 | 4.29 | 4.3 | 4.27 | 4.28 | 4.26 | 4.3 | 4.3 | 4.31 | 4.31 | 4.31 | 4.3 | 4.27 | 4.36 | 4.35 | 4.27 | 4.26 | 4.3 | 4.32 | 4.3 | 4.29 | 4.3 | 4.29 | 4.28
|  | 130 | 4.145 | 4.06 | 4.04 | 4.1 | 4.04 | 4.059 | 4 | 4.08 | 4.02 | 3.98 | 4.01 | 4 | 3.99 | 4.03 | 3.96 | 3.98 | 4.07 | 3.99 | 3.96 | 4.02 | 4 | 3.98 | 3.97 | 4.01 | 4.01 | 4.02 | 4.02 | 4.02 | 4 | 4 | 4 | 3.99 | 4.02 | 4.03 | 4.05 | 4.03 | 4.04 | 4.03 | 3.99 | 4.1 | 4.08 | 4 | 4 | 4.01 | 4.04 | 4.01 | 4.02 | 4.01 | 4.01 | 4.01
|  | 140 | 3.919 | 3.89 | 3.86 | 3.92 | 3.86 | 3.846 | 3.77 | 3.89 | 3.85 | 3.78 | 3.79 | 3.8 | 3.79 | 3.81 | 3.76 | 3.77 | 3.85 | 3.79 | 3.8 | 3.79 | 3.81 | 3.79 | 3.78 | 3.8 | 3.82 | 3.84 | 3.81 | 3.82 | 3.82 | 3.79 | 3.79 | 3.78 | 3.81 | 3.81 | 3.85 | 3.86 | 3.84 | 3.83 | 3.79 | 3.92 | 3.9 | 3.78 | 3.77 | 3.8 | 3.84 | 3.79 | 3.8 | 3.8 | 3.8 | 3.79
|  | 150 | 3.748 | 3.73 | 3.63 | 3.76 | 3.7 | 3.684 | 3.56 | 3.7 | 3.67 | 3.6 | 3.63 | 3.62 | 3.59 | 3.63 | 3.58 | 3.6 | 3.69 | 3.62 | 3.64 | 3.64 | 3.64 | 3.63 | 3.6 | 3.64 | 3.65 | 3.67 | 3.66 | 3.63 | 3.64 | 3.63 | 3.62 | 3.61 | 3.65 | 3.66 | 3.69 | 3.69 | 3.68 | 3.68 | 3.62 | 3.76 | 3.74 | 3.62 | 3.6 | 3.63 | 3.67 | 3.62 | 3.64 | 3.65 | 3.62 | 3.62
|  | 160 | 3.552 | 3.56 | 3.41 | 3.58 | 3.49 | 3.469 | 3.3 | 3.53 | 3.48 | 3.4 | 3.42 | 3.42 | 3.41 | 3.42 | 3.38 | 3.39 | 3.47 | 3.42 | 3.42 | 3.43 | 3.43 | 3.42 | 3.39 | 3.44 | 3.45 | 3.47 | 3.44 | 3.43 | 3.44 | 3.43 | 3.43 | 3.41 | 3.46 | 3.48 | 3.5 | 3.51 | 3.5 | 3.49 | 3.43 | 3.58 | 3.56 | 3.4 | 3.39 | 3.42 | 3.47 | 3.42 | 3.43 | 3.43 | 3.42 | 3.41
|  | 170 | 3.358 | 3.38 | 3.26 | 3.39 | 3.31 | 3.293 | 3.12 | 3.35 | 3.3 | 3.22 | 3.25 | 3.26 | 3.22 | 3.25 | 3.19 | 3.21 | 3.28 | 3.23 | 3.25 | 3.27 | 3.26 | 3.24 | 3.21 | 3.25 | 3.28 | 3.29 | 3.28 | 3.26 | 3.27 | 3.24 | 3.25 | 3.24 | 3.3 | 3.28 | 3.33 | 3.34 | 3.32 | 3.31 | 3.25 | 3.42 | 3.41 | 3.23 | 3.22 | 3.24 | 3.29 | 3.25 | 3.26 | 3.27 | 3.25 | 3.24
|  | 180 | 3.246 | 3.29 | 3.11 | 3.3 | 3.21 | 3.19 | 2.95 | 3.23 | 3.2 | 3.1 | 3.14 | 3.14 | 3.11 | 3.12 | 3.07 | 3.09 | 3.17 | 3.13 | 3.13 | 3.15 | 3.15 | 3.13 | 3.09 | 3.15 | 3.17 | 3.18 | 3.16 | 3.14 | 3.17 | 3.13 | 3.13 | 3.11 | 3.18 | 3.18 | 3.22 | 3.22 | 3.2 | 3.19 | 3.13 | 3.33 | 3.29 | 3.11 | 3.1 | 3.12 | 3.17 | 3.13 | 3.15 | 3.16 | 3.13 | 3.12
|  | 190 | 3.142 | 3.2 | 2.99 | 3.21 | 3.12 | 3.09 | 2.81 | 3.12 | 3.1 | 3 | 3.02 | 3.03 | 2.99 | 3 | 2.96 | 2.97 | 3.07 | 3.02 | 3.03 | 3.05 | 3.04 | 3.03 | 2.99 | 3.04 | 3.06 | 3.08 | 3.07 | 3.03 | 3.07 | 3.02 | 3.01 | 3.01 | 3.06 | 3.07 | 3.12 | 3.12 | 3.09 | 3.1 | 3.02 | 3.23 | 3.21 | 3 | 2.98 | 3.02 | 3.08 | 3.03 | 3.08 | 3.06 | 3.03 | 3.01
|  | 200 | 3.059 | 3.2 | 2.94 | 3.2 | 3.09 | 3.039 | 2.66 | 3.09 | 3.06 | 2.92 | 2.95 | 2.98 | 2.92 | 2.9 | 2.86 | 2.88 | 3 | 2.95 | 2.95 | 2.99 | 2.96 | 2.96 | 2.91 | 2.95 | 2.99 | 3.01 | 2.99 | 2.96 | 3.03 | 2.92 | 2.94 | 2.93 | 3.02 | 3.02 | 3.06 | 3.07 | 3.04 | 3.02 | 2.93 | 3.24 | 3.19 | 2.9 | 2.89 | 2.93 | 3.01 | 2.94 | 3 | 2.99 | 2.95 | 2.93
|  | 210 | 2.929 | 3.15 | 2.91 | 3.19 | 3.01 | 2.951 | 2.52 | 3 | 2.95 | 2.79 | 2.83 | 2.88 | 2.82 | 2.76 | 2.72 | 2.75 | 2.86 | 2.82 | 2.83 | 2.86 | 2.84 | 2.82 | 2.78 | 2.84 | 2.88 | 2.9 | 2.89 | 2.84 | 2.93 | 2.8 | 2.8 | 2.8 | 2.91 | 2.9 | 2.98 | 2.98 | 2.95 | 2.91 | 2.8 | 3.22 | 3.15 | 2.77 | 2.76 | 2.81 | 2.89 | 2.81 | 2.9 | 2.88 | 2.82 | 2.8
|  | 220 | 2.831 | 3.18 | 2.97 | 3.18 | 2.97 | 2.904 | 2.51 | 2.99 | 2.91 | 2.7 | 2.77 | 2.83 | 2.75 | 2.68 | 2.63 | 2.66 | 2.79 | 2.74 | 2.75 | 2.8 | 2.76 | 2.75 | 2.69 | 2.78 | 2.82 | 2.84 | 2.83 | 2.78 | 2.9 | 2.71 | 2.73 | 2.74 | 2.89 | 2.87 | 2.97 | 2.96 | 2.93 | 2.88 | 2.71 | 3.33 | 3.26 | 2.74 | 2.72 | 2.79 | 2.85 | 2.75 | 2.87 | 2.85 | 2.77 | 2.78
|  | 226 | 2.76 | 3.19 | 3.01 | 3.17 | 2.95 | 2.895 | 2.52 | 2.96 | 2.94 | 2.69 | 2.77 | 2.85 | 2.72 | 2.65 | 2.59 | 2.64 | 2.79 | 2.73 | 2.73 | 2.8 | 2.76 | 2.73 | 2.66 | 2.77 | 2.81 | 2.84 | 2.82 | 2.76 | 2.89 | 2.71 | 2.7 | 2.72 | 2.89 | 2.89 | 2.97 | 2.96 | 2.94 | 2.89 | 2.7 | 3.4 | 3.32 | 2.72 | 2.7 | 2.77 | 2.83 | 2.72 | 2.85 | 2.85 | 2.75 | 2.76
Deviation from reference [%]
|  | E [MeV] | OOS | 2020-11-18: After deflector adjustment due to IDDC measurements showing deviations; 2021-01-11: SMPS replacement; 2021-02-14: cyclo PM_run2
|  |  | 2018-05-29 | [5 empty] | 2019-09-11 | 2019-11-12 | 2020-02-02 |  | 2020-03-25 | 2020-05-05 | 2020-06-02 | 2020-07-28 | 2020-09-01 | 2020-10-06 | 2020-10-11 | 2020-11-18 | 2020-12-10 | 2021-01-05 | 2021-01-11 | 2021-02-04 | 2021-02-14 | 2021-03-09 | 2021-04-06 | 2021-06-22 | 2021-07-20 | 2021-08-17 | 2021-10-12 | 2021-10-17 | 2021-11-08 | 2021-12-09 | 2022-01-21 | 2022-02-08 | 2022-03-08 | 2022-03-13 | 2022-04-04 | 2022-05-16 | 2022-07-08 | 2022-07-12 | 2022-08-30 | 2022-10-09 | 2022-11-08 | 2022-12-01 | 2023-02-04 | 2023-03-10 | 2023-03-20 | 2023-04-22 | 2023-05-27 | 2023-06-20 | 0 | 0
|  |  | sigma y [mm] | [5 empty] | sigma y [mm] | sigma y [mm] | sigma y [mm] |  | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm]
|  | 60 | #REF! | [5 empty] | 0.9072% | 0.7676% | -0.6281% |  | -0.3489% | -1.326% | -0.4885% | -0.6281% | -1.186% | -0.2094% | 1.047% | -1.047% | -0.4885% | -0.06978% | -0.4885% | -0.3489% | -0.4885% | -0.7676% | -0.6281% | -0.06978% | 0.06978% | 0.06978% | -0.4885% | 0.06978% | -0.2094% | -0.9072% | -0.6281% | -0.4885% | -0.06978% | 0.3489% | 0.4885% | -0.4885% | -0.06978% | 0.06978% | 0.9072% | -0.06978% | -0.2094% | 0.4885% | 0.4885% | 0.2094% | -0.2094% | 0.2094% | 0.9072% | 0.4885% | -100% | -100%
|  | 70 | #REF! | [5 empty] | 2.145% | 1.99% | 1.219% |  | 1.373% | 1.99% | 1.065% | 1.219% | 0.6018% | 1.065% | 2.453% | 1.219% | 1.528% | 1.065% | 1.065% | 0.6018% | 0.1389% | 0.7561% | 1.065% | 1.99% | 1.528% | 1.219% | 1.219% | 1.219% | 1.836% | 1.065% | 1.373% | 1.373% | 1.682% | 1.99% | 1.528% | 1.99% | 0.9104% | 1.373% | 2.145% | 1.682% | 1.373% | 1.682% | 1.99% | 1.528% | 1.528% | 1.528% | 1.836% | 1.836% | -100% | -100%
|  | 80 | #REF! | [5 empty] | 0.722% | -0.1585% | -1.039% |  | -1.039% | -1.919% | -1.039% | -0.3346% | -1.215% | -1.039% | -0.1585% | -1.215% | -1.039% | -1.567% | -1.039% | -0.1585% | -0.6867% | -1.039% | -0.3346% | -0.5107% | -0.3346% | -1.039% | -1.919% | -0.6867% | -0.3346% | -0.8628% | -0.5107% | -0.3346% | -0.1585% | 0.01761% | -0.3346% | -0.3346% | -0.3346% | 0.01761% | -0.1585% | -1.039% | -0.8628% | -0.1585% | -0.3346% | -0.6867% | -0.6867% | -0.3346% | -0.3346% | -0.5107% | -100% | -100%
|  | 90 | #REF! | [5 empty] | -2.256% | -2.806% | -3.539% |  | -2.989% | -3.539% | -3.723% | -2.989% | -3.356% | -3.173% | -2.622% | -3.723% | -3.723% | -2.989% | -2.806% | -2.806% | -3.723% | -2.989% | -3.539% | -2.622% | -3.173% | -3.356% | -3.723% | -3.173% | -3.173% | -2.989% | -3.173% | -2.989% | -2.989% | -2.806% | -2.622% | -3.173% | -3.356% | -2.439% | -2.622% | -3.173% | -2.989% | -2.806% | -2.439% | -2.989% | -3.356% | -2.256% | -2.989% | -3.173% | -100% | -100%
|  | 100 | #REF! | [5 empty] | -1.707% | -2.492% | -3.08% |  | -3.08% | -3.669% | -3.669% | -3.08% | -3.473% | -3.473% | -2.099% | -3.08% | -3.276% | -3.276% | -3.473% | -3.473% | -3.669% | -3.08% | -3.08% | -2.884% | -3.473% | -3.473% | -3.276% | -3.669% | -2.884% | -3.08% | -3.08% | -2.884% | -2.884% | -2.688% | -2.884% | -2.688% | -3.276% | -2.295% | -2.492% | -3.08% | -3.473% | -3.08% | -2.688% | -2.884% | -3.276% | -2.884% | -3.276% | -2.884% | -100% | -100%
|  | 110 | #REF! | [5 empty] | -1.878% | -2.3% | -2.933% |  | -2.933% | -2.933% | -3.144% | -2.722% | -3.777% | -2.933% | -2.089% | -2.933% | -3.566% | -3.144% | -3.566% | -3.355% | -3.355% | -2.933% | -2.722% | -2.933% | -2.933% | -2.722% | -3.355% | -3.355% | -2.933% | -2.933% | -2.511% | -2.511% | -2.3% | -2.511% | -2.3% | -2.3% | -2.933% | -2.089% | -2.089% | -2.511% | -3.144% | -2.511% | -2.089% | -2.722% | -2.933% | -2.511% | -2.511% | -2.511% | -100% | -100%
|  | 120 | #REF! | [5 empty] | -0.9295% | -2.29% | -2.97% |  | -2.743% | -2.97% | -3.423% | -2.97% | -3.65% | -3.197% | -1.836% | -3.197% | -2.743% | -2.97% | -2.743% | -2.97% | -3.65% | -2.97% | -2.743% | -2.516% | -2.743% | -2.743% | -2.516% | -3.197% | -2.97% | -3.423% | -2.516% | -2.516% | -2.29% | -2.29% | -2.29% | -2.516% | -3.197% | -1.156% | -1.383% | -3.197% | -3.423% | -2.516% | -2.063% | -2.516% | -2.743% | -2.516% | -2.743% | -2.97% | -100% | -100%
|  | 130 | #REF! | [5 empty] | -1.568% | -3.016% | -3.981% |  | -3.257% | -3.498% | -3.739% | -2.774% | -4.463% | -3.981% | -1.809% | -3.739% | -4.463% | -3.016% | -3.498% | -3.981% | -4.222% | -3.257% | -3.257% | -3.016% | -3.016% | -3.016% | -3.498% | -3.498% | -3.498% | -3.739% | -3.016% | -2.774% | -2.292% | -2.774% | -2.533% | -2.774% | -3.739% | -1.086% | -1.568% | -3.498% | -3.498% | -3.257% | -2.533% | -3.257% | -3.016% | -3.257% | -3.257% | -3.257% | -100% | -100%
|  | 140 | #REF! | [5 empty] | -0.74% | -1.761% | -3.547% |  | -3.292% | -3.036% | -3.292% | -2.781% | -4.057% | -3.802% | -1.761% | -3.292% | -3.036% | -3.292% | -2.781% | -3.292% | -3.547% | -3.036% | -2.526% | -2.016% | -2.781% | -2.526% | -2.526% | -3.292% | -3.292% | -3.547% | -2.781% | -2.781% | -1.761% | -1.505% | -2.016% | -2.271% | -3.292% | 0.02552% | -0.4848% | -3.547% | -3.802% | -3.036% | -2.016% | -3.292% | -3.036% | -3.036% | -3.036% | -3.292% | -100% | -100%
|  | 150 | #REF! | [5 empty] | -1.281% | -2.081% | -3.949% |  | -3.148% | -3.415% | -4.216% | -3.148% | -4.482% | -3.949% | -1.547% | -3.415% | -2.882% | -2.882% | -2.882% | -3.148% | -3.949% | -2.882% | -2.615% | -2.081% | -2.348% | -3.148% | -2.882% | -3.148% | -3.415% | -3.682% | -2.615% | -2.348% | -1.547% | -1.547% | -1.814% | -1.814% | -3.415% | 0.3202% | -0.2134% | -3.415% | -3.949% | -3.148% | -2.081% | -3.415% | -2.882% | -2.615% | -3.415% | -3.415% | -100% | -100%
|  | 160 | #REF! | [5 empty] | -0.6194% | -2.027% | -4.279% |  | -3.716% | -3.716% | -3.998% | -3.716% | -4.842% | -4.561% | -2.309% | -3.716% | -3.716% | -3.435% | -3.435% | -3.716% | -4.561% | -3.153% | -2.872% | -2.309% | -3.153% | -3.435% | -3.153% | -3.435% | -3.435% | -3.998% | -2.59% | -2.027% | -1.464% | -1.182% | -1.464% | -1.745% | -3.435% | 0.7883% | 0.2252% | -4.279% | -4.561% | -3.716% | -2.309% | -3.716% | -3.435% | -3.435% | -3.716% | -3.998% | -100% | -100%
|  | 170 | #REF! | [5 empty] | -0.2382% | -1.727% | -4.11% |  | -3.216% | -2.918% | -4.11% | -3.216% | -5.003% | -4.407% | -2.323% | -3.812% | -3.216% | -2.621% | -2.918% | -3.514% | -4.407% | -3.216% | -2.323% | -2.025% | -2.323% | -2.918% | -2.621% | -3.514% | -3.216% | -3.514% | -1.727% | -2.323% | -0.8338% | -0.536% | -1.132% | -1.429% | -3.216% | 1.846% | 1.549% | -3.812% | -4.11% | -3.514% | -2.025% | -3.216% | -2.918% | -2.621% | -3.216% | -3.514% | -100% | -100%
|  | 180 | #REF! | [5 empty] | -0.4929% | -1.417% | -4.498% |  | -3.266% | -3.266% | -4.19% | -3.882% | -5.422% | -4.806% | -2.341% | -3.574% | -3.574% | -2.957% | -2.957% | -3.574% | -4.806% | -2.957% | -2.341% | -2.033% | -2.649% | -3.266% | -2.341% | -3.574% | -3.574% | -4.19% | -2.033% | -2.033% | -0.801% | -0.801% | -1.417% | -1.725% | -3.574% | 2.588% | 1.356% | -4.19% | -4.498% | -3.882% | -2.341% | -3.574% | -2.957% | -2.649% | -3.574% | -3.882% | -100% | -100%
|  | 190 | #REF! | [5 empty] | -0.7002% | -1.337% | -4.519% |  | -3.883% | -3.565% | -4.838% | -4.519% | -5.792% | -5.474% | -2.292% | -3.883% | -3.565% | -2.928% | -3.246% | -3.565% | -4.838% | -3.246% | -2.61% | -1.973% | -2.292% | -3.565% | -2.292% | -3.883% | -4.201% | -4.201% | -2.61% | -2.292% | -0.7002% | -0.7002% | -1.655% | -1.337% | -3.883% | 2.801% | 2.164% | -4.519% | -5.156% | -3.883% | -1.973% | -3.565% | -1.973% | -2.61% | -3.565% | -4.201% | -100% | -100%
|  | 200 | #REF! | [5 empty] | 1.013% | 0.03269% | -4.544% |  | -3.563% | -2.583% | -4.544% | -5.198% | -6.505% | -5.852% | -1.929% | -3.563% | -3.563% | -2.256% | -3.236% | -3.236% | -4.871% | -3.563% | -2.256% | -1.602% | -2.256% | -3.236% | -0.948% | -4.544% | -3.89% | -4.217% | -1.275% | -1.275% | 0.03269% | 0.3596% | -0.6211% | -1.275% | -4.217% | 5.917% | 4.282% | -5.198% | -5.525% | -4.217% | -1.602% | -3.89% | -1.929% | -2.256% | -3.563% | -4.217% | -100% | -100%
|  | 210 | #REF! | [5 empty] | 2.424% | 0.717% | -4.746% |  | -3.38% | -1.673% | -3.721% | -5.77% | -7.136% | -6.111% | -2.356% | -3.721% | -3.38% | -2.356% | -3.039% | -3.721% | -5.087% | -3.039% | -1.673% | -0.9901% | -1.332% | -3.039% | 0.03414% | -4.404% | -4.404% | -4.404% | -0.6487% | -0.9901% | 1.741% | 1.741% | 0.717% | -0.6487% | -4.404% | 9.935% | 7.545% | -5.428% | -5.77% | -4.063% | -1.332% | -4.063% | -0.9901% | -1.673% | -3.721% | -4.404% | -100% | -100%
|  | 220 | #REF! | [5 empty] | 5.616% | 2.791% | -4.627% |  | -2.155% | -0.03532% | -2.861% | -5.334% | -7.1% | -6.04% | -1.448% | -3.214% | -2.861% | -1.095% | -2.508% | -2.861% | -4.981% | -1.801% | -0.3886% | 0.3179% | -0.03532% | -1.801% | 2.437% | -4.274% | -3.568% | -3.214% | 2.084% | 1.378% | 4.91% | 4.557% | 3.497% | 1.731% | -4.274% | 17.63% | 15.15% | -3.214% | -3.921% | -1.448% | 0.6711% | -2.861% | 1.378% | 0.6711% | -2.155% | -1.801% | -100% | -100%
|  | 226 | #REF! | [5 empty] | 7.246% | 6.522% | -2.536% |  | 0.3623% | 3.261% | -1.449% | -3.986% | -6.159% | -4.348% | 1.087% | -1.087% | -1.087% | 1.449% | 0% | -1.087% | -3.623% | 0.3623% | 1.812% | 2.899% | 2.174% | 0% | 4.71% | -1.812% | -2.174% | -1.449% | 4.71% | 4.71% | 7.609% | 7.246% | 6.522% | 4.71% | -2.174% | 23.19% | 20.29% | -1.449% | -2.174% | 0.3623% | 2.536% | -1.449% | 3.261% | 3.261% | -0.3623% | 0% | -100% | -100%

## GTR1 Skewness
GTR1 skewness at iso for G315º every 10 MeV
Lynx PT
warrning
action level
Sel parameter [%] for ellipse
[14 empty] | After deflector adjustment due to IDDC measurements showing deviations |  |  | SMPS replacement | cyclo PM_run1 | cyclo PM_run2 | [5 empty] | RoomMatching/OldLynx |  |  | Lynx SN 63021 | Lynx SN 63021 | [6 empty] | validation of new set point for deflector
| E [MeV] | 2019-01-10 | 2019-08-06 | 2019-11-12 | 2020-02-02 | 2020-03-25 | 2020-05-05 | 2020-06-02 | 2020-07-28 | 2020-09-01 | 2020-10-06 | 2020-10-11 | 2020-11-18 | 2020-11-18 | 2020-12-10 | 2021-01-05 | 2021-01-11 | 2021-02-04 | 2021-02-14 | 2021-03-09 | 2021-04-06 | 2021-06-22 | 2021-07-20 | 2021-08-17 | 2021-09-21 | 2021-10-12 | 2021-10-17 | 2021-11-08 | 2021-12-09 | 2022-01-21 | 2022-02-08 | 2022-03-08 | 2022-03-13 | 2022-04-04 | 2022-05-16 | 2022-07-08 | 2022-07-12 | 2022-08-30 | 2022-10-09 | 2022-11-08 | 2022-12-01 | 2023-02-04 | 2023-03-10 | 2023-03-20 | 2023-04-22 | 2023-05-27 | 2023-06-20
| 60 | 3.8 | 3.5 | 4.1 | 3.7 | 3.6 | 4.5 | 3.8 | 3.9 | 3.6 | 4 | 4.1 | 3.8 | 3.5 | 3.7 | 3.9 | 3.8 | 3.8 | 4.6 | 3.706 | 3.741 | 3.776 | 4.019 | 4.277 | 0.406 | 4.512 | 4.46 | 4.246 | 4.036 | 4.14 | 4.294 | 4.244 | 4.405 | 4.609 | 4.151 | 4.333 | 4.332 | 4.901 | 4.36 | 4.465 | 4.565 | 4.606 | 4.453 | 4.411 | 4.59 | 5.111 | 4.79
| 70 | 4.9 | 4.1 | 4.8 | 4.5 | 4.6 | 5 | 4.6 | 4.8 | 4.5 | 4.6 | 4.7 | 4.7 | 4.6 | 5 | 4.5 | 4.6 | 4.2 | 4.1 | 4.581 | 4.559 | 4.769 | 4.839 | 4.895 | 3.42 | 4.969 | 4.888 | 5.085 | 4.787 | 4.874 | 4.858 | 4.826 | 4.86 | 4.94 | 4.799 | 4.496 | 4.356 | 5.004 | 5.259 | 5.218 | 5.159 | 5.206 | 5.029 | 5.028 | 5.102 | 5.373 | 5.488
| 80 | 3.7 | 3.4 | 3.5 | 3.8 | 3.8 | 4.4 | 4 | 4.3 | 4 | 3.9 | 3.4 | 3.9 | 3.7 | 3.6 | 3.5 | 3.9 | 4 | 1.3 | 4.017 | 4.132 | 3.776 | 4.193 | 4.227 | 1.386 | 4.425 | 4.64 | 5.176 | 5.003 | 5.027 | 5.225 | 4.99 | 4.922 | 5.095 | 4.808 | 5.07 | 4.728 | 4.848 | 5.266 | 5.55 | 5.502 | 4.889 | 5.098 | 5.001 | 5.188 | 5.281 | 5.403
| 90 | 1 | 0.7 | 0.8 | 1 | 1.3 | 1.9 | 1.3 | 1.6 | 1.4 | 1.1 | 1.1 | 1.1 | 1.1 | 1 | 1.1 | 1.3 | 1.1 | 5.3 | 1.369 | 1.264 | 1.216 | 1.33 | 1.29 | 0.806 | 1.945 | 1.795 | 2.356 | 2.442 | 2.286 | 2.352 | 2.288 | 2.098 | 2.387 | 2.269 | 2.362 | 2.054 | 2.252 | 2.679 | 2.841 | 2.79 | 2.329 | 2.356 | 2.2 | 2.743 | 2.531 | 2.831
| 100 | 5.1 | 4.8 | 5 | 5.2 | 5.3 | 6.2 | 5.3 | 5.9 | 5.6 | 5.5 | 5.1 | 4.9 | 5.3 | 5.2 | 5.4 | 5.3 | 5.1 | 5 | 5.858 | 5.645 | 5.594 | 5.448 | 5.495 | 1.104 | 6.227 | 5.938 | 6.743 | 6.678 | 6.576 | 6.699 | 6.536 | 6.524 | 6.935 | 6.675 | 6.471 | 6.337 | 6.548 | 7.178 | 7.278 | 7.243 | 6.778 | 6.761 | 6.645 | 6.907 | 7.045 | 7.472
| 110 | 5.4 | 4.8 | 5 | 5.2 | 5.2 | 6 | 5.4 | 5.9 | 5.4 | 5.5 | 4.9 | 5.3 | 5.3 | 5 | 5.2 | 5.1 | 4.6 | 5.5 | 5.601 | 5.747 | 5.228 | 5.302 | 5.469 | 1.368 | 6.025 | 5.847 | 6.573 | 6.497 | 6.669 | 6.64 | 6.355 | 6.373 | 6.629 | 6.656 | 6.372 | 6.418 | 6.535 | 7.023 | 7.078 | 7.15 | 6.601 | 6.568 | 6.438 | 6.716 | 6.821 | 7.141
| 120 | 5.6 | 5.2 | 5.2 | 5.3 | 5.4 | 6 | 5.5 | 6.1 | 5.9 | 5.8 | 5.3 | 5.3 | 5.3 | 5.3 | 4.9 | 5.4 | 5.6 | 5.8 | 5.774 | 5.803 | 5.514 | 5.458 | 5.561 | 1.565 | 6.045 | 6.04 | 7.165 | 6.718 | 6.737 | 6.952 | 6.75 | 6.564 | 6.92 | 7.006 | 6.493 | 6.508 | 6.858 | 7.294 | 7.766 | 7.456 | 6.679 | 7.173 | 6.741 | 6.953 | 7.341 | 7.453
| 130 | 2.6 | 2.4 | 1.9 | 2.5 | 2.7 | 3.4 | 2.7 | 3.4 | 2.8 | 2.6 | 2.4 | 2.4 | 2.6 | 2.4 | 2 | 2.8 | 2.7 | 2.8 | 3.087 | 2.887 | 2.752 | 2.627 | 2.638 | 1.424 | 3.383 | 3.054 | 4.022 | 4.039 | 3.949 | 4.222 | 4.106 | 3.763 | 4.145 | 3.807 | 3.481 | 3.935 | 4.166 | 4.283 | 4.553 | 4.394 | 4.014 | 4.034 | 4.059 | 4.089 | 4.315 | 4.581
| 140 | 1.3 | 0.2 | 0.3 | 0.9 | 1.1 | 1.6 | 0.8 | 1.4 | 1.1 | 0.9 | 0.5 | 0.6 | 0.8 | 0.8 | 0.8 | 0.8 | 1 | 1.1 | 0.958 | 1.009 | 0.772 | 0.862 | 1.011 | 1.434 | 1.624 | 1.388 | 2.645 | 2.536 | 2.57 | 2.745 | 2.35 | 2.044 | 2.292 | 2.446 | 2.17 | 1.953 | 2.375 | 2.773 | 3.09 | 2.988 | 2.143 | 2.673 | 2.357 | 2.546 | 2.667 | 2.933
| 150 | 0.8 | 0.9 | 1.2 | 1.3 | 0.7 | 0.9 | 1 | 0.4 | 0.7 | 0.6 | 0.9 | 1 | 0.7 | 0.6 | 0.5 | 0.5 | 0.6 | 0.7 | 0.262 | 0.237 | 0.457 | 0.583 | 0.533 | 1.171 | 0.939 | 0.727 | 1.859 | 1.801 | 1.802 | 1.821 | 1.375 | 1.365 | 1.41 | 1.492 | 1.834 | 1.22 | 1.6 | 1.662 | 1.908 | 1.839 | 1.225 | 1.591 | 1.587 | 1.555 | 1.753 | 1.725
| 160 | 2.3 | 2.8 | 3.4 | 2.4 | 2.4 | 1.7 | 2.3 | 1.4 | 2 | 2 | 2.7 | 2.6 | 2.4 | 2.3 | 2.1 | 2.3 | 2.1 | 1.7 | 1.816 | 2.018 | 2.317 | 2.335 | 1.985 | 1.145 | 1.685 | 2.019 | 1.519 | 1.88 | 1.944 | 1.555 | 1.705 | 1.835 | 1.142 | 1.6 | 2.404 | 1.615 | 1.217 | 1.28 | 1.201 | 1.102 | 1.628 | 1.217 | 1.857 | 1.699 | 1.425 | 1.192
| 170 | 3 | 3.2 | 4.1 | 3.1 | 2.8 | 2.1 | 3 | 1.9 | 2.9 | 2.7 | 3.2 | 3.2 | 3.1 | 2.5 | 2.9 | 2.8 | 2.8 | 2.5 | 2.442 | 2.458 | 2.862 | 3.059 | 2.594 | 0.866 | 2.118 | 3.313 | 2.594 | 2.388 | 2.351 | 2.301 | 2.529 | 2.526 | 1.786 | 2.404 | 3.256 | 2.021 | 1.568 | 1.663 | 1.717 | 1.503 | 2.264 | 1.766 | 2.274 | 2.155 | 1.804 | 1.507
| 180 | 4.7 | 5.4 | 6.1 | 4.5 | 4.5 | 3.7 | 4.5 | 3.5 | 4.5 | 4.2 | 4.8 | 4.8 | 4.7 | 4.6 | 4.9 | 4.5 | 4.5 | 4.3 | 4.206 | 4.491 | 4.909 | 4.931 | 4.527 | 0.669 | 3.726 | 4.283 | 3.449 | 3.738 | 3.317 | 3.494 | 3.955 | 4.203 | 3.46 | 3.782 | 4.945 | 3.442 | 3.411 | 2.868 | 2.838 | 2.903 | 3.699 | 3.282 | 3.829 | 3.718 | 3.316 | 3.389
| 190 | 5.2 | 5.7 | 7.1 | 5.3 | 4.9 | 4.3 | 5.5 | 4.3 | 5.5 | 4.6 | 5.5 | 5.7 | 5.3 | 4.9 | 5.6 | 5.2 | 5 | 5.1 | 4.94 | 5.21 | 5.378 | 5.778 | 5.041 | 0.881 | 4.304 | 5.15 | 4.219 | 4.456 | 4.324 | 4.299 | 5.25 | 5.259 | 3.927 | 4.824 | 5.951 | 4.122 | 3.784 | 3.393 | 3.254 | 3.444 | 4.69 | 3.858 | 4.874 | 4.326 | 3.961 | 3.584
| 200 | 5.1 | 5.9 | 7.7 | 4.9 | 5 | 5 | 5.6 | 4.2 | 5.1 | 4.5 | 5.4 | 5.8 | 5 | 5 | 5.2 | 5 | 5.2 | 4.6 | 4.892 | 4.937 | 5.732 | 5.586 | 5.262 | 0.943 | 4.968 | 5.003 | 4.082 | 4.311 | 4.229 | 4.333 | 5.256 | 5.377 | 3.813 | 4.786 | 6.07 | 4.585 | 3.79 | 3.488 | 3.037 | 3.011 | 4.484 | 3.755 | 4.49 | 4.275 | 3.891 | 3.416
| 210 | 4.2 | 5.7 | 6.8 | 4.7 | 4.7 | 4.9 | 5.1 | 4.1 | 5.4 | 4.3 | 5.1 | 5.6 | 5 | 4.5 | 4.8 | 5 | 5.1 | 4.9 | 4.472 | 4.944 | 5.104 | 5.604 | 5.32 | 0.498 | 4.926 | 5.214 | 4.843 | 4.76 | 4.406 | 4.598 | 5.145 | 5.3 | 3.684 | 4.782 | 6.492 | 4.271 | 3.961 | 3.727 | 3.361 | 3.546 | 4.613 | 3.637 | 4.85 | 4.581 | 4.324 | 3.998
| 220 | 2.8 | 2.8 | 5.4 | 3.1 | 3.1 | 2.8 | 3.8 | 2.7 | 4.4 | 2.7 | 3.1 | 4.1 | 3.2 | 2.5 | 2.8 | 3 | 3.1 | 3.3 | 3.235 | 2.785 | 2.933 | 3.641 | 3.579 | 1.957 | 2.799 | 4.282 | 4.508 | 4.69 | 2.986 | 3.748 | 2.935 | 3.529 | 2.374 | 3.101 | 5.426 | 1.072 | 0.999 | 3.305 | 3.21 | 2.865 | 3.635 | 3.249 | 4.125 | 3.414 | 3.566 | 3.07
| 226 | 2.8 | 3.3 | 5.7 | 3.1 | 3.4 | 3.1 | 3.2 | 3.7 | 4.6 | 3.5 | 3.5 | 4.9 | 4.3 | 2.5 | 3.4 | 2.7 | 4.4 | 3.3 | 3.342 | 2.607 | 4.371 | 4.229 | 4.552 | 2.744 | 3.089 | 5.093 | 5.579 | 5.637 | 3.665 | 3.931 | 3.542 | 4.971 | 3.407 | 4.014 | 7.008 | 2.568 | 1.895 | 4.502 | 4.481 | 4.319 | 4.539 | 4.817 | 4.814 | 4.35 | 5.272 | 4.538
Correlation coefficient
| E [MeV] | 2019-01-10 | 2019-08-06 | 2019-11-12 | 2020-02-02 | 2020-03-25 | 2020-05-05 | 2020-06-02 | 2020-07-28 | 2020-09-01 | 2020-10-06 | 2020-10-11 | 2020-11-18 | 2020-11-18 | 2020-12-10 | 2021-01-05 | 2021-01-11 | 2021-02-04 | 2021-02-14 | 2021-03-09 | 2021-04-06 | 2021-06-22 | 2021-07-20 | 2021-08-17 | 2021-09-21 | 2021-10-12 | 2021-10-17 | 2021-11-08 | 2021-12-09 | 2022-01-21 | 2022-02-08 | 2022-03-08 | 2022-03-13 | 2022-04-04 | 2022-05-16 | 2022-07-08 | 2022-07-12 | 2022-08-30 | 2022-10-09 | 2022-11-08 | 2022-12-01 | 2023-02-04 | 2023-03-10 | 2023-03-20 | 2023-04-22 | 2023-05-27 | 2023-06-20
| 60 | 0.069 | 0.065 | 0.07 | 0.069 | 0.067 | 0.086 | 0.071 | 0.074 | 0.068 | 0.072 | 0.073 | 0.071 | 0.067 | 0.067 | 0.071 | 0.069 | 0.071 | 0.057 | 0.069 | 0.07 | 0.07 | 0.074 | 0.077 | -0.005 | 0.086 | 0.084 | 0.083 | 0.079 | 0.081 | 0.083 | 0.081 | 0.084 | 0.087 | 0.081 | 0.083 | 0.083 | -0.093 | -0.084 | -0.086 | -0.086 | -0.086 | -0.084 | -0.083 | -0.086 | -0.095 | -0.09
| 70 | 0.052 | 0.046 | 0.048 | 0.053 | 0.055 | 0.069 | 0.055 | 0.061 | 0.054 | 0.052 | 0.051 | 0.055 | 0.054 | 0.057 | 0.051 | 0.052 | 0.049 | 0.081 | 0.055 | 0.056 | 0.057 | 0.058 | 0.057 | 0.01 | 0.069 | 0.065 | 0.074 | 0.07 | 0.071 | 0.07 | 0.067 | 0.067 | 0.071 | 0.067 | 0.066 | 0.061 | -0.07 | -0.077 | -0.078 | -0.075 | -0.071 | -0.071 | -0.07 | -0.073 | -0.076 | -0.08
| 80 | 0.073 | 0.068 | 0.07 | 0.073 | 0.074 | 0.081 | 0.076 | 0.085 | 0.077 | 0.075 | 0.067 | 0.076 | 0.071 | 0.07 | 0.068 | 0.073 | 0.079 | 0.025 | 0.078 | 0.08 | 0.073 | 0.082 | 0.081 | -0.004 | 0.081 | 0.089 | 0.099 | 0.097 | 0.096 | 0.099 | 0.095 | 0.092 | 0.098 | 0.091 | 0.097 | 0.087 | -0.09 | -0.098 | -0.105 | -0.105 | -0.093 | -0.095 | -0.095 | -0.099 | -0.102 | -0.105
| 90 | 0.018 | 0.013 | 0.012 | 0.018 | 0.023 | 0.034 | 0.022 | 0.032 | 0.027 | 0.022 | 0.015 | 0.022 | 0.015 | 0.014 | 0.019 | 0.025 | 0.021 | 0.105 | 0.027 | 0.022 | 0.023 | 0.023 | 0.024 | 0.01 | 0.034 | 0.033 | 0.044 | 0.044 | 0.04 | 0.043 | 0.036 | 0.035 | 0.043 | 0.037 | 0.036 | 0.031 | -0.036 | -0.048 | -0.051 | -0.052 | -0.041 | -0.042 | -0.039 | -0.053 | -0.046 | -0.054
| 100 | 0.101 | 0.096 | 0.1 | 0.104 | 0.105 | 0.124 | 0.106 | 0.117 | 0.112 | 0.109 | 0.102 | 0.098 | 0.107 | 0.104 | 0.108 | 0.106 | 0.102 | 0.1 | 0.116 | 0.113 | 0.112 | 0.108 | 0.109 | 0.009 | 0.124 | 0.118 | 0.133 | 0.132 | 0.13 | 0.133 | 0.129 | 0.128 | 0.137 | 0.132 | 0.127 | 0.124 | -0.128 | -0.142 | -0.144 | -0.143 | -0.134 | -0.134 | -0.131 | -0.136 | -0.139 | -0.148
| 110 | 0.106 | 0.095 | 0.1 | 0.103 | 0.105 | 0.12 | 0.108 | 0.118 | 0.107 | 0.109 | 0.097 | 0.105 | 0.105 | 0.1 | 0.104 | 0.101 | 0.092 | 0.11 | 0.111 | 0.114 | 0.104 | 0.106 | 0.109 | 0.018 | 0.12 | 0.116 | 0.13 | 0.129 | 0.132 | 0.131 | 0.126 | 0.126 | 0.131 | 0.132 | 0.126 | 0.127 | -0.129 | -0.139 | -0.14 | -0.142 | -0.131 | -0.13 | -0.127 | -0.133 | -0.135 | -0.142
| 120 | 0.11 | 0.104 | 0.104 | 0.104 | 0.108 | 0.12 | 0.11 | 0.121 | 0.117 | 0.115 | 0.106 | 0.105 | 0.105 | 0.105 | 0.097 | 0.108 | 0.112 | 0.115 | 0.115 | 0.116 | 0.11 | 0.108 | 0.111 | 0.023 | 0.12 | 0.119 | 0.14 | 0.131 | 0.131 | 0.136 | 0.131 | 0.127 | 0.135 | 0.137 | 0.125 | 0.127 | -0.134 | -0.142 | -0.152 | -0.146 | -0.131 | -0.141 | -0.132 | -0.136 | -0.144 | -0.146
| 130 | 0.05 | 0.048 | 0.037 | 0.047 | 0.053 | 0.067 | 0.053 | 0.067 | 0.055 | 0.05 | 0.048 | 0.046 | 0.05 | 0.046 | 0.04 | 0.055 | 0.052 | 0.055 | 0.062 | 0.058 | 0.055 | 0.052 | 0.052 | 0.024 | 0.067 | 0.059 | 0.075 | 0.076 | 0.074 | 0.081 | 0.079 | 0.072 | 0.08 | 0.072 | 0.061 | 0.077 | -0.081 | -0.081 | -0.086 | -0.084 | -0.078 | -0.076 | -0.078 | -0.078 | -0.082 | -0.088
| 140 | 0.019 | 0.003 | -0.004 | 0.009 | 0.015 | 0.031 | 0.011 | 0.028 | 0.018 | 0.014 | 0.007 | 0.008 | 0.013 | 0.016 | 0.015 | 0.017 | 0.018 | 0.02 | 0.019 | 0.02 | 0.015 | 0.015 | 0.019 | 0.019 | 0.031 | 0.023 | 0.044 | 0.04 | 0.042 | 0.048 | 0.042 | 0.031 | 0.041 | 0.042 | 0.027 | 0.034 | -0.043 | -0.047 | -0.055 | -0.054 | -0.038 | -0.047 | -0.039 | -0.042 | -0.046 | -0.054
| 150 | -0.008 | -0.018 | -0.024 | -0.013 | -0.008 | 0.009 | -0.009 | 0.004 | -0.007 | -0.008 | -0.016 | -0.017 | -0.01 | -0.011 | -0.008 | -0.009 | -0.007 | -0.007 | -0.001 | -0.003 | -0.007 | -0.008 | -0.002 | 0.013 | 0.009 | -0.004 | 0.014 | 0.016 | 0.017 | 0.022 | 0.011 | 0.009 | 0.019 | 0.01 | -0.001 | 0.016 | -0.024 | -0.022 | -0.024 | -0.028 | -0.013 | -0.018 | -0.014 | -0.015 | -0.02 | -0.022
| 160 | -0.043 | -0.057 | -0.068 | -0.044 | -0.043 | -0.03 | -0.044 | -0.028 | -0.039 | -0.039 | -0.054 | -0.05 | -0.046 | -0.045 | -0.041 | -0.046 | -0.041 | -0.034 | -0.036 | -0.04 | -0.046 | -0.046 | -0.038 | 0.012 | -0.03 | -0.038 | -0.012 | -0.018 | -0.017 | -0.016 | -0.025 | -0.026 | -0.016 | -0.022 | -0.032 | -0.024 | 0.012 | 0.01 | 0.005 | 0.006 | 0.027 | 0.015 | 0.024 | 0.022 | 0.014 | 0.01
| 170 | -0.058 | -0.064 | -0.081 | -0.055 | -0.053 | -0.037 | -0.053 | -0.038 | -0.056 | -0.051 | -0.062 | -0.06 | -0.058 | -0.049 | -0.057 | -0.056 | -0.054 | -0.045 | -0.047 | -0.049 | -0.056 | -0.059 | -0.05 | 0.014 | -0.037 | -0.061 | -0.031 | -0.026 | -0.03 | -0.018 | -0.04 | -0.038 | -0.029 | -0.036 | -0.048 | -0.032 | 0.024 | 0.02 | 0.018 | 0.017 | 0.039 | 0.027 | 0.035 | 0.032 | 0.025 | 0.019
| 180 | -0.091 | -0.107 | -0.122 | -0.088 | -0.088 | -0.073 | -0.088 | -0.07 | -0.09 | -0.084 | -0.096 | -0.096 | -0.094 | -0.092 | -0.098 | -0.09 | -0.089 | -0.085 | -0.084 | -0.09 | -0.098 | -0.098 | -0.09 | 0.013 | -0.073 | -0.083 | -0.059 | -0.064 | -0.058 | -0.062 | -0.074 | -0.079 | -0.067 | -0.07 | -0.091 | -0.067 | 0.064 | 0.056 | 0.055 | 0.056 | 0.073 | 0.065 | 0.074 | 0.072 | 0.064 | 0.065
| 190 | -0.102 | -0.114 | -0.141 | -0.104 | -0.097 | -0.084 | -0.106 | -0.087 | -0.109 | -0.092 | -0.108 | -0.112 | -0.105 | -0.099 | -0.112 | -0.104 | -0.1 | -0.101 | -0.099 | -0.104 | -0.107 | -0.115 | -0.1 | 0.018 | -0.084 | -0.101 | -0.075 | -0.08 | -0.076 | -0.077 | -0.1 | -0.098 | -0.075 | -0.092 | -0.111 | -0.078 | 0.073 | 0.067 | 0.063 | 0.067 | 0.093 | 0.076 | 0.095 | 0.084 | 0.077 | 0.07
| 200 | -0.096 | -0.117 | -0.152 | -0.097 | -0.099 | -0.098 | -0.109 | -0.083 | -0.101 | -0.09 | -0.107 | -0.115 | -0.099 | -0.099 | -0.104 | -0.099 | -0.103 | -0.092 | -0.098 | -0.098 | -0.114 | -0.111 | -0.105 | 0.019 | -0.098 | -0.097 | -0.073 | -0.076 | -0.078 | -0.079 | -0.1 | -0.102 | -0.074 | -0.091 | -0.112 | -0.091 | 0.075 | 0.067 | 0.059 | 0.058 | 0.089 | 0.074 | 0.086 | 0.082 | 0.075 | 0.067
| 210 | -0.076 | -0.113 | -0.135 | -0.084 | -0.087 | -0.094 | -0.096 | -0.079 | -0.104 | -0.083 | -0.096 | -0.108 | -0.096 | -0.089 | -0.093 | -0.097 | -0.099 | -0.092 | -0.086 | -0.097 | -0.1 | -0.109 | -0.103 | 0.01 | -0.094 | -0.096 | -0.073 | -0.068 | -0.065 | -0.062 | -0.093 | -0.095 | -0.067 | -0.082 | -0.109 | -0.085 | 0.079 | 0.059 | 0.049 | 0.059 | 0.085 | 0.063 | 0.084 | 0.081 | 0.075 | 0.067
| 220 | -0.003 | -0.056 | -0.105 | -0.02 | -0.034 | -0.041 | -0.052 | -0.028 | -0.074 | -0.033 | -0.042 | -0.064 | -0.046 | -0.035 | -0.035 | -0.043 | -0.04 | -0.048 | -0.052 | -0.044 | -0.046 | -0.062 | -0.058 | 0.035 | -0.041 | -0.058 | -0.024 | -0.03 | -0.008 | 0.004 | -0.027 | -0.036 | -0.005 | -0.029 | -0.061 | -0.004 | 0.01 | 0.007 | 0.001 | 0 | 0.044 | 0.014 | 0.041 | 0.029 | 0.026 | 0.017
| 226 | -0.001 | -0.066 | -0.113 | -0.003 | -0.04 | -0.029 | -0.025 | -0.041 | -0.067 | -0.045 | -0.046 | -0.076 | -0.064 | -0.021 | -0.046 | -0.023 | -0.064 | -0.031 | -0.046 | -0.023 | -0.073 | -0.066 | -0.073 | 0.046 | -0.029 | -0.074 | -0.047 | -0.045 | -0.008 | -0.006 | -0.043 | -0.068 | -0.038 | -0.043 | -0.099 | -0.032 | 0.016 | 0.017 | 0.013 | 0.026 | 0.044 | 0.039 | 0.035 | 0.051 | 0.052 | 0.04
S parameter [%] sigma comparision
| E [MeV] | 2019-01-10 | 2019-08-06 | 2019-11-12 | 2020-02-02 | 2020-03-25 | 2020-05-05 | 2020-06-02 | 2020-07-28 | 2020-09-01 | 2020-10-06 | 2020-10-11 | 2020-11-18 | 2020-11-18 | 2020-12-10 | 2021-01-05 | 2021-01-11 | 2021-02-04 | 2021-02-14 | 2021-03-09 | 2021-04-06 | 2021-06-22 | 2021-07-20 | 2021-08-17 | 2021-09-21 | 2021-10-12 | 2021-10-17 | 2021-11-08 | 2021-12-09 | 2022-01-21 | 2022-02-08 | 2022-03-08 | 2022-03-13 | 2022-04-04 | 2022-05-16 | 2022-07-08 | 2022-07-12 | 2022-08-30 | 2022-10-09 | 2022-11-08 | 2022-12-01 | 2023-02-04 | 2023-03-10 | 2023-03-20 | 2023-04-22 | 2023-05-27 | 2023-06-20
| 60 | 1.5 | 7 | 2.1 | 1.3 | 1.4 | 1.3 | 1.5 | 1.3 | 1.2 | 1.7 | 1.9 | 1.5 | 1 | 1.6 | 1.5 | 1.5 | 1.4 | 3.7 | 1.375 | 1.279 | 1.358 | 1.594 | 1.847 | 0.31 | 1.295 | 1.538 | 0.92 | 0.738 | 0.813 | 0.983 | 1.185 | 1.339 | 1.432 | 0.931 | 1.197 | 1.152 | 1.593 | 1.131 | 1.112 | 1.463 | 1.586 | 1.449 | 1.404 | 1.539 | 1.919 | 1.599
| 70 | 4.2 | 1.9 | 4.1 | 3.6 | 3.7 | 3.6 | 3.7 | 3.7 | 3.6 | 3.8 | 3.9 | 3.8 | 3.7 | 4.1 | 3.7 | 3.7 | 3.4 | -0.7 | 3.68 | 3.61 | 3.81 | 3.856 | 3.96 | 3.386 | 3.552 | 3.629 | 3.505 | 3.243 | 3.353 | 3.348 | 3.459 | 3.528 | 3.435 | 3.41 | 3.076 | 3.101 | 3.592 | 3.574 | 3.487 | 3.532 | 3.794 | 3.563 | 3.619 | 3.574 | 3.782 | 3.75
| 80 | -0.2 | 2.2 | -0.7 | -0.9 | -1.1 | -1.7 | -1.2 | -0.7 | -1 | -1 | -0.7 | -0.8 | -1 | -0.9 | -1 | -1.2 | -0.8 | -0.2 | -0.842 | -0.993 | -0.944 | -0.89 | -1.241 | -1.375 | -1.731 | -1.354 | -1.536 | -1.266 | -1.401 | -1.645 | -1.439 | -1.737 | -1.366 | -1.476 | -1.456 | -1.799 | -1.731 | -1.879 | -1.755 | -1.538 | -1.544 | -1.819 | -1.565 | -1.468 | -1.234 | -1.305
| 90 | 0.5 | 0.1 | -0.5 | -0.5 | -0.6 | -1 | -0.8 | -0.3 | -0.3 | -0.3 | -0.7 | -0.4 | -0.8 | -0.7 | -0.5 | -0.1 | -0.2 | 0.1 | -0.166 | -0.629 | -0.378 | -0.688 | -0.488 | -0.649 | -0.988 | -0.75 | -0.866 | -1.05 | -1.106 | -0.976 | -1.39 | -1.174 | -1.041 | -1.28 | -1.5 | -1.332 | -1.351 | -1.171 | -1.219 | -1.051 | -1.069 | -1.037 | -1.059 | -0.714 | -1.094 | -0.908
| 100 | 0.3 | 8.7 | 0.1 | -0.1 | -0.3 | -0.3 | -0.2 | 0.2 | 0.1 | 0 | 0.1 | 0.2 | 0.2 | 0.3 | 0.2 | 0.2 | 0 | 0.2 | 0.489 | 0.105 | 0.095 | -0.312 | -0.21 | -1 | -0.321 | -0.515 | -0.794 | -0.917 | -0.863 | -0.746 | -0.914 | -1.068 | -0.862 | -0.96 | -1.003 | -1.16 | -1.131 | -0.913 | -0.822 | -0.903 | -0.882 | -0.721 | -1.042 | -0.967 | -0.716 | -0.638
| 110 | 0.6 | 8.9 | 0.2 | 0.2 | -0.2 | -0.2 | 0.1 | 0.2 | -0.1 | 0.3 | 0 | 0.4 | 0.1 | 0 | 0.3 | 0.1 | 0.1 | 0.2 | 0.326 | 0.314 | -0.011 | -0.161 | 0.091 | -1.04 | -0.192 | -0.373 | -0.849 | -0.701 | -0.681 | -0.803 | -0.828 | -0.799 | -0.727 | -0.666 | -0.952 | -0.837 | -0.771 | -0.694 | -0.913 | -0.555 | -0.637 | -0.693 | -0.787 | -0.752 | -0.759 | -0.634
| 120 | 0.4 | 8.7 | -0.1 | -0.6 | -0.5 | -0.6 | -0.5 | -0.3 | -0.4 | -0.3 | -0.2 | -0.2 | -0.6 | -0.2 | -0.4 | -0.2 | -0.1 | -0.4 | -0.268 | -0.148 | -0.291 | -0.475 | -0.328 | -1.057 | -0.644 | -0.803 | -1.237 | -1.468 | -1.479 | -1.37 | -1.414 | -1.454 | -1.335 | -1.241 | -1.662 | -1.23 | -1.21 | -1.46 | -1.448 | -1.275 | -1.145 | -1.169 | -1.346 | -1.183 | -1.226 | -1.18
| 130 | 0.5 | 2 | -0.4 | -0.8 | -0.4 | -0.6 | -0.5 | 0 | -0.5 | -0.5 | -0.1 | -0.4 | -0.6 | -0.7 | -0.1 | -0.1 | -0.4 | -0.5 | -0.076 | -0.195 | -0.22 | -0.454 | -0.328 | -0.777 | -0.57 | -0.718 | -1.438 | -1.295 | -1.389 | -1.153 | -1.099 | -1.116 | -0.97 | -1.256 | -1.652 | -0.711 | -1.005 | -1.364 | -1.504 | -1.245 | -0.986 | -1.278 | -1.14 | -1.254 | -1.288 | -1.236
| 140 | 0.9 | -1.3 | -0.2 | -0.8 | -0.8 | -0.5 | -0.5 | -0.3 | -0.6 | -0.5 | -0.4 | -0.4 | -0.4 | -0.3 | -0.4 | 0 | -0.4 | -0.5 | -0.181 | -0.129 | -0.155 | -0.4 | -0.376 | -1.078 | -0.527 | -0.796 | -1.461 | -1.581 | -1.449 | -1.339 | -1.083 | -1.356 | -1.057 | -1.285 | -1.684 | -0.968 | -1.001 | -1.428 | -1.399 | -1.224 | -1.008 | -1.285 | -1.319 | -1.429 | -1.319 | -1.185
| 150 | 0.7 | -2.8 | -0.1 | -1.1 | -0.6 | -0.8 | -0.9 | -0.4 | -0.7 | -0.5 | -0.3 | -0.6 | -0.6 | -0.2 | -0.2 | -0.3 | -0.5 | -0.6 | -0.259 | -0.169 | -0.272 | -0.401 | -0.521 | -0.957 | -0.817 | -0.699 | -1.724 | -1.616 | -1.586 | -1.456 | -1.268 | -1.286 | -1.038 | -1.4 | -1.834 | -0.93 | -1.065 | -1.224 | -1.464 | -1.175 | -1.047 | -1.296 | -1.429 | -1.363 | -1.434 | -1.335
| 160 | 0.8 | -5.8 | -0.2 | -1 | -1 | -0.8 | -0.6 | -0.2 | -0.5 | -0.6 | -0.4 | -0.5 | -0.6 | -0.2 | -0.3 | -0.3 | -0.4 | -0.4 | -0.201 | -0.249 | -0.343 | -0.493 | -0.582 | -0.968 | -0.807 | -0.67 | -1.404 | -1.651 | -1.744 | -1.336 | -1.161 | -1.285 | -0.826 | -1.166 | -1.793 | -1.097 | -1.063 | -1.177 | -1.171 | -1.063 | -0.932 | -0.952 | -1.413 | -1.311 | -1.245 | -1.091
| 170 | 0.5 | -5.9 | -0.6 | -1.5 | -1 | -1 | -1.3 | -0.4 | -0.8 | -0.8 | -0.9 | -1 | -0.9 | -0.2 | -0.6 | -0.4 | -0.8 | -1.1 | -0.691 | -0.386 | -0.579 | -0.81 | -0.705 | -0.525 | -1.05 | -1.244 | -2.086 | -1.993 | -1.793 | -2.115 | -1.54 | -1.647 | -1.074 | -1.566 | -2.225 | -1.21 | -1.018 | -1.32 | -1.463 | -1.246 | -1.167 | -1.161 | -1.481 | -1.431 | -1.318 | -1.169
| 180 | 1.2 | -9.5 | -0.1 | -1 | -0.8 | -0.6 | -0.6 | 0.1 | -0.2 | -0.3 | -0.4 | -0.2 | -0.1 | 0.2 | -0.2 | 0.2 | -0.3 | -0.5 | 0.112 | -0.083 | -0.217 | -0.491 | -0.21 | -0.233 | -0.618 | -0.989 | -1.741 | -1.968 | -1.576 | -1.605 | -1.322 | -1.355 | -0.748 | -1.365 | -1.946 | -0.743 | -1.108 | -0.576 | -0.673 | -0.769 | -0.635 | -0.421 | -0.931 | -0.894 | -0.855 | -0.946
| 190 | 0.9 | -9.9 | 0.1 | -1.1 | -1 | -0.8 | -1.1 | 0 | -0.2 | -0.3 | -0.6 | -0.6 | -0.2 | 0 | -0.2 | -0.2 | -0.3 | -0.6 | -0.166 | -0.311 | -0.203 | -0.425 | -0.465 | 0.076 | -0.822 | -0.858 | -1.97 | -1.996 | -2.096 | -1.871 | -1.584 | -1.847 | -1.076 | -1.426 | -2.074 | -1.28 | -0.988 | -0.653 | -0.821 | -0.721 | -0.681 | -0.615 | -1.136 | -1.098 | -0.931 | -0.64
| 200 | 1.8 | -10 | 0.5 | -0.8 | -0.5 | -0.7 | -1.2 | -0.2 | -0.3 | -0.2 | -0.3 | -0.2 | 0 | 0.2 | 0 | 0 | -0.2 | -0.5 | -0.027 | 0.043 | -0.191 | -0.406 | -0.176 | 0.175 | -0.725 | -1.167 | -1.842 | -1.999 | -1.618 | -1.744 | -1.579 | -1.657 | -0.789 | -1.482 | -2.293 | -0.214 | -0.467 | -0.981 | -0.756 | -0.846 | -0.618 | -0.741 | -1.259 | -1.102 | -0.974 | -0.686
| 210 | 1.9 | -9.5 | 0.3 | -2.1 | -1.8 | -1.4 | -1.5 | -1.2 | -1.3 | -1.2 | -1.5 | -1.5 | -1.4 | -0.9 | -1 | -1 | -1.4 | -1.5 | -1.23 | -0.968 | -0.999 | -1.201 | -1.198 | 0.127 | -1.363 | -1.953 | -3.194 | -3.352 | -2.979 | -3.376 | -2.139 | -2.37 | -1.528 | -2.448 | -3.52 | -0.104 | -0.252 | -2.294 | -2.324 | -1.975 | -1.755 | -1.813 | -2.427 | -2.166 | -2.117 | -2.178
| 220 | 2.8 | -7.1 | 1 | -2.9 | -2.5 | -1.9 | -2.8 | -2.2 | -2.4 | -2.1 | -2.3 | -2.5 | -2.3 | -1.8 | -2.1 | -2.1 | -2.3 | -2.3 | -1.956 | -1.702 | -1.812 | -1.936 | -2.09 | 0.857 | -1.917 | -3.147 | -4.347 | -4.436 | -2.959 | -3.743 | -2.602 | -3.05 | -2.362 | -2.731 | -4.501 | 1.054 | 0.857 | -3.287 | -3.209 | -2.865 | -2.886 | -3.178 | -3.567 | -3.098 | -3.327 | -2.951
| 226 | 2.8 | -7.3 | 0.6 | -3.1 | -2.8 | -2.7 | -3 | -3.1 | -3.1 | -2.7 | -2.6 | -3.1 | -2.9 | -2.3 | -2.5 | -2.4 | -3 | -2.9 | -2.443 | -2.351 | -2.402 | -2.643 | -2.706 | 1.47 | -2.737 | -3.508 | -5.06 | -5.173 | -3.642 | -3.918 | -2.804 | -3.618 | -2.832 | -3.387 | -4.924 | 2.023 | 1.718 | -4.418 | -4.435 | -4.121 | -3.976 | -4.401 | -4.486 | -3.524 | -4.575 | -4.084
el

## GTR2 G315 z0 x
GTR2 Spot size in x axis at iso for G315º every 10 MeV
Lynx PT
warrning | 8%
action level | 10%
[32 empty] | validation of new set point for deflector |  |  |  | After ECPM |  |  |  | After Cyclo PM
|  | E [MeV] | REF | [7 empty] | After Cyclo PM | After waterleak to cyclo | [12 empty] | old Lynx | old Lynx | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021
|  |  |  | 2019-10-22 | 2019-12-07 | 2019-12-17 | 2020-01-14 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-08-02 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27
|  |  |  | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm]
|  | 60 | 7.085 | 6.73 | 6.75 | 6.79 | 6.75 | 6.77 | 6.78 | 6.76 | 6.77 | 6.81 | 6.81 | 6.8 | 6.82 | 6.8 | 6.79 | 6.75 | 6.75 | 6.78 | 6.77 | 6.74 | 6.76 | 6.76 | 6.8 | 6.88 | 6.83 | 6.8 | 6.83 | 6.84 | 6.87 | 6.84 | 6.84 | 6.82 | 6.85 | 6.85 | 6.83 | 6.8 | 6.8 | 6.8 | 6.8 | 6.8
|  | 70 | 6.23 | 6.1 | 6.09 | 6.16 | 6.08 | 6.11 | 6.12 | 6.08 | 6.12 | 6.13 | 6.14 | 6.13 | 6.14 | 6.14 | 6.11 | 6.09 | 6.09 | 6.13 | 6.12 | 6.08 | 6.11 | 6.11 | 6.13 | 6.2 | 6.17 | 6.14 | 6.15 | 6.17 | 6.16 | 6.17 | 6.18 | 6.15 | 6.18 | 6.16 | 6.15 | 6.13 | 6.14 | 6.12 | 6.12 | 6.14
|  | 80 | 5.93 | 5.46 | 5.48 | 5.52 | 5.49 | 5.48 | 5.51 | 5.48 | 5.5 | 5.52 | 5.52 | 5.49 | 5.52 | 5.52 | 5.51 | 5.47 | 5.47 | 5.5 | 5.49 | 5.47 | 5.47 | 5.47 | 5.5 | 5.59 | 5.53 | 5.52 | 5.55 | 5.56 | 5.55 | 5.57 | 5.57 | 5.58 | 5.57 | 5.54 | 5.53 | 5.52 | 5.53 | 5.53 | 5.54 | 5.53
|  | 90 | 5.519 | 5.07 | 5.07 | 5.1 | 5.08 | 5.07 | 5.08 | 5.09 | 5.08 | 5.1 | 5.09 | 5.08 | 5.11 | 5.11 | 5.08 | 5.07 | 5.07 | 5.1 | 5.09 | 5.07 | 5.07 | 5.05 | 5.09 | 5.17 | 5.13 | 5.12 | 5.15 | 5.15 | 5.15 | 5.16 | 5.18 | 5.17 | 5.17 | 5.15 | 5.14 | 5.1 | 5.14 | 5.14 | 5.15 | 5.14
|  | 100 | 5.11 | 4.81 | 4.82 | 4.87 | 4.84 | 4.82 | 4.84 | 4.84 | 4.84 | 4.84 | 4.84 | 4.82 | 4.85 | 4.87 | 4.86 | 4.83 | 4.84 | 4.84 | 4.84 | 4.83 | 4.83 | 4.8 | 4.85 | 4.93 | 4.89 | 4.89 | 4.91 | 4.92 | 4.93 | 4.94 | 4.95 | 4.94 | 4.95 | 4.89 | 4.89 | 4.89 | 4.93 | 4.91 | 4.91 | 4.89
|  | 110 | 4.754 | 4.45 | 4.5 | 4.51 | 4.49 | 4.47 | 4.48 | 4.48 | 4.48 | 4.48 | 4.48 | 4.46 | 4.47 | 4.5 | 4.49 | 4.47 | 4.48 | 4.49 | 4.49 | 4.47 | 4.48 | 4.45 | 4.49 | 4.57 | 4.54 | 4.55 | 4.57 | 4.59 | 4.57 | 4.57 | 4.59 | 4.58 | 4.58 | 4.56 | 4.55 | 4.55 | 4.56 | 4.56 | 4.56 | 4.55
|  | 120 | 4.457 | 4.18 | 4.19 | 4.21 | 4.21 | 4.18 | 4.19 | 4.19 | 4.2 | 4.19 | 4.16 | 4.19 | 4.21 | 4.2 | 4.21 | 4.21 | 4.19 | 4.21 | 4.22 | 4.2 | 4.18 | 4.18 | 4.19 | 4.28 | 4.25 | 4.26 | 4.29 | 4.31 | 4.32 | 4.28 | 4.33 | 4.34 | 4.32 | 4.24 | 4.25 | 4.24 | 4.31 | 4.29 | 4.29 | 4.27
|  | 130 | 4.199 | 3.93 | 3.94 | 3.96 | 3.93 | 3.92 | 3.92 | 3.95 | 3.93 | 3.93 | 3.91 | 3.92 | 3.93 | 3.95 | 3.94 | 3.91 | 3.93 | 3.95 | 3.93 | 3.92 | 3.92 | 3.92 | 3.94 | 4.02 | 3.99 | 4 | 4.03 | 4.04 | 4.03 | 4.02 | 4.06 | 4.06 | 4.05 | 3.98 | 4 | 3.98 | 4.02 | 4.01 | 4.01 | 4.01
|  | 140 | 4.011 | 3.7 | 3.72 | 3.74 | 3.7 | 3.69 | 3.7 | 3.71 | 3.71 | 3.71 | 3.68 | 3.69 | 3.71 | 3.73 | 3.72 | 3.71 | 3.73 | 3.73 | 3.72 | 3.73 | 3.68 | 3.7 | 3.71 | 3.77 | 3.76 | 3.78 | 3.83 | 3.83 | 3.82 | 3.79 | 3.86 | 3.86 | 3.85 | 3.75 | 3.76 | 3.75 | 3.81 | 3.78 | 3.8 | 3.78
|  | 150 | 3.833 | 3.54 | 3.54 | 3.58 | 3.55 | 3.52 | 3.52 | 3.53 | 3.56 | 3.52 | 3.51 | 3.54 | 3.55 | 3.56 | 3.54 | 3.54 | 3.53 | 3.57 | 3.56 | 3.56 | 3.54 | 3.53 | 3.55 | 3.62 | 3.61 | 3.61 | 3.65 | 3.67 | 3.66 | 3.62 | 3.73 | 3.68 | 3.67 | 3.58 | 3.59 | 3.6 | 3.64 | 3.61 | 3.64 | 3.61
|  | 160 | 3.614 | 3.37 | 3.32 | 3.4 | 3.36 | 3.33 | 3.34 | 3.39 | 3.37 | 3.37 | 3.32 | 3.35 | 3.37 | 3.37 | 3.37 | 3.37 | 3.36 | 3.37 | 3.39 | 3.39 | 3.34 | 3.36 | 3.39 | 3.44 | 3.42 | 3.44 | 3.49 | 3.48 | 3.47 | 3.45 | 3.55 | 3.52 | 3.54 | 3.38 | 3.4 | 3.41 | 3.45 | 3.42 | 3.45 | 3.43
|  | 170 | 3.45 | 3.22 | 3.24 | 3.26 | 3.22 | 3.21 | 3.21 | 3.21 | 3.22 | 3.21 | 3.18 | 3.22 | 3.24 | 3.24 | 3.22 | 3.23 | 3.23 | 3.25 | 3.26 | 3.27 | 3.19 | 3.23 | 3.21 | 3.29 | 3.3 | 3.31 | 3.37 | 3.35 | 3.34 | 3.31 | 3.42 | 3.39 | 3.4 | 3.23 | 3.26 | 3.25 | 3.31 | 3.28 | 3.31 | 3.28
|  | 180 | 3.303 | 3.06 | 3.09 | 3.1 | 3.06 | 3.03 | 3.07 | 3.08 | 3.07 | 3.05 | 3.01 | 3.06 | 3.1 | 3.09 | 3.08 | 3.08 | 3.08 | 3.12 | 3.11 | 3.12 | 3.06 | 3.06 | 3.08 | 3.14 | 3.16 | 3.15 | 3.23 | 3.23 | 3.22 | 3.17 | 3.27 | 3.26 | 3.27 | 3.07 | 3.1 | 3.08 | 3.17 | 3.14 | 3.17 | 3.14
|  | 190 | 3.184 | 2.94 | 2.97 | 3.01 | 2.95 | 3.05 | 2.95 | 2.97 | 2.95 | 2.93 | 2.88 | 2.95 | 2.98 | 2.96 | 2.99 | 2.99 | 2.98 | 3.02 | 3.02 | 3.01 | 2.96 | 2.98 | 2.97 | 3.01 | 3.05 | 3.05 | 3.13 | 3.13 | 3.14 | 3.07 | 3.22 | 3.2 | 3.19 | 2.95 | 2.98 | 2.97 | 3.06 | 3.03 | 3.08 | 3.03
|  | 200 | 3.087 | 2.89 | 2.94 | 2.94 | 2.89 | 2.84 | 2.89 | 2.93 | 2.9 | 2.86 | 2.81 | 2.89 | 2.93 | 2.92 | 2.92 | 2.94 | 2.93 | 2.95 | 2.97 | 2.99 | 2.88 | 2.91 | 2.88 | 2.96 | 2.99 | 3 | 3.1 | 3.12 | 3.1 | 3.02 | 3.24 | 3.23 | 3.18 | 2.87 | 2.92 | 2.89 | 3.03 | 3 | 3.04 | 2.97
|  | 210 | 2.965 | 2.85 | 2.87 | 2.89 | 2.82 | 2.78 | 2.87 | 2.9 | 2.85 | 2.82 | 2.77 | 2.85 | 2.89 | 2.87 | 2.91 | 2.92 | 2.91 | 2.94 | 2.95 | 2.97 | 2.85 | 2.9 | 2.85 | 2.92 | 2.99 | 2.98 | 3.11 | 3.13 | 3.1 | 2.99 | 3.3 | 3.33 | 3.21 | 2.84 | 2.9 | 2.9 | 3.03 | 2.99 | 3.04 | 2.94
|  | 220 | 2.888 | 2.84 | 2.86 | 2.89 | 2.82 | 2.8 | 2.89 | 2.95 | 2.9 | 2.84 | 2.8 | 2.87 | 2.93 | 2.88 | 2.95 | 2.97 | 2.97 | 2.99 | 3.01 | 3.08 | 2.88 | 2.98 | 2.88 | 2.98 | 3.06 | 3.04 | 3.24 | 3.27 | 3.21 | 3.04 | 3.42 | 3.38 | 3.35 | 2.91 | 3.04 | 2.99 | 3.14 | 3.09 | 3.15 | 3.03
|  | 226 | 2.9 | 2.81 | 2.82 | 2.84 | 2.76 | 2.74 | 2.85 | 2.88 | 2.87 | 2.79 | 2.78 | 2.83 | 2.9 | 2.85 | 2.9 | 2.94 | 2.92 | 2.96 | 2.96 | 3.03 | 2.84 | 2.94 | 2.86 | 2.93 | 3.01 | 3 | 3.2 | 3.2 | 3.15 | 2.98 | 3.34 | 3.33 | 3.27 | 2.9 | 2.99 | 2.94 | 3.07 | 3.04 | 3.1 | 2.97
Deviation from reference [%]
|  | E [MeV] | note | 2020-06-16: After Cyclo PM; 2020-07-05: After waterleak to cyclo; 2021-10-19: old Lynx; 2021-11-15: old Lynx; 2021-12-10: Lynx SN 63021
|  |  |  | 2019-10-22 | 2019-12-07 | 2019-12-17 | 2020-01-14 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-07-02 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm] | sigma x [mm]
|  | 60 |  | -5.011% | -4.728% | -4.164% | -4.728% | -4.446% | -4.305% | -4.587% | -4.446% | -3.881% | -3.881% | -4.023% | -3.74% | -4.023% | -4.164% | -4.728% | -4.728% | -4.305% | -4.446% | -4.869% | -4.587% | -4.587% | -4.023% | -2.893% | -3.599% | -4.023% | -3.599% | -3.458% | -3.035% | -3.458% | -3.458% | -3.74% | -3.317% | -3.317% | -3.599% | -4.023% | -4.023% | -4.023% | -4.023% | -4.023%
|  | 70 |  | -2.087% | -2.247% | -1.124% | -2.408% | -1.926% | -1.766% | -2.408% | -1.766% | -1.605% | -1.445% | -1.605% | -1.445% | -1.445% | -1.926% | -2.247% | -2.247% | -1.605% | -1.766% | -2.408% | -1.926% | -1.926% | -1.605% | -0.4815% | -0.9631% | -1.445% | -1.284% | -0.9631% | -1.124% | -0.9631% | -0.8026% | -1.284% | -0.8026% | -1.124% | -1.284% | -1.605% | -1.445% | -1.766% | -1.766% | -1.445%
|  | 80 |  | -7.926% | -7.589% | -6.914% | -7.42% | -7.589% | -7.083% | -7.589% | -7.251% | -6.914% | -6.914% | -7.42% | -6.914% | -6.914% | -7.083% | -7.757% | -7.757% | -7.251% | -7.42% | -7.757% | -7.757% | -7.757% | -7.251% | -5.734% | -6.745% | -6.914% | -6.408% | -6.239% | -6.408% | -6.071% | -6.071% | -5.902% | -6.071% | -6.577% | -6.745% | -6.914% | -6.745% | -6.745% | -6.577% | -6.745%
|  | 90 |  | -8.136% | -8.136% | -7.592% | -7.954% | -8.136% | -7.954% | -7.773% | -7.954% | -7.592% | -7.773% | -7.954% | -7.411% | -7.411% | -7.954% | -8.136% | -8.136% | -7.592% | -7.773% | -8.136% | -8.136% | -8.498% | -7.773% | -6.324% | -7.048% | -7.23% | -6.686% | -6.686% | -6.686% | -6.505% | -6.142% | -6.324% | -6.324% | -6.686% | -6.867% | -7.592% | -6.867% | -6.867% | -6.686% | -6.867%
|  | 100 |  | -5.871% | -5.675% | -4.697% | -5.284% | -5.675% | -5.284% | -5.284% | -5.284% | -5.284% | -5.284% | -5.675% | -5.088% | -4.697% | -4.892% | -5.479% | -5.284% | -5.284% | -5.284% | -5.479% | -5.479% | -6.067% | -5.088% | -3.523% | -4.305% | -4.305% | -3.914% | -3.718% | -3.523% | -3.327% | -3.131% | -3.327% | -3.131% | -4.305% | -4.305% | -4.305% | -3.523% | -3.914% | -3.914% | -4.305%
|  | 110 |  | -6.395% | -5.343% | -5.133% | -5.553% | -5.974% | -5.764% | -5.764% | -5.764% | -5.764% | -5.764% | -6.184% | -5.974% | -5.343% | -5.553% | -5.974% | -5.764% | -5.553% | -5.553% | -5.974% | -5.764% | -6.395% | -5.553% | -3.87% | -4.501% | -4.291% | -3.87% | -3.45% | -3.87% | -3.87% | -3.45% | -3.66% | -3.66% | -4.081% | -4.291% | -4.291% | -4.081% | -4.081% | -4.081% | -4.291%
|  | 120 |  | -6.215% | -5.991% | -5.542% | -5.542% | -6.215% | -5.991% | -5.991% | -5.766% | -5.991% | -6.664% | -5.991% | -5.542% | -5.766% | -5.542% | -5.542% | -5.991% | -5.542% | -5.317% | -5.766% | -6.215% | -6.215% | -5.991% | -3.971% | -4.644% | -4.42% | -3.747% | -3.298% | -3.074% | -3.971% | -2.849% | -2.625% | -3.074% | -4.869% | -4.644% | -4.869% | -3.298% | -3.747% | -3.747% | -4.196%
|  | 130 |  | -6.406% | -6.168% | -5.692% | -6.406% | -6.644% | -6.644% | -5.93% | -6.406% | -6.406% | -6.883% | -6.644% | -6.406% | -5.93% | -6.168% | -6.883% | -6.406% | -5.93% | -6.406% | -6.644% | -6.644% | -6.644% | -6.168% | -4.263% | -4.977% | -4.739% | -4.025% | -3.787% | -4.025% | -4.263% | -3.31% | -3.31% | -3.548% | -5.216% | -4.739% | -5.216% | -4.263% | -4.501% | -4.501% | -4.501%
|  | 140 |  | -7.754% | -7.255% | -6.756% | -7.754% | -8.003% | -7.754% | -7.504% | -7.504% | -7.504% | -8.252% | -8.003% | -7.504% | -7.006% | -7.255% | -7.504% | -7.006% | -7.006% | -7.255% | -7.006% | -8.252% | -7.754% | -7.504% | -6.008% | -6.258% | -5.759% | -4.513% | -4.513% | -4.762% | -5.51% | -3.765% | -3.765% | -4.014% | -6.507% | -6.258% | -6.507% | -5.011% | -5.759% | -5.261% | -5.759%
|  | 150 |  | -7.644% | -7.644% | -6.601% | -7.383% | -8.166% | -8.166% | -7.905% | -7.122% | -8.166% | -8.427% | -7.644% | -7.383% | -7.122% | -7.644% | -7.644% | -7.905% | -6.861% | -7.122% | -7.122% | -7.644% | -7.905% | -7.383% | -5.557% | -5.818% | -5.818% | -4.774% | -4.253% | -4.513% | -5.557% | -2.687% | -3.992% | -4.253% | -6.601% | -6.34% | -6.079% | -5.035% | -5.818% | -5.035% | -5.818%
|  | 160 |  | -6.752% | -8.135% | -5.921% | -7.028% | -7.858% | -7.582% | -6.198% | -6.752% | -6.752% | -8.135% | -7.305% | -6.752% | -6.752% | -6.752% | -6.752% | -7.028% | -6.752% | -6.198% | -6.198% | -7.582% | -7.028% | -6.198% | -4.815% | -5.368% | -4.815% | -3.431% | -3.708% | -3.985% | -4.538% | -1.771% | -2.601% | -2.048% | -6.475% | -5.921% | -5.645% | -4.538% | -5.368% | -4.538% | -5.091%
|  | 170 |  | -6.667% | -6.087% | -5.507% | -6.667% | -6.957% | -6.957% | -6.957% | -6.667% | -6.957% | -7.826% | -6.667% | -6.087% | -6.087% | -6.667% | -6.377% | -6.377% | -5.797% | -5.507% | -5.217% | -7.536% | -6.377% | -6.957% | -4.638% | -4.348% | -4.058% | -2.319% | -2.899% | -3.188% | -4.058% | -0.8696% | -1.739% | -1.449% | -6.377% | -5.507% | -5.797% | -4.058% | -4.928% | -4.058% | -4.928%
|  | 180 |  | -7.357% | -6.449% | -6.146% | -7.357% | -8.265% | -7.054% | -6.751% | -7.054% | -7.66% | -8.871% | -7.357% | -6.146% | -6.449% | -6.751% | -6.751% | -6.751% | -5.54% | -5.843% | -5.54% | -7.357% | -7.357% | -6.751% | -4.935% | -4.329% | -4.632% | -2.21% | -2.21% | -2.513% | -4.027% | -0.9991% | -1.302% | -0.9991% | -7.054% | -6.146% | -6.751% | -4.027% | -4.935% | -4.027% | -4.935%
|  | 190 |  | -7.663% | -6.721% | -5.465% | -7.349% | -4.209% | -7.349% | -6.721% | -7.349% | -7.977% | -9.548% | -7.349% | -6.407% | -7.035% | -6.093% | -6.093% | -6.407% | -5.151% | -5.151% | -5.465% | -7.035% | -6.407% | -6.721% | -5.465% | -4.209% | -4.209% | -1.696% | -1.696% | -1.382% | -3.58% | 1.131% | 0.5025% | 0.1884% | -7.349% | -6.407% | -6.721% | -3.894% | -4.837% | -3.266% | -4.837%
|  | 200 |  | -6.382% | -4.762% | -4.762% | -6.382% | -8.001% | -6.382% | -5.086% | -6.058% | -7.353% | -8.973% | -6.382% | -5.086% | -5.41% | -5.41% | -4.762% | -5.086% | -4.438% | -3.79% | -3.142% | -6.706% | -5.734% | -6.706% | -4.114% | -3.142% | -2.818% | 0.4211% | 1.069% | 0.4211% | -2.17% | 4.956% | 4.632% | 3.013% | -7.029% | -5.41% | -6.382% | -1.846% | -2.818% | -1.523% | -3.79%
|  | 210 |  | -3.879% | -3.204% | -2.53% | -4.89% | -6.239% | -3.204% | -2.192% | -3.879% | -4.89% | -6.577% | -3.879% | -2.53% | -3.204% | -1.855% | -1.518% | -1.855% | -0.8432% | -0.5059% | 0.1686% | -3.879% | -2.192% | -3.879% | -1.518% | 0.8432% | 0.5059% | 4.89% | 5.565% | 4.553% | 0.8432% | 11.3% | 12.31% | 8.263% | -4.216% | -2.192% | -2.192% | 2.192% | 0.8432% | 2.53% | -0.8432%
|  | 220 |  | -1.662% | -0.9695% | 0.06925% | -2.355% | -3.047% | 0.06925% | 2.147% | 0.4155% | -1.662% | -3.047% | -0.6233% | 1.454% | -0.277% | 2.147% | 2.839% | 2.839% | 3.532% | 4.224% | 6.648% | -0.277% | 3.186% | -0.277% | 3.186% | 5.956% | 5.263% | 12.19% | 13.23% | 11.15% | 5.263% | 18.42% | 17.04% | 16% | 0.7618% | 5.263% | 3.532% | 8.726% | 6.994% | 9.072% | 4.917%
|  | 226 |  | -3.103% | -2.759% | -2.069% | -4.828% | -5.517% | -1.724% | -0.6897% | -1.034% | -3.793% | -4.138% | -2.414% | 0% | -1.724% | 0% | 1.379% | 0.6897% | 2.069% | 2.069% | 4.483% | -2.069% | 1.379% | -1.379% | 1.034% | 3.793% | 3.448% | 10.34% | 10.34% | 8.621% | 2.759% | 15.17% | 14.83% | 12.76% | 0% | 3.103% | 1.379% | 5.862% | 4.828% | 6.897% | 2.414%

## GTR2 G315 z0 y
GTR2 Spot size in y axis at iso for G315º every 10 MeV
Lynx PT
warrning | 8%
action level | 10%
[35 empty] | validation of new set point for deflector |  |  |  | After ECPM |  |  |  | After Cyclo PM
|  | E [MeV] | REF | [9 empty] | After Cyclo PM | After waterleak to cyclo | [12 empty] | old Lynx | old Lynx run 2 | old Lynx | Lynx SN 63021 | [6 empty] | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021 | Lynx SN 63021
|  |  |  | 2019-10-22 | 2019-12-07 | 2019-12-17 | 2020-01-14 | 2020-01-14 | 2020-02-06 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-08-02 | 2021-10-19 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27
|  |  |  | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm]
|  | 60 | 7.165 | 6.59 | 6.66 | 6.7 | 6.62 | 6.62 | 6.6 | 6.6 | 6.6 | 6.66 | 6.66 | 6.62 | 6.66 | 6.67 | 6.64 | 6.74 | 6.7 | 6.59 | 6.64 | 6.67 | 6.67 | 6.59 | 6.63 | 6.63 | 6.66 | 6.68 | 6.61 | 6.65 | 6.57 | 6.65 | 6.61 | 6.57 | 6.63 | 6.61 | 6.59 | 6.57 | 6.54 | 6.59 | 6.61 | 6.64 | 6.6 | 6.68 | 6.6
|  | 70 | 6.481 | 5.99 | 5.97 | 5.99 | 5.96 | 5.96 | 5.97 | 5.97 | 5.97 | 6.02 | 5.97 | 5.94 | 5.99 | 5.97 | 5.99 | 6.01 | 6.02 | 5.94 | 5.97 | 5.95 | 5.97 | 5.97 | 5.98 | 5.98 | 5.9 | 5.91 | 5.95 | 5.99 | 5.94 | 5.92 | 5.91 | 5.86 | 5.94 | 5.92 | 5.98 | 5.92 | 5.88 | 5.88 | 5.89 | 5.93 | 5.9 | 5.89 | 5.86
|  | 80 | 5.679 | 5.42 | 5.4 | 5.47 | 5.41 | 5.41 | 5.4 | 5.4 | 5.43 | 5.42 | 5.41 | 5.39 | 5.39 | 5.42 | 5.42 | 5.45 | 5.39 | 5.35 | 5.39 | 5.37 | 5.41 | 5.37 | 5.4 | 5.4 | 5.38 | 5.36 | 5.41 | 5.36 | 5.4 | 5.34 | 5.36 | 5.4 | 5.38 | 5.45 | 5.39 | 5.36 | 5.34 | 5.35 | 5.32 | 5.38 | 5.34 | 5.36 | 5.31
|  | 90 | 5.453 | 4.96 | 4.98 | 5 | 4.99 | 4.99 | 4.98 | 4.98 | 4.96 | 4.97 | 4.95 | 5 | 4.99 | 4.97 | 4.96 | 5 | 4.97 | 4.97 | 4.98 | 4.98 | 5 | 4.92 | 4.97 | 4.97 | 4.95 | 4.93 | 4.95 | 4.97 | 4.9 | 4.93 | 4.93 | 4.92 | 4.97 | 5.01 | 4.97 | 4.95 | 4.92 | 4.92 | 4.88 | 4.95 | 4.93 | 4.91 | 4.93
|  | 100 | 5.097 | 4.71 | 4.7 | 4.74 | 4.71 | 4.71 | 4.7 | 4.7 | 4.69 | 4.71 | 4.68 | 4.67 | 4.72 | 4.7 | 4.69 | 4.74 | 4.7 | 4.67 | 4.68 | 4.7 | 4.7 | 4.69 | 4.68 | 4.69 | 4.66 | 4.68 | 4.7 | 4.7 | 4.65 | 4.65 | 4.67 | 4.68 | 4.67 | 4.69 | 4.66 | 4.68 | 4.65 | 4.67 | 4.66 | 4.69 | 4.68 | 4.66 | 4.67
|  | 110 | 4.739 | 4.34 | 4.37 | 4.39 | 4.35 | 4.35 | 4.35 | 4.35 | 4.34 | 4.33 | 4.36 | 4.34 | 4.36 | 4.34 | 4.34 | 4.38 | 4.33 | 4.33 | 4.33 | 4.34 | 4.34 | 4.32 | 4.34 | 4.33 | 4.31 | 4.32 | 4.33 | 4.32 | 4.31 | 4.32 | 4.33 | 4.31 | 4.33 | 4.34 | 4.32 | 4.34 | 4.31 | 4.32 | 4.29 | 4.32 | 4.31 | 4.32 | 4.31
|  | 120 | 4.411 | 4.09 | 4.08 | 4.13 | 4.1 | 4.1 | 4.07 | 4.07 | 4.07 | 4.06 | 4.08 | 4.07 | 4.05 | 4.06 | 4.06 | 4.08 | 4.05 | 4.06 | 4.03 | 4.08 | 4.07 | 4.07 | 4.04 | 4.05 | 4.02 | 4.01 | 4.04 | 4.03 | 4.06 | 4.02 | 4.05 | 4.05 | 4.04 | 4.09 | 4.09 | 4.06 | 4.01 | 4.01 | 4.01 | 4.08 | 4.04 | 4.03 | 4.02
|  | 130 | 4.145 | 3.83 | 3.83 | 3.85 | 3.83 | 3.83 | 3.81 | 3.81 | 3.8 | 3.84 | 3.8 | 3.8 | 3.79 | 3.82 | 3.79 | 3.83 | 3.8 | 3.78 | 3.78 | 3.81 | 3.8 | 3.8 | 3.81 | 3.83 | 3.75 | 3.76 | 3.79 | 3.78 | 3.77 | 3.78 | 3.76 | 3.77 | 3.77 | 3.83 | 3.82 | 3.82 | 3.74 | 3.79 | 3.74 | 3.78 | 3.76 | 3.75 | 3.75
|  | 140 | 3.919 | 3.64 | 3.65 | 3.68 | 3.63 | 3.63 | 3.57 | 3.57 | 3.61 | 3.6 | 3.6 | 3.6 | 3.59 | 3.61 | 3.6 | 3.63 | 3.59 | 3.6 | 3.61 | 3.62 | 3.63 | 3.67 | 3.59 | 3.63 | 3.58 | 3.57 | 3.58 | 3.59 | 3.59 | 3.61 | 3.6 | 3.61 | 3.61 | 3.7 | 3.7 | 3.67 | 3.55 | 3.6 | 3.58 | 3.6 | 3.59 | 3.58 | 3.58
|  | 150 | 3.748 | 3.44 | 3.47 | 3.47 | 3.44 | 3.44 | 3.4 | 3.4 | 3.41 | 3.42 | 3.43 | 3.4 | 3.4 | 3.42 | 3.41 | 3.44 | 3.42 | 3.42 | 3.42 | 3.44 | 3.44 | 3.45 | 3.43 | 3.43 | 3.37 | 3.39 | 3.4 | 3.42 | 3.4 | 3.41 | 3.43 | 3.41 | 3.42 | 3.53 | 3.51 | 3.49 | 3.37 | 3.38 | 3.38 | 3.42 | 3.4 | 3.39 | 3.39
|  | 160 | 3.552 | 3.27 | 3.28 | 3.32 | 3.26 | 3.26 | 3.2 | 3.2 | 3.23 | 3.26 | 3.25 | 3.22 | 3.22 | 3.25 | 3.23 | 3.25 | 3.24 | 3.24 | 3.24 | 3.26 | 3.28 | 3.27 | 3.25 | 3.26 | 3.21 | 3.21 | 3.23 | 3.23 | 3.22 | 3.25 | 3.25 | 3.26 | 3.24 | 3.38 | 3.32 | 3.32 | 3.2 | 3.21 | 3.2 | 3.23 | 3.23 | 3.23 | 3.21
|  | 170 | 3.358 | 3.1 | 3.1 | 3.14 | 3.09 | 3.09 | 3.04 | 3.04 | 3.05 | 3.06 | 3.07 | 3.05 | 3.05 | 3.06 | 3.06 | 3.08 | 3.06 | 3.06 | 3.06 | 3.08 | 3.09 | 3.11 | 3.07 | 3.1 | 3.02 | 3.02 | 3.04 | 3.06 | 3.05 | 3.09 | 3.07 | 3.07 | 3.05 | 3.2 | 3.13 | 3.15 | 3.01 | 3.02 | 3.01 | 3.05 | 3.04 | 3.05 | 3.04
|  | 180 | 3.246 | 3.02 | 3.05 | 3.06 | 3.01 | 3.01 | 2.94 | 2.94 | 2.97 | 2.98 | 2.98 | 2.95 | 2.95 | 2.97 | 2.97 | 2.99 | 2.98 | 2.98 | 2.99 | 3.02 | 3.02 | 3.02 | 2.99 | 2.97 | 2.93 | 2.95 | 2.96 | 2.99 | 2.98 | 3.03 | 3.03 | 3.01 | 2.96 | 3.12 | 3.09 | 3.11 | 2.92 | 2.94 | 2.94 | 2.99 | 2.95 | 2.97 | 2.96
|  | 190 | 3.142 | 2.95 | 2.98 | 2.98 | 2.92 | 2.92 | 2.94 | 2.94 | 2.87 | 2.89 | 2.9 | 2.84 | 2.85 | 2.88 | 2.89 | 2.9 | 2.9 | 2.91 | 2.9 | 2.94 | 2.94 | 2.95 | 2.9 | 2.93 | 2.84 | 2.85 | 2.87 | 2.89 | 2.91 | 2.94 | 2.97 | 2.94 | 2.87 | 3.1 | 3.05 | 3.05 | 2.82 | 2.85 | 2.85 | 2.91 | 2.88 | 2.89 | 2.88
|  | 200 | 3.059 | 2.91 | 2.94 | 2.94 | 2.85 | 2.85 | 2.78 | 2.78 | 2.81 | 2.83 | 2.82 | 2.77 | 2.76 | 2.81 | 2.82 | 2.83 | 2.85 | 2.85 | 2.84 | 2.89 | 2.89 | 2.91 | 2.83 | 2.87 | 2.76 | 2.77 | 2.78 | 2.83 | 2.84 | 2.9 | 2.92 | 2.89 | 2.8 | 3.1 | 3.08 | 3.05 | 2.74 | 2.78 | 2.78 | 2.85 | 2.82 | 2.84 | 2.8
|  | 210 | 2.929 | 2.84 | 2.88 | 2.86 | 2.76 | 2.76 | 2.68 | 2.68 | 2.74 | 2.76 | 2.73 | 2.68 | 2.66 | 2.71 | 2.73 | 2.73 | 2.74 | 2.76 | 2.77 | 2.8 | 2.81 | 2.85 | 2.74 | 2.8 | 2.66 | 2.66 | 2.68 | 2.73 | 2.75 | 2.83 | 2.86 | 2.82 | 2.69 | 3.13 | 3.08 | 3.04 | 2.63 | 2.68 | 2.67 | 2.78 | 2.74 | 2.76 | 2.72
|  | 220 | 2.831 | 2.83 | 2.9 | 2.86 | 2.73 | 2.73 | 2.62 | 2.62 | 2.7 | 2.74 | 2.71 | 2.62 | 2.6 | 2.67 | 2.7 | 2.69 | 2.71 | 2.74 | 2.75 | 2.78 | 2.8 | 2.88 | 2.69 | 2.78 | 2.6 | 2.59 | 2.63 | 2.7 | 2.72 | 2.84 | 2.85 | 2.8 | 2.63 | 3.28 | 3.23 | 3.17 | 2.61 | 2.68 | 2.66 | 2.76 | 2.73 | 2.76 | 2.68
|  | 226 | 2.76 | 2.77 | 2.78 | 2.78 | 2.66 | 2.66 | 2.56 | 2.56 | 2.65 | 2.69 | 2.66 | 2.58 | 2.54 | 2.63 | 2.66 | 2.64 | 2.67 | 2.68 | 2.68 | 2.73 | 2.77 | 2.81 | 2.74 | 2.73 | 2.53 | 2.54 | 2.56 | 2.64 | 2.65 | 2.76 | 2.78 | 2.71 | 2.56 | 3.23 | 3.17 | 3.12 | 2.54 | 2.62 | 2.61 | 2.7 | 2.67 | 2.68 | 2.61
Deviation from reference [%]
|  | E [MeV] | note | [9 empty] | After Cyclo PM | After waterleak to cyclo | [15 empty] | Lynx SN 63021 | Lynx SN 63021
|  |  |  | 2019-10-22 | 2019-12-07 | 2019-12-17 | 2020-01-14 | 2020-01-14 | 2020-02-06 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-08-02 | 2021-10-19 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm] | sigma y [mm]
|  | 60 |  | -8.025% | -7.048% | -6.49% | -7.606% | -7.606% | -7.886% | -7.886% | -7.886% | -7.048% | -7.048% | -7.606% | -7.048% | -6.909% | -7.327% | -5.932% | -6.49% | -8.025% | -7.327% | -6.909% | -6.909% | -8.025% | -7.467% | -7.467% | -7.048% | -6.769% | -7.746% | -7.188% | -8.304% | -7.188% | -7.746% | -8.304% | -7.467% | -7.746% | -8.025% | -8.304% | -8.723% | -8.025% | -7.746% | -7.327% | -7.886% | -6.769% | -7.886%
|  | 70 |  | -7.576% | -7.885% | -7.576% | -8.039% | -8.039% | -7.885% | -7.885% | -7.885% | -7.113% | -7.885% | -8.347% | -7.576% | -7.885% | -7.576% | -7.267% | -7.113% | -8.347% | -7.885% | -8.193% | -7.885% | -7.885% | -7.73% | -7.73% | -8.965% | -8.81% | -8.193% | -7.576% | -8.347% | -8.656% | -8.81% | -9.582% | -8.347% | -8.656% | -7.73% | -8.656% | -9.273% | -9.273% | -9.119% | -8.502% | -8.965% | -9.119% | -9.582%
|  | 80 |  | -4.561% | -4.913% | -3.68% | -4.737% | -4.737% | -4.913% | -4.913% | -4.385% | -4.561% | -4.737% | -5.089% | -5.089% | -4.561% | -4.561% | -4.032% | -5.089% | -5.793% | -5.089% | -5.441% | -4.737% | -5.441% | -4.913% | -4.913% | -5.265% | -5.617% | -4.737% | -5.617% | -4.913% | -5.969% | -5.617% | -4.913% | -5.265% | -4.032% | -5.089% | -5.617% | -5.969% | -5.793% | -6.322% | -5.265% | -5.969% | -5.617% | -6.498%
|  | 90 |  | -9.041% | -8.674% | -8.307% | -8.491% | -8.491% | -8.674% | -8.674% | -9.041% | -8.858% | -9.224% | -8.307% | -8.491% | -8.858% | -9.041% | -8.307% | -8.858% | -8.858% | -8.674% | -8.674% | -8.307% | -9.774% | -8.858% | -8.858% | -9.224% | -9.591% | -9.224% | -8.858% | -10.14% | -9.591% | -9.591% | -9.774% | -8.858% | -8.124% | -8.858% | -9.224% | -9.774% | -9.774% | -10.51% | -9.224% | -9.591% | -9.958% | -9.591%
|  | 100 |  | -7.593% | -7.789% | -7.004% | -7.593% | -7.593% | -7.789% | -7.789% | -7.985% | -7.593% | -8.181% | -8.377% | -7.397% | -7.789% | -7.985% | -7.004% | -7.789% | -8.377% | -8.181% | -7.789% | -7.789% | -7.985% | -8.181% | -7.985% | -8.574% | -8.181% | -7.789% | -7.789% | -8.77% | -8.77% | -8.377% | -8.181% | -8.377% | -7.985% | -8.574% | -8.181% | -8.77% | -8.377% | -8.574% | -7.985% | -8.181% | -8.574% | -8.377%
|  | 110 |  | -8.419% | -7.786% | -7.364% | -8.208% | -8.208% | -8.208% | -8.208% | -8.419% | -8.631% | -7.997% | -8.419% | -7.997% | -8.419% | -8.419% | -7.575% | -8.631% | -8.631% | -8.631% | -8.419% | -8.419% | -8.842% | -8.419% | -8.631% | -9.053% | -8.842% | -8.631% | -8.842% | -9.053% | -8.842% | -8.631% | -9.053% | -8.631% | -8.419% | -8.842% | -8.419% | -9.053% | -8.842% | -9.475% | -8.842% | -9.053% | -8.842% | -9.053%
|  | 120 |  | -7.277% | -7.504% | -6.37% | -7.051% | -7.051% | -7.731% | -7.731% | -7.731% | -7.957% | -7.504% | -7.731% | -8.184% | -7.957% | -7.957% | -7.504% | -8.184% | -7.957% | -8.637% | -7.504% | -7.731% | -7.731% | -8.411% | -8.184% | -8.864% | -9.091% | -8.411% | -8.637% | -7.957% | -8.864% | -8.184% | -8.184% | -8.411% | -7.277% | -7.277% | -7.957% | -9.091% | -9.091% | -9.091% | -7.504% | -8.411% | -8.637% | -8.864%
|  | 130 |  | -7.6% | -7.6% | -7.117% | -7.6% | -7.6% | -8.082% | -8.082% | -8.323% | -7.358% | -8.323% | -8.323% | -8.565% | -7.841% | -8.565% | -7.6% | -8.323% | -8.806% | -8.806% | -8.082% | -8.323% | -8.323% | -8.082% | -7.6% | -9.53% | -9.288% | -8.565% | -8.806% | -9.047% | -8.806% | -9.288% | -9.047% | -9.047% | -7.6% | -7.841% | -7.841% | -9.771% | -8.565% | -9.771% | -8.806% | -9.288% | -9.53% | -9.53%
|  | 140 |  | -7.119% | -6.864% | -6.098% | -7.374% | -7.374% | -8.905% | -8.905% | -7.885% | -8.14% | -8.14% | -8.14% | -8.395% | -7.885% | -8.14% | -7.374% | -8.395% | -8.14% | -7.885% | -7.629% | -7.374% | -6.354% | -8.395% | -7.374% | -8.65% | -8.905% | -8.65% | -8.395% | -8.395% | -7.885% | -8.14% | -7.885% | -7.885% | -5.588% | -5.588% | -6.354% | -9.416% | -8.14% | -8.65% | -8.14% | -8.395% | -8.65% | -8.65%
|  | 150 |  | -8.218% | -7.417% | -7.417% | -8.218% | -8.218% | -9.285% | -9.285% | -9.018% | -8.751% | -8.485% | -9.285% | -9.285% | -8.751% | -9.018% | -8.218% | -8.751% | -8.751% | -8.751% | -8.218% | -8.218% | -7.951% | -8.485% | -8.485% | -10.09% | -9.552% | -9.285% | -8.751% | -9.285% | -9.018% | -8.485% | -9.018% | -8.751% | -5.816% | -6.35% | -6.884% | -10.09% | -9.819% | -9.819% | -8.751% | -9.285% | -9.552% | -9.552%
|  | 160 |  | -7.939% | -7.658% | -6.532% | -8.221% | -8.221% | -9.91% | -9.91% | -9.065% | -8.221% | -8.502% | -9.347% | -9.347% | -8.502% | -9.065% | -8.502% | -8.784% | -8.784% | -8.784% | -8.221% | -7.658% | -7.939% | -8.502% | -8.221% | -9.628% | -9.628% | -9.065% | -9.065% | -9.347% | -8.502% | -8.502% | -8.221% | -8.784% | -4.842% | -6.532% | -6.532% | -9.91% | -9.628% | -9.91% | -9.065% | -9.065% | -9.065% | -9.628%
|  | 170 |  | -7.683% | -7.683% | -6.492% | -7.981% | -7.981% | -9.47% | -9.47% | -9.172% | -8.874% | -8.577% | -9.172% | -9.172% | -8.874% | -8.874% | -8.279% | -8.874% | -8.874% | -8.874% | -8.279% | -7.981% | -7.385% | -8.577% | -7.683% | -10.07% | -10.07% | -9.47% | -8.874% | -9.172% | -7.981% | -8.577% | -8.577% | -9.172% | -4.705% | -6.79% | -6.194% | -10.36% | -10.07% | -10.36% | -9.172% | -9.47% | -9.172% | -9.47%
|  | 180 |  | -6.962% | -6.038% | -5.73% | -7.27% | -7.27% | -9.427% | -9.427% | -8.503% | -8.195% | -8.195% | -9.119% | -9.119% | -8.503% | -8.503% | -7.887% | -8.195% | -8.195% | -7.887% | -6.962% | -6.962% | -6.962% | -7.887% | -8.503% | -9.735% | -9.119% | -8.811% | -7.887% | -8.195% | -6.654% | -6.654% | -7.27% | -8.811% | -3.882% | -4.806% | -4.19% | -10.04% | -9.427% | -9.427% | -7.887% | -9.119% | -8.503% | -8.811%
|  | 190 |  | -6.111% | -5.156% | -5.156% | -7.066% | -7.066% | -6.429% | -6.429% | -8.657% | -8.02% | -7.702% | -9.612% | -9.293% | -8.339% | -8.02% | -7.702% | -7.702% | -7.384% | -7.702% | -6.429% | -6.429% | -6.111% | -7.702% | -6.747% | -9.612% | -9.293% | -8.657% | -8.02% | -7.384% | -6.429% | -5.474% | -6.429% | -8.657% | -1.337% | -2.928% | -2.928% | -10.25% | -9.293% | -9.293% | -7.384% | -8.339% | -8.02% | -8.339%
|  | 200 |  | -4.871% | -3.89% | -3.89% | -6.832% | -6.832% | -9.121% | -9.121% | -8.14% | -7.486% | -7.813% | -9.448% | -9.774% | -8.14% | -7.813% | -7.486% | -6.832% | -6.832% | -7.159% | -5.525% | -5.525% | -4.871% | -7.486% | -6.178% | -9.774% | -9.448% | -9.121% | -7.486% | -7.159% | -5.198% | -4.544% | -5.525% | -8.467% | 1.34% | 0.6865% | -0.2942% | -10.43% | -9.121% | -9.121% | -6.832% | -7.813% | -7.159% | -8.467%
|  | 210 |  | -3.039% | -1.673% | -2.356% | -5.77% | -5.77% | -8.501% | -8.501% | -6.453% | -5.77% | -6.794% | -8.501% | -9.184% | -7.477% | -6.794% | -6.794% | -6.453% | -5.77% | -5.428% | -4.404% | -4.063% | -2.697% | -6.453% | -4.404% | -9.184% | -9.184% | -8.501% | -6.794% | -6.111% | -3.38% | -2.356% | -3.721% | -8.16% | 6.862% | 5.155% | 3.79% | -10.21% | -8.501% | -8.843% | -5.087% | -6.453% | -5.77% | -7.136%
|  | 220 |  | -0.03532% | 2.437% | 1.024% | -3.568% | -3.568% | -7.453% | -7.453% | -4.627% | -3.214% | -4.274% | -7.453% | -8.16% | -5.687% | -4.627% | -4.981% | -4.274% | -3.214% | -2.861% | -1.801% | -1.095% | 1.731% | -4.981% | -1.801% | -8.16% | -8.513% | -7.1% | -4.627% | -3.921% | 0.3179% | 0.6711% | -1.095% | -7.1% | 15.86% | 14.09% | 11.97% | -7.806% | -5.334% | -6.04% | -2.508% | -3.568% | -2.508% | -5.334%
|  | 226 |  | 0.3623% | 0.7246% | 0.7246% | -3.623% | -3.623% | -7.246% | -7.246% | -3.986% | -2.536% | -3.623% | -6.522% | -7.971% | -4.71% | -3.623% | -4.348% | -3.261% | -2.899% | -2.899% | -1.087% | 0.3623% | 1.812% | -0.7246% | -1.087% | -8.333% | -7.971% | -7.246% | -4.348% | -3.986% | 0% | 0.7246% | -1.812% | -7.246% | 17.03% | 14.86% | 13.04% | -7.971% | -5.072% | -5.435% | -2.174% | -3.261% | -2.899% | -5.435%

## GTR2 Skewness
GTR2 skewness at iso for G315º every 10 MeV
Lynx PT
warrning
action level
Sel parameter [%] for ellipse
[33 empty] | validation of new set point for deflector | [7 empty] | After Cyclo PM
| E [MeV] | 2019-04-22 | 2019-07-01 | 2019-12-17 | 2020-01-14 | 2020-01-14 | 2020-02-06 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-08-02 | 2021-10-19 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27
| 60 | 1.4 | 2.2 | 0.9 | 1.3 | 1.3 | 1.4 | 1.5 | 1.5 | 0.9 | 1.1 | 1.7 | 1.419 | 1.1 | 1.4 | 0.6 | 0.8 | 1.5 | 1.3 | 1.245 | 1.317 | 1.597 | 1.398 | 1.155 | 1.397 | 1.383 | 2.794 | 2.049 | 2.471 | 2.156 | 2.459 | 2.968 | 2.332 | 2.438 | 2.57 | 2.679 | 3.195 | 2.621 | 2.617 | 2.309 | 2.63 | 2.173 | 2.52
| 70 | 2.2 | 2.4 | 1.7 | 1.5 | 1.5 | 1.4 | 1.4 | 1.6 | 0.8 | 1.5 | 1.9 | 1.612 | 1.5 | 1.372 | 1.3 | 1 | 1.6 | 1.6 | 2.006 | 1.971 | 1.296 | 1.563 | 1.28 | 2.516 | 2.523 | 3.046 | 2.3 | 2.604 | 2.925 | 3.144 | 3.528 | 2.743 | 2.989 | 2.32 | 3.035 | 3.374 | 3.341 | 3.358 | 3.022 | 3.132 | 3.334 | 3.51
| 80 | 2.9 | 3.1 | 2 | 2.4 | 2.4 | 2.1 | 2.3 | 2.3 | 2 | 2.4 | 2.7 | 2.767 | 2 | 2.089 | 2.1 | 2.5 | 2.8 | 2.5 | 2.742 | 2.746 | 2.548 | 2.621 | 2.271 | 3.095 | 3.208 | 4.049 | 3.791 | 3.409 | 3.932 | 3.998 | 3.541 | 3.742 | 2.956 | 3.57 | 3.819 | 4.161 | 3.988 | 4.649 | 3.88 | 4.361 | 4.195 | 4.542
| 90 | 1.1 | 1.7 | 1.1 | 1 | 1 | 1.5 | 0.9 | 1.3 | 1.3 | 1.3 | 1 | 1.05 | 1.1 | 1.463 | 1.1 | 1.1 | 1 | 0.9 | 1.344 | 1.111 | 1.681 | 1.045 | 0.874 | 1.624 | 1.874 | 2.706 | 1.996 | 2.718 | 2.687 | 2.591 | 2.682 | 2.287 | 1.927 | 2.33 | 2.535 | 2.949 | 2.857 | 3.185 | 2.555 | 2.785 | 3.138 | 2.88
| 100 | 4.4 | 4.8 | 3.8 | 3.9 | 3.9 | 3.7 | 3.8 | 4 | 3.7 | 4 | 4.5 | 4.004 | 3.6 | 3.678 | 3.6 | 4 | 4.1 | 4.2 | 4.119 | 4.423 | 3.844 | 4.252 | 3.961 | 4.87 | 4.7 | 5.673 | 5.016 | 5.624 | 5.474 | 5.525 | 5.499 | 5.551 | 5.166 | 5.76 | 5.605 | 5.713 | 5.883 | 6.224 | 5.88 | 5.658 | 5.972 | 5.735
| 110 | 4.6 | 5.3 | 4.1 | 4.3 | 4.3 | 3.9 | 4 | 4.2 | 4.2 | 3.8 | 4.3 | 4.185 | 3.8 | 3.922 | 3.8 | 4.4 | 4.2 | 4.4 | 4.245 | 4.698 | 4.124 | 4.37 | 4.286 | 5.237 | 4.764 | 5.895 | 5.525 | 5.888 | 5.841 | 5.817 | 5.925 | 5.718 | 5.772 | 6.13 | 5.817 | 6.059 | 6.103 | 6.826 | 6.336 | 6.276 | 6.347 | 6.38
| 120 | 5.1 | 5.6 | 4.4 | 4.7 | 4.7 | 4.6 | 4.6 | 4.6 | 4.8 | 4.6 | 5 | 5.068 | 4.5 | 4.56 | 4.6 | 5.1 | 4.8 | 5.2 | 4.775 | 5.195 | 4.685 | 5.123 | 4.958 | 5.704 | 5.528 | 6.805 | 6.391 | 6.105 | 6.799 | 6.476 | 6.497 | 6.213 | 6.148 | 6.42 | 6.621 | 7.01 | 6.992 | 7.283 | 6.564 | 7.007 | 6.936 | 7.149
| 130 | 3.3 | 3.7 | 2.5 | 2.4 | 2.4 | 2.8 | 2.4 | 2.6 | 2.1 | 2.8 | 2.7 | 2.819 | 2.2 | 2.77 | 2.4 | 2.8 | 2.9 | 3.2 | 2.91 | 3.22 | 2.727 | 2.746 | 2.325 | 3.77 | 3.573 | 4.666 | 4.262 | 4.61 | 4.872 | 5.288 | 4.927 | 4.626 | 4.627 | 4.91 | 4.763 | 4.965 | 4.49 | 5.43 | 5.052 | 5.244 | 5.518 | 5.375
| 140 | 1.7 | 2 | 0.8 | 1 | 1 | 1.5 | 1.7 | 1.3 | 1.6 | 1.5 | 1.6 | 1.624 | 1.1 | 1.55 | 1.4 | 2 | 1.6 | 1.8 | 1.573 | 1.403 | 0.706 | 1.333 | 1.07 | 2.073 | 2.299 | 3.346 | 2.784 | 3.221 | 3.489 | 3.857 | 3.411 | 2.754 | 2.67 | 2.75 | 3.105 | 3.674 | 3.074 | 3.599 | 3.628 | 3.582 | 3.979 | 3.679
| 150 | 1.5 | 2.1 | 1.5 | 1.6 | 1.6 | 2.1 | 1.7 | 1.6 | 1.6 | 1.9 | 1.8 | 1.667 | 1.7 | 2.09 | 1.7 | 1.8 | 1.7 | 1.7 | 1.84 | 1.697 | 1.573 | 1.573 | 1.407 | 2.567 | 2.251 | 3.354 | 2.903 | 3.328 | 3.691 | 3.67 | 3.76 | 3.058 | 3.018 | 2.8 | 3.064 | 3.535 | 3.611 | 3.985 | 3.488 | 3.561 | 4.262 | 3.7
| 160 | 2 | 1.6 | 2.1 | 2.2 | 2.2 | 2.6 | 2.4 | 2.3 | 2.7 | 2.7 | 2.7 | 1.873 | 2.6 | 2.622 | 2.6 | 2.4 | 2.3 | 2.2 | 2.229 | 2.11 | 2.33 | 2.078 | 2.102 | 2.482 | 2.954 | 3.135 | 2.795 | 3.305 | 3.523 | 3.478 | 3.113 | 3.093 | 2.461 | 2.98 | 3.216 | 2.818 | 2.888 | 3.2 | 3.267 | 2.949 | 3.332 | 3.3
| 170 | 3.1 | 2.3 | 2.8 | 2.8 | 2.8 | 3.6 | 3.5 | 3.2 | 3.1 | 3.4 | 3.2 | 2.679 | 3.7 | 3.554 | 3.4 | 3 | 3 | 3 | 3.279 | 3.044 | 3.146 | 2.641 | 3.128 | 3.277 | 3.351 | 3.962 | 3.862 | 4.15 | 4.328 | 4.353 | 4.102 | 4.303 | 3.202 | 4.04 | 4.027 | 3.556 | 3.831 | 3.798 | 4.116 | 3.93 | 4.177 | 3.772
| 180 | 2.9 | 2.9 | 3.7 | 3.4 | 3.4 | 4.4 | 3.8 | 4 | 3.7 | 4.2 | 3.8 | 3.393 | 4.4 | 4.401 | 4.4 | 3.6 | 3.5 | 3.5 | 4.018 | 3.37 | 3.946 | 3.616 | 3.638 | 3.565 | 3.916 | 3.506 | 3.523 | 3.425 | 3.64 | 3.634 | 3.735 | 4.277 | 2.532 | 2.77 | 2.708 | 2.945 | 2.986 | 2.345 | 3.309 | 3.216 | 3.448 | 3.258
| 190 | 3.3 | 2.6 | 4.2 | 3.6 | 3.6 | 4.3 | 3.6 | 3.9 | 3.6 | 4 | 4 | 3.246 | 4.5 | 4.306 | 4.3 | 3.7 | 3.5 | 3.3 | 4.089 | 3.45 | 3.759 | 3.596 | 3.573 | 3.19 | 4.042 | 2.944 | 3.616 | 3.218 | 3.527 | 3.623 | 3.8 | 4.722 | 2.443 | 2.59 | 2.537 | 2.444 | 2.453 | 2.045 | 2.902 | 2.934 | 3.348 | 2.865
| 200 | 3 | 2.3 | 3.7 | 2.9 | 2.9 | 4 | 3.1 | 3.2 | 3.4 | 3.5 | 3.3 | 2.95 | 4.1 | 4.026 | 3.9 | 2.8 | 3.1 | 2.8 | 3.531 | 2.823 | 3.565 | 3.067 | 2.793 | 3.139 | 3.269 | 3.225 | 3.279 | 3.052 | 3.574 | 3.802 | 3.805 | 4.769 | 2.317 | 2.41 | 2.074 | 2.263 | 2.432 | 2.198 | 3.168 | 3.156 | 3.418 | 2.855
| 210 | 2.3 | 1.9 | 2.6 | 1.8 | 1.8 | 3.4 | 2.4 | 3 | 2.8 | 2.6 | 3 | 2.715 | 3.5 | 3.365 | 3.4 | 3.2 | 3.1 | 2.7 | 2.777 | 2.564 | 2.411 | 2.649 | 2.106 | 3.505 | 3.687 | 4.171 | 4.501 | 4.041 | 4.672 | 4.545 | 4.598 | 5.489 | 2.974 | 4.03 | 3.673 | 4.049 | 4.005 | 4.366 | 4.384 | 4.606 | 4.952 | 4.018
| 220 | 2.9 | 3.7 | 1.2 | 3 | 3 | 4.4 | 3.5 | 3.8 | 4.4 | 3.9 | 4 | 3.76 | 3.7 | 4.414 | 3.5 | 4.7 | 4.9 | 4.6 | 4.212 | 4.673 | 3.853 | 3.501 | 4.13 | 5.443 | 5.408 | 6.722 | 7.298 | 6.883 | 8.625 | 8.013 | 8.024 | 7.498 | 7.56 | 7.97 | 8.461 | 7.576 | 8.24 | 8.246 | 8.2 | 8.36 | 8.455 | 7.322
| 226 | 3 | 3.9 | 1.2 | 2.8 | 2.8 | 4.4 | 3.4 | 3.8 | 4.3 | 4.1 | 4 | 4.477 | 3.7 | 4.413 | 3.8 | 4.7 | 5.1 | 4.9 | 4.367 | 4.67 | 4.415 | 2.594 | 4.188 | 5.591 | 5.954 | 7.088 | 7.421 | 7.382 | 9.357 | 8.696 | 8.677 | 7.688 | 8.778 | 9.09 | 9.389 | 8.473 | 8.757 | 8.965 | 8.99 | 9.029 | 9.65 | 7.908
Correlation coefficient
| E [MeV] | 2019-04-22 | 2019-07-01 | 2019-12-17 | 2020-01-14 | 2020-01-14 | 2020-02-06 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-08-02 | 2021-10-19 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27
| 60 | 0.022 | 0.032 | 0.012 | 0.018 | 0.018 | 0.011 | 0.016 | 0.017 | 0.01 | 0.013 | 0.02 | 0.016 | 0.01 | 0.011 | 0.007 | 0.01 | 0.017 | 0.019 | 0.019 | 0.022 | 0.022 | 0.02 | 0.017 | 0.023 | 0.021 | 0.039 | 0.031 | 0.036 | 0.034 | 0.036 | 0.038 | 0.034 | 0.034 |  | -0.035 | -0.043 | -0.039 | -0.044 | -0.04 | -0.043 | -0.04 | -0.04
| 70 | 0.033 | 0.035 | 0.02 | 0.021 | 0.021 | 0.013 | 0.017 | 0.02 | 0.014 | 0.015 | 0.024 | 0.02 | 0.013 | 0.013 | 0.014 | 0.014 | 0.021 | 0.023 | 0.027 | 0.03 | 0.02 | 0.022 | 0.018 | 0.037 | 0.035 | 0.045 | 0.036 | 0.04 | 0.044 | 0.046 | 0.05 | 0.04 | 0.04 |  | -0.043 | -0.049 | -0.05 | -0.055 | -0.05 | -0.051 | -0.054 | -0.053
| 80 | 0.055 | 0.056 | 0.039 | 0.045 | 0.045 | 0.038 | 0.043 | 0.044 | 0.039 | 0.045 | 0.048 | 0.051 | 0.037 | 0.038 | 0.04 | 0.046 | 0.05 | 0.047 | 0.05 | 0.053 | 0.048 | 0.05 | 0.045 | 0.059 | 0.059 | 0.074 | 0.069 | 0.064 | 0.069 | 0.071 | 0.065 | 0.066 | 0.055 |  | -0.066 | -0.074 | -0.073 | -0.085 | -0.072 | -0.079 | -0.077 | -0.082
| 90 | 0.013 | 0.016 | 0.004 | 0.004 | 0.004 | 0.006 | 0.001 | 0.008 | 0.008 | 0.007 | 0.005 | 0.01 | 0 | 0.006 | 0.003 | 0.003 | 0.005 | 0.005 | 0.011 | 0.012 | 0.017 | 0.009 | 0.006 | 0.021 | 0.018 | 0.034 | 0.024 | 0.033 | 0.031 | 0.029 | 0.029 | 0.024 | 0.017 |  | -0.026 | -0.038 | -0.038 | -0.046 | -0.035 | -0.038 | -0.042 | -0.039
| 100 | 0.084 | 0.089 | 0.07 | 0.074 | 0.074 | 0.066 | 0.071 | 0.074 | 0.069 | 0.074 | 0.083 | 0.076 | 0.068 | 0.066 | 0.066 | 0.072 | 0.074 | 0.077 | 0.076 | 0.083 | 0.071 | 0.079 | 0.075 | 0.091 | 0.086 | 0.103 | 0.092 | 0.101 | 0.095 | 0.097 | 0.097 | 0.097 | 0.088 |  | -0.096 | -0.103 | -0.107 | -0.115 | -0.106 | -0.103 | -0.107 | -0.105
| 110 | 0.085 | 0.099 | 0.077 | 0.079 | 0.079 | 0.072 | 0.074 | 0.078 | 0.077 | 0.071 | 0.08 | 0.08 | 0.07 | 0.072 | 0.07 | 0.079 | 0.078 | 0.081 | 0.078 | 0.088 | 0.075 | 0.082 | 0.081 | 0.097 | 0.087 | 0.105 | 0.099 | 0.104 | 0.103 | 0.101 | 0.103 | 0.1 | 0.1 |  | -0.102 | -0.107 | -0.11 | -0.122 | -0.114 | -0.112 | -0.114 | -0.114
| 120 | 0.097 | 0.104 | 0.085 | 0.089 | 0.089 | 0.084 | 0.089 | 0.087 | 0.091 | 0.086 | 0.096 | 0.098 | 0.085 | 0.084 | 0.086 | 0.093 | 0.089 | 0.096 | 0.089 | 0.097 | 0.088 | 0.096 | 0.094 | 0.108 | 0.102 | 0.123 | 0.116 | 0.111 | 0.119 | 0.114 | 0.113 | 0.11 | 0.109 |  | -0.117 | -0.128 | -0.127 | -0.134 | -0.119 | -0.126 | -0.124 | -0.129
| 130 | 0.055 | 0.061 | 0.04 | 0.039 | 0.039 | 0.041 | 0.037 | 0.042 | 0.032 | 0.044 | 0.043 | 0.048 | 0.033 | 0.041 | 0.037 | 0.043 | 0.048 | 0.051 | 0.045 | 0.054 | 0.044 | 0.046 | 0.041 | 0.063 | 0.053 | 0.073 | 0.067 | 0.07 | 0.072 | 0.076 | 0.073 | 0.066 | 0.071 |  | -0.074 | -0.078 | -0.071 | -0.09 | -0.08 | -0.083 | -0.087 | -0.083
| 140 | 0.022 | 0.021 | 0.004 | 0.005 | 0.005 | 0.009 | 0.013 | 0.009 | 0.013 | 0.007 | 0.014 | 0.021 | 0 | 0.009 | 0.007 | 0.017 | 0.016 | 0.016 | 0.012 | 0.016 | 0 | 0.014 | 0.01 | 0.022 | 0.024 | 0.041 | 0.033 | 0.038 | 0.037 | 0.044 | 0.038 | 0.025 | 0.033 |  | -0.039 | -0.049 | -0.043 | -0.054 | -0.047 | -0.049 | -0.052 | -0.05
| 150 | 0.002 | 0.016 | -0.001 | -0.002 | -0.002 | -0.005 | -0.004 | -0.003 | -0.003 | -0.013 | 0.003 | 0.007 | -0.01 | -0.005 | -0.008 | 0.003 | -0.001 | 0.003 | -0.004 | 0.003 | -0.007 | 0.004 | -0.004 | 0.017 | 0.002 | 0.027 | 0.02 | 0.026 | 0.029 | 0.028 | 0.029 | 0.016 | 0.026 |  | -0.035 | -0.038 | -0.038 | -0.048 | -0.031 | -0.038 | -0.045 | -0.038
| 160 | -0.023 | -0.016 | -0.035 | -0.031 | -0.031 | -0.033 | -0.026 | -0.03 | -0.037 | -0.04 | -0.033 | -0.023 | -0.042 | -0.033 | -0.038 | -0.028 | -0.022 | -0.025 | -0.03 | -0.024 | -0.03 | -0.03 | -0.03 | -0.016 | -0.027 | 0.001 | -0.005 | -0.002 | 0.003 | 0.001 | 0 | -0.015 | -0.002 |  | -0.01 | -0.007 | -0.007 | -0.017 | -0.011 | -0.011 | -0.013 | -0.01
| 170 | -0.035 | -0.013 | -0.041 | -0.039 | -0.039 | -0.041 | -0.043 | -0.038 | -0.039 | -0.047 | -0.038 | -0.033 | -0.053 | -0.041 | -0.047 | -0.028 | -0.032 | -0.03 | -0.037 | -0.027 | -0.041 | -0.034 | -0.047 | -0.02 | -0.027 | -0.005 | -0.009 | -0.007 | 0 | -0.003 | -0.003 | -0.021 | 0.003 |  | -0.015 | -0.004 | -0.005 | -0.015 | -0.008 | -0.011 | -0.013 | -0.007
| 180 | -0.052 | -0.049 | -0.073 | -0.066 | -0.066 | -0.077 | -0.07 | -0.072 | -0.067 | -0.079 | -0.068 | -0.064 | -0.084 | -0.077 | -0.08 | -0.064 | -0.062 | -0.063 | -0.074 | -0.061 | -0.071 | -0.069 | -0.067 | -0.054 | -0.064 | -0.036 | -0.046 | -0.041 | -0.038 | -0.037 | -0.034 | -0.054 | -0.023 |  | 0.018 | 0.03 | 0.028 | 0.011 | 0.03 | 0.021 | 0.022 | 0.026
| 190 | -0.065 | -0.049 | -0.084 | -0.071 | -0.071 | -0.08 | -0.063 | -0.074 | -0.068 | -0.078 | -0.074 | -0.064 | -0.087 | -0.08 | -0.083 | -0.068 | -0.064 | -0.06 | -0.077 | -0.063 | -0.073 | -0.069 | -0.07 | -0.054 | -0.07 | -0.035 | -0.048 | -0.043 | -0.035 | -0.049 | -0.041 | -0.066 | -0.028 |  | 0.022 | 0.023 | 0.025 | 0.01 | 0.027 | 0.025 | 0.021 | 0.027
| 200 | -0.056 | -0.045 | -0.074 | -0.058 | -0.058 | -0.07 | -0.059 | -0.059 | -0.061 | -0.065 | -0.059 | -0.057 | -0.077 | -0.07 | -0.071 | -0.048 | -0.053 | -0.047 | -0.067 | -0.05 | -0.066 | -0.059 | -0.054 | -0.047 | -0.053 | -0.025 | -0.032 | -0.025 | -0.02 | -0.032 | -0.032 | -0.056 | -0.016 |  | 0.004 | 0.003 | 0.004 | -0.012 | 0.017 | 0.01 | 0.006 | 0.01
| 210 | -0.027 | -0.022 | -0.052 | -0.027 | -0.027 | -0.037 | -0.033 | -0.034 | -0.024 | -0.031 | -0.034 | -0.035 | -0.051 | -0.037 | -0.047 | -0.024 | -0.024 | -0.019 | -0.03 | -0.014 | -0.028 | -0.035 | -0.024 | -0.024 | -0.03 | 0.002 | -0.002 | 0.006 | 0.018 | 0 | 0.001 | -0.031 | 0.031 |  | -0.047 | -0.026 | -0.025 | -0.033 | -0.017 | -0.024 | -0.019 | -0.015
| 220 | 0.042 | 0.069 | 0.023 | 0.051 | 0.051 | 0.027 | 0.024 | 0.035 | 0.048 | 0.036 | 0.021 | 0.021 | 0.012 | 0.027 | 0.017 | 0.044 | 0.052 | 0.05 | 0.038 | 0.06 | 0.038 | 0.017 | 0.047 | 0.031 | 0.029 | 0.048 | 0.075 | 0.082 | 0.109 | 0.087 | 0.086 | 0.033 | 0.145 |  | -0.158 | -0.105 | -0.106 | -0.116 | -0.102 | -0.111 | -0.108 | -0.082
| 226 | 0.023 | 0.066 | 0.011 | 0.041 | 0.041 | 0.019 | 0.016 | 0.025 | 0.049 | 0.031 | 0.005 | -0.01 | 0.006 | 0.019 | 0.005 | 0.045 | 0.045 | 0.05 | 0.027 | 0.064 | 0.043 | -0.034 | 0.041 | 0.019 | 0.011 | 0.05 | 0.073 | 0.076 | 0.115 | 0.101 | 0.088 | 0.024 | 0.171 |  | -0.18 | -0.106 | -0.114 | -0.135 | -0.124 | -0.123 | -0.129 | -0.09
S parameter [%] sigma comparision
| E [MeV] | 2019-04-22 | 2019-07-01 | 2019-12-17 | 2020-01-14 | 2020-01-14 | 2020-02-06 | 2020-02-06 | 2020-04-28 | 2020-05-28 | 2020-06-16 | 2020-07-05 | 2020-07-27 | 2020-10-27 | 2020-11-09 | 2020-11-10 | 2020-12-08 | 2021-01-12 | 2021-02-09 | 2021-03-31 | 2021-04-20 | 2021-05-25 | 2021-07-01 | 2021-08-02 | 2021-10-19 | 2021-10-19 | 2021-11-15 | 2021-12-10 | 2022-01-03 | 2022-02-14 | 2022-03-14 | 2022-05-09 | 2022-07-08 | 2022-07-11 | 2022-07-15 | 2022-09-09 | 2022-10-08 | 2022-11-09 | 2022-12-15 | 2023-02-04 | 2023-03-21 | 2023-04-24 | 2023-05-27
| 60 | -0.9 | -1.5 | -0.7 | -0.9 | -0.9 | -1.3 | -1.3 | -1.3 | -0.8 | -0.8 | -1.4 | -1.186 | -1 | -1.294 | -0.5 | -0.7 | -1.2 | -0.8 | -0.815 | -0.722 | -1.147 | -0.987 | -0.796 | -0.815 | -0.907 | -1.977 | -1.335 | -1.705 | -1.318 | -1.695 | -2.288 | -1.607 | -1.728 | -1.73 | -2.027 | -2.373 | -1.764 | -1.42 | -1.161 | -1.494 | -0.847 | -1.528
| 70 | -1.5 | -1.7 | -1.4 | -1 | -1 | -1.2 | -1.1 | -1.2 | -0.5 | -1.3 | -1.5 | -1.269 | -1.3 | -1.218 | -1.1 | -0.8 | -1.2 | -1 | -1.498 | -1.276 | -0.853 | -1.093 | -0.901 | -1.697 | -1.815 | -2.029 | -1.461 | -1.647 | -1.922 | -2.15 | -2.51 | -1.894 | -2.2 | -1.4 | -2.125 | -2.311 | -2.235 | -1.951 | -1.706 | -1.799 | -1.937 | -2.277
| 80 | -0.8 | -1.2 | -0.4 | -0.7 | -0.7 | -0.9 | -0.7 | -0.7 | -0.5 | -0.8 | -1.2 | -1.125 | -0.7 | -0.863 | -0.6 | -1.1 | -1.2 | -0.7 | -1.167 | -0.717 | -0.896 | -0.719 | -0.363 | -0.895 | -1.253 | -1.67 | -1.604 | -1.124 | -1.899 | -1.843 | -1.412 | -1.712 | -1.111 | -1.69 | -1.92 | -1.859 | -1.608 | -1.831 | -1.391 | -1.81 | -1.637 | -1.983
| 90 | -0.9 | -1.6 | -1.1 | -1 | -1 | -1.4 | -0.9 | -1.2 | -1.2 | -1.3 | -1 | -0.919 | -1.1 | -1.433 | -1.1 | -1.1 | -1 | -0.9 | -1.226 | -0.929 | -1.446 | -0.954 | -0.829 | -1.238 | -1.64 | -2.1 | -1.583 | -2.17 | -2.188 | -2.148 | -2.254 | -1.961 | -1.739 | -1.95 | -2.182 | -2.258 | -2.135 | -2.195 | -1.882 | -2.037 | -2.32 | -2.113
| 100 | -1.3 | -1.9 | -1.4 | -1.4 | -1.4 | -1.6 | -1.2 | -1.6 | -1.3 | -1.6 | -1.8 | -1.183 | -1.3 | -1.605 | -1.4 | -1.7 | -1.7 | -1.7 | -1.545 | -1.542 | -1.478 | -1.518 | -1.188 | -1.704 | -1.834 | -2.319 | -1.969 | -2.479 | -2.68 | -2.625 | -2.581 | -2.719 | -2.674 | -2.95 | -2.867 | -2.448 | -2.357 | -2.354 | -2.491 | -2.369 | -2.592 | -2.319
| 110 | -1.6 | -1.9 | -1.3 | -1.5 | -1.5 | -1.5 | -1.4 | -1.6 | -1.7 | -1.4 | -1.6 | -1.271 | -1.4 | -1.512 | -1.4 | -1.8 | -1.6 | -1.7 | -1.706 | -1.662 | -1.658 | -1.532 | -1.408 | -1.998 | -1.918 | -2.679 | -2.465 | -2.713 | -2.777 | -2.842 | -2.866 | -2.737 | -2.878 | -2.9 | -2.751 | -2.79 | -2.573 | -2.98 | -2.736 | -2.776 | -2.747 | -2.783
| 120 | -1.5 | -2 | -1 | -1.4 | -1.4 | -1.8 | -1.3 | -1.5 | -1.6 | -1.5 | -1.5 | -1.273 | -1.6 | -1.79 | -1.5 | -2 | -1.8 | -1.9 | -1.647 | -1.801 | -1.582 | -1.7 | -1.572 | -1.864 | -2.09 | -2.827 | -2.694 | -2.516 | -3.224 | -3.082 | -3.187 | -2.876 | -2.808 | -2.98 | -3.109 | -2.851 | -2.862 | -2.733 | -2.752 | -3.012 | -3.109 | -3.011
| 130 | -1.8 | -2.2 | -1.5 | -1.4 | -1.4 | -1.9 | -1.5 | -1.5 | -1.4 | -1.7 | -1.6 | -1.505 | -1.4 | -1.853 | -1.6 | -1.8 | -1.7 | -1.9 | -1.851 | -1.769 | -1.598 | -1.462 | -1.108 | -2.054 | -2.38 | -2.893 | -2.601 | -2.975 | -3.27 | -3.652 | -3.317 | -3.259 | -2.939 | -3.05 | -3.011 | -3.079 | -2.719 | -3.042 | -3.063 | -3.211 | -3.366 | -3.417
| 140 | -1.3 | -1.7 | -0.8 | -1 | -1 | -1.5 | -1.5 | -1.3 | -1.5 | -1.5 | -1.4 | -1.256 | -1.1 | -1.48 | -1.4 | -1.8 | -1.4 | -1.7 | -1.457 | -1.154 | -0.705 | -1.141 | -0.957 | -1.743 | -1.97 | -2.625 | -2.255 | -2.61 | -2.945 | -3.161 | -2.849 | -2.465 | -2.111 | -2.14 | -2.428 | -2.723 | -2.211 | -2.392 | -2.773 | -2.618 | -3.014 | -2.701
| 150 | -1.5 | -1.9 | -1.5 | -1.6 | -1.6 | -2.1 | -1.7 | -1.6 | -1.6 | -1.8 | -1.8 | -1.627 | -1.6 | -2.075 | -1.7 | -1.8 | -1.7 | -1.7 | -1.826 | -1.689 | -1.538 | -1.561 | -1.396 | -2.426 | -2.25 | -3.08 | -2.717 | -3.058 | -3.384 | -3.386 | -3.474 | -2.951 | -2.719 | -2.25 | -2.525 | -2.976 | -3.072 | -3.18 | -3.117 | -3.02 | -3.627 | -3.159
| 160 | -1.7 | -1.3 | -1.2 | -1.6 | -1.6 | -2 | -2 | -1.7 | -1.9 | -1.8 | -2.2 | -1.495 | -1.5 | -2.045 | -1.8 | -2 | -2 | -1.8 | -1.645 | -1.728 | -1.789 | -1.415 | -1.494 | -2.351 | -2.624 | -3.134 | -2.784 | -3.303 | -3.519 | -3.477 | -3.113 | -2.997 | -2.458 | -2.91 | -3.173 | -2.794 | -2.868 | -3.087 | -3.218 | -2.898 | -3.271 | -3.262
| 170 | -2.6 | -2.2 | -1.9 | -2.1 | -2.1 | -2.9 | -2.8 | -2.6 | -2.4 | -2.4 | -2.5 | -2.103 | -2.5 | -2.89 | -2.5 | -2.6 | -2.6 | -2.6 | -2.696 | -2.726 | -2.381 | -2.035 | -2.036 | -3.128 | -3.072 | -3.955 | -3.833 | -4.136 | -4.328 | -4.351 | -4.1 | -4.173 | -3.199 | -3.95 | -3.958 | -3.551 | -3.822 | -3.723 | -4.095 | -3.894 | -4.122 | -3.757
| 180 | -1.2 | -1.6 | -0.6 | -0.9 | -0.9 | -2.1 | -1.5 | -1.7 | -1.6 | -1.5 | -1.7 | -1.086 | -1.5 | -2.104 | -1.7 | -1.6 | -1.7 | -1.5 | -1.58 | -1.451 | -1.702 | -1.135 | -1.421 | -2.333 | -2.24 | -2.999 | -2.656 | -2.746 | -3.112 | -3.141 | -3.313 | -3.304 | -2.264 | -2.65 | -2.549 | -2.536 | -2.625 | -2.284 | -2.937 | -3.031 | -3.273 | -2.983
| 190 | -0.7 | -0.9 | -0.4 | -0.5 | -0.5 | -1.6 | -1.8 | -1.2 | -1.3 | -1 | -1.6 | -0.463 | -1.2 | -1.56 | -1 | -1.5 | -1.5 | -1.3 | -1.33 | -1.385 | -0.914 | -0.921 | -0.783 | -1.714 | -1.974 | -2.38 | -2.695 | -2.381 | -3.051 | -2.646 | -3.199 | -3.394 | -2.002 | -2.34 | -2.288 | -2.157 | -2.099 | -1.986 | -2.564 | -2.651 | -3.185 | -2.531
| 200 | -1.1 | -0.5 | -0.1 | -0.6 | -0.6 | -1.9 | -1 | -1.3 | -1.6 | -1.3 | -1.6 | -0.817 | -1.4 | -1.947 | -1.6 | -1.4 | -1.6 | -1.5 | -1.164 | -1.266 | -1.333 | -0.791 | -0.735 | -2.069 | -1.898 | -2.98 | -2.865 | -2.792 | -3.427 | -3.457 | -3.457 | -3.875 | -2.17 | -2.27 | -2.064 | -2.259 | -2.426 | -2.108 | -3.049 | -3.115 | -3.405 | -2.809
| 210 | -1.8 | -1.5 | -0.4 | -1.2 | -1.2 | -2.8 | -1.8 | -2.5 | -2.5 | -2.2 | -2.5 | -2.072 | -2.5 | -2.819 | -2.5 | -3 | -2.8 | -2.5 | -2.341 | -2.47 | -1.947 | -1.971 | -1.716 | -3.297 | -3.365 | -4.17 | -4.5 | -4.031 | -4.585 | -4.545 | -4.598 | -5.271 | -2.543 | -3.75 | -2.831 | -3.835 | -3.8 | -4.046 | -4.298 | -4.441 | -4.858 | -3.947
| 220 | -1.9 | -1.2 | -0.4 | -1.5 | -1.5 | -4.2 | -3.2 | -3.3 | -3.7 | -3.5 | -3.9 | -3.613 | -3.6 | -4.202 | -3.4 | -4.1 | -4.1 | -3.9 | -3.758 | -3.56 | -3.35 | -3.394 | -3.409 | -5.223 | -5.212 | -6.28 | -6.255 | -5.537 | -6.668 | -6.732 | -6.768 | -7.316 | -2.038 | -2.21 | -2.828 | -5.454 | -6.31 | -5.832 | -6.423 | -6.252 | -6.483 | -6.062
| 226 | -2.8 | -2.2 | -1.1 | -1.9 | -1.9 | -4.3 | -3.3 | -3.6 | -3.5 | -3.8 | -4 | -4.448 | -3.7 | -4.305 | -3.8 | -4.1 | -4.5 | -4.2 | -4.156 | -3.389 | -3.857 | -1.954 | -3.645 | -5.507 | -5.93 | -6.641 | -6.47 | -6.333 | -7.378 | -7.084 | -7.476 | -7.591 | -1.683 | -2.39 | -2.436 | -6.608 | -6.626 | -5.846 | -6.466 | -6.605 | -7.161 | -6.494