## GTR1
Daily cubes parameters | [13 empty] | Reference | from measurements for Matrixx 2 | [6 empty] | first half of 2019 |  |  |  |  | first half of 2019
MatriXX PT | [14 empty] | R15_SOBP10_2Gy |  | R31_SOBP10 |  |  |  |  | R15_SOBP10_2Gy |  | R31_SOBP10 |  |  | R15_SOBP10_2Gy |  | R31_SOBP10
[11 empty] | S | F |  | D (from meas) | 2.003 |  | 2 |  |  | P |  | 9.0208 |  | 11.6688 |  | FWHM | 119.7792 |  | 119.5792
Set-up: MatriXX position adjusted using lasers or X-rays (chambers at iso); MyQA Project: GTR2_MAtrixMorgonQA | [10 empty] | 0.5 | 1 |  | 1% | 0.02003 | 2.023 | 0.02 | 2.02 |  | 1 |  | 10.0208 |  | 12.6688 |  | 1 | 120.7792 |  | 120.5792
| F | S | Dose | FWHM [mm] | Penumbra [mm] | [5 empty] | -0.5 | -1 |  | -1% | -0.02003 | 1.983 | -0.02 | 1.98 |  | -1 |  | 8.0208 |  | 10.6688 |  | -1 | 118.7792 |  | 118.5792
warrning | 1% | 0.5% | 1% | 1 | 1 | [5 empty] | 1 | 2 |  | 2% | 0.04006 | 2.0431 | 0.04 | 2.04 |  | 2 |  | 11.0208 |  | 13.6688 |  | 2 | 121.7792 |  | 121.5792
action level | 2% | 1% | 2% | 2 | 2 | [5 empty] | -1 | -2 |  | -2% | -0.04006 | 1.9629 | -0.04 | 1.96 |  | -2 |  | 7.0208 |  | 9.6688 |  | -2 | 117.7792 |  | 117.5792
Date | R15_SOBP10_2Gy, 10 cm WET, 9.2 SW | [13 empty] | R31_SOBP10_2Gy , 26 cm WET, 24.8 SW | [13 empty] | Notes | From Daily QA
| Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) |  |  |  | ref R15_SOBP10_2Gy | Diff R15_SOBP10_2Gy | ref R31_SOBP10_2Gy | Diff R31_SOBP10_2Gy | 1% | -1% | 2% | -2% | Ref.
20190206 | 0.38 | 0.23 | 0.18 | 0.01 | 0.52 | 0.1 | 0.07 | 2 | 9.5 | 8.8 | 9.1 | 9 | 119.6 | 119.8 | 0.32 | 0.25 | 0.14 | -0.17 | 0.54 | 0.07 | -0.17 | 1.999 | 11.6 | 11.6 | 11.9 | 12.2 | 119.4 | 119.5 | LS |  | 2019-02-06 | 2.003 | 0.9985 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190220 | 0.28 | 0.25 | 0.13 | 0.01 | 0.51 | 0.13 | 0.1 | 1.998 | 9.2 | 9 | 9.2 | 8.9 | 119.6 | 119.8 | 0.29 | 0.25 | 0.1 | -0.17 | 0.51 | 0.06 | -0.16 | 1.995 | 11.6 | 11.6 | 11.6 | 11.6 | 119.4 | 119.6 | LS |  | 2019-02-20 | 2.003 | 0.9975 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190305 | 0.31 | 0.18 | 0.1 | 0.01 | 0.54 | 0.08 | 0.07 | 2.014 | 9.5 | 8.8 | 9.1 | 8.9 | 119.6 | 119.8 | 0.23 | 0.26 | 0.06 | -0.13 | 0.52 | 0.06 | -0.11 | 2.008 | 11.6 | 11.6 | 11.7 | 11.7 | 119.3 | 119.6 | LS |  | 2019-03-05 | 2.003 | 1.0055 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190312 | 0.42 | 0.23 | 0.17 | 0 | 0.48 | 0.07 | 0.09 | 1.994 | 9.4 | 8.9 | 9.2 | 8.9 | 119.7 | 119.9 | 0.26 | 0.19 | 0.09 | -0.11 | 0.44 | 0.06 | -0.1 | 2.003 | 11.9 | 11.7 | 11.7 | 11.9 | 119.6 | 119.7 | LS |  | 2019-03-12 | 2.003 | 0.99551 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190317 | 0.29 | 0.18 | 0.11 | 0.03 | 0.6 | 0.07 | 0.03 | 2.007 | 8.9 | 9.3 | 9.6 | 8.2 | 119.6 | 119.8 | 0.24 | 0.31 | 0.11 | -0.18 | 0.58 | 0.1 | -0.13 | 2.006 | 11.8 | 11.7 | 11.8 | 12 | 119.3 | 119.6 | LS |  | 2019-03-17 | 2.003 | 1.002 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190320 | 0.41 | 0.49 | 0.17 | -0.34 | 0.69 | 0.1 | -0.26 | 2.006 | 9 | 9.1 | 8.8 | 8.9 | 119.8 | 120 | 0.4 | 0.69 | 0.13 | -0.56 | 0.83 | 0.09 | -0.48 | 1.995 | 11.5 | 11.7 | 11.5 | 11.8 | 119.6 | 119.8 | LS |  | 2019-03-20 | 2.003 | 1.0015 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190326 | 0.28 | 0.23 | 0.1 | 0.12 | 0.58 | 0.04 | 0.13 | 2 | 9 | 9.2 | 9.2 | 9 | 119.7 | 119.9 | 0.3 | 0.26 | 0.14 | -0.16 | 0.5 | 0.09 | -0.14 | 1.997 | 11.7 | 11.6 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2019-03-26 | 2.003 | 0.9985 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190422 | 0.33 | 0.29 | 0.18 | 0.05 | 0.57 | 0.09 | 0.05 | 2.001 | 9.2 | 8.9 | 9 | 8.6 | 119.7 | 120 | 0.28 | 0.29 | 0.13 | -0.21 | 0.55 | 0.09 | -0.18 | 1.993 | 11.5 | 11.5 | 11.6 | 12 | 119.6 | 119.9 | LS |  | 2019-04-22 | 2.003 | 0.999 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190507 | 0.35 | 0.14 | 0.15 | 0.04 | 0.49 | 0.08 | 0.07 | 2.003 | 9.2 | 8.9 | 8.8 | 8.8 | 119.6 | 119.9 | 0.39 | 0.28 | 0.18 | -0.2 | 0.64 | 0.12 | -0.18 | 1.994 | 11.3 | 11.3 | 11.5 | 11.4 | 119.5 | 119.8 | LS |  | 2019-05-07 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190514 | 0.39 | 0.18 | 0.21 | -0.02 | 0.5 | 0.13 | 0.08 | 2.003 | 9.1 | 9 | 9.2 | 8.9 | 119.8 | 119.9 | 0.29 | 0.26 | 0.12 | -0.19 | 0.54 | 0.09 | -0.13 | 2.012 | 11.6 | 11.6 | 11.7 | 11.9 | 119.5 | 119.7 | LS |  | 2019-05-14 | 2.003 | 1 | 2 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190523 | 0.33 | 0.19 | 0.17 | 0.04 | 0.51 | 0.1 | 0.09 | 2.02 | 9.3 | 8.9 | 9 | 8.9 | 119.7 | 119.8 | 0.29 | 0.17 | 0.14 | -0.12 | 0.47 | 0.1 | -0.11 | 2.015 | 11.6 | 11.6 | 11.8 | 12 | 119.5 | 119.5 | LS |  | 2019-05-23 | 2.003 | 1.0085 | 2 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190619 | 0.25 | 0.25 | 0.16 | 0.06 | 0.49 | 0.12 | 0.07 | 1.999 | 8.8 | 9.3 | 8.9 | 8.7 | 119.7 | 120 | 0.22 | 0.22 | 0.1 | -0.16 | 0.46 | 0.08 | -0.11 | 1.995 | 11.5 | 11.7 | 11.5 | 11.4 | 119.5 | 119.8 | LS |  | 2019-06-19 | 2.003 | 0.998 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190702 | 0.48 | 0.23 | 0.22 | -0.02 | 0.5 | 0.14 | 0.01 | 1.994 | 9.2 | 8.9 | 8.9 | 8.6 | 119.7 | 120 | 0.38 | 0.31 | 0.2 | -0.26 | 0.7 | 0.15 | -0.24 | 1.995 | 11.4 | 11.4 | 11.5 | 11.7 | 119.5 | 119.8 | LS |  | 2019-07-02 | 2.003 | 0.99551 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190709 | 0.27 | 0.45 | 0.12 | 0.08 | 0.68 | 0.09 | 0.12 | 1.998 | 9.1 | 8.9 | 8.5 | 9 | 119.7 | 120 | 0.34 | 0.29 | 0.1 | -0.23 | 0.64 | 0.1 | -0.2 | 1.999 | 11.5 | 11.6 | 11.6 | 11.5 | 119.5 | 119.9 | LS |  | 2019-07-09 | 2.003 | 0.9975 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190723 | 0.42 | 0.29 | 0.18 | 0.03 | 0.54 | 0.12 | 0.06 | 2.001 | 9.5 | 8.6 | 9.1 | 8.6 | 119.6 | 119.9 | 0.32 | 0.4 | 0.11 | -0.24 | 0.68 | 0.11 | -0.24 | 1.997 | 11.5 | 11.5 | 11.7 | 11.9 | 119.4 | 119.7 | LS |  | 2019-07-23 | 2.003 | 0.999 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190801 | 0.39 | 0.19 | 0.18 | 0.01 | 0.56 | 0.12 | 0.04 | 2.004 | 9.5 | 8.6 | 9.1 | 8.6 | 119.6 | 120 | 0.32 | 0.29 | 0.12 | -0.2 | 0.55 | 0.08 | -0.17 | 2.003 | 11.6 | 11.8 | 11.6 | 11.9 | 119.5 | 119.8 | LS |  | 2019-08-01 | 2.003 | 1.0005 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190807 | 0.33 | 0.49 | -0.05 | 0.03 | 0.68 | -0.06 | 0.03 | 2.019 | 9.6 | 8.4 | 9.5 | 8.1 | 119.5 | 119.8 | [14 empty] | LS |  | 2019-08-07 | 2.003 | 1.008 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190814 | 0.36 | 0.14 | 0.1 | -0.03 | 0.44 | 0.09 | 0.04 | 2.003 | 9.6 | 8.3 | 8.8 | 8.8 | 119.6 | 120 | 0.34 | 0.25 | 0.11 | -0.16 | 0.55 | 0.1 | -0.1 | 2.006 | 11.8 | 12 | 11.7 | 11.9 | 119.4 | 119.6 | LS |  | 2019-08-14 | 2.003 | 1 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190820 | 0.41 | 0.16 | 0.19 | 0.04 | 0.62 | 0.13 | 0.14 | 2.004 | 9.6 | 8.3 | 9.1 | 8.4 | 119.6 | 120 | 0.39 | 0.32 | 0.19 | -0.18 | 0.58 | 0.14 | -0.13 | 2.001 | 11.6 | 11.8 | 11.7 | 12 | 119.4 | 119.7 | LS |  | 2019-08-20 | 2.003 | 1.0005 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190828 | 0.42 | 0.09 | 0.2 | -0.01 | 0.52 | 0.15 | 0.09 | 2.004 | 9.7 | 8.3 | 9.3 | 8.2 | 119.6 | 119.9 | 0.42 | 0.4 | 0.17 | -0.22 | 0.64 | 0.1 | -0.21 | 2 | 11.7 | 11.9 | 11.8 | 12 | 119.3 | 119.6 | LS |  | 2019-08-28 | 2.003 | 1.0005 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190911 | 0.49 | 0.16 | 0.24 | -0.02 | 0.57 | 0.13 | 0.08 | 2 | 9.5 | 8.5 | 9.3 | 8.3 | 119.6 | 119.9 | 0.39 | 0.29 | 0.18 | -0.2 | 0.66 | 0.14 | -0.19 | 2.003 | 11.6 | 11.7 | 11.7 | 11.9 | 119.3 | 119.6 | LS |  | 2019-09-11 | 2.003 | 0.9985 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190918 | 0.45 | 0.21 | 0.26 | -0.01 | 0.67 | 0.17 | 0.1 | 2.002 | 9.7 | 8.2 | 9.1 | 8.8 | 119.6 | 119.9 | 0.42 | 0.26 | 0.19 | -0.14 | 0.73 | 0.15 | -0.14 | 2.004 | 11.5 | 11.4 | 11.8 | 12.1 | 119.4 | 119.7 | LS |  | 2019-09-18 | 2.003 | 0.9995 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191009 | 0.32 | 0.2 | 0.2 | 0.01 | 0.53 | 0.15 | 0.07 | 2.005 | 9 | 9.2 | 8.9 | 8.9 | 119.6 | 119.9 | 0.29 | 0.28 | 0.13 | -0.17 | 0.55 | 0.1 | -0.13 | 2.002 | 11.4 | 11.5 | 11.5 | 11.4 | 119.4 | 119.7 | LS |  | 2019-10-09 | 2.003 | 1.001 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191013 | 0.52 | 0.23 | 0.32 | -0.02 | 0.67 | 0.24 | 0.02 | 1.997 | 9.1 | 9.1 | 9 | 8.7 | 119.6 | 119.9 | 0.44 | 0.31 | 0.23 | -0.22 | 0.82 | 0.17 | -0.23 | 1.997 | 11.4 | 11.5 | 11.6 | 11.7 | 119.4 | 119.6 | LS |  | 2019-10-13 | 2.003 | 0.997 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191016 | 0.47 | 0.22 | 0.22 | 0 | 0.56 | 0.12 | 0.02 | 1.993 | 9.8 | 8 | 9.5 | 7.8 | 119.6 | 119.9 | 0.42 | 0.24 | 0.17 | -0.18 | 0.6 | 0.13 | -0.15 | 1.986 | 11.7 | 11.9 | 11.8 | 12.1 | 119.4 | 119.5 | LS |  | 2019-10-16 | 2.003 | 0.99501 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191030 | 0.45 | 0.14 | 0.28 | -0.03 | 0.6 | 0.2 | 0.03 | 2.005 | 9.5 | 8.5 | 9.1 | 8.9 | 119.6 | 119.7 | 0.39 | 0.32 | 0.2 | -0.24 | 0.71 | 0.14 | -0.23 | 2.002 | 11.5 | 11.6 | 11.7 | 11.7 | 119.4 | 119.4 | LS |  | 2019-10-30 | 2.003 | 1.001 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191107 | 0.36 | 0.17 | 0.18 | -0.04 | 0.53 | 0.12 | 0.04 | 2 | 9.7 | 8.2 | 9.2 | 8.7 | 119.6 | 119.8 | 0.28 | 0.21 | 0.13 | -0.14 | 0.51 | 0.1 | -0.15 | 1.997 | 11.6 | 11.8 | 11.7 | 11.9 | 119.4 | 119.6 | LS |  | 2019-11-07 | 2.003 | 0.9985 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191114 | 0.32 | 0.29 | 0.17 | 0.04 | 0.53 | 0.09 | 0.02 | 1.999 | 9.3 | 9 | 8.9 | 9 | 119.6 | 119.8 | 0.26 | 0.25 | 0.12 | -0.07 | 0.47 | 0.09 | -0.07 | 1.997 | 11.4 | 11.4 | 11.5 | 11.4 | 119.5 | 119.6 | LS |  | 2019-11-14 | 2.003 | 0.998 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191119 | 0.36 | 0.09 | 0.15 | -0.01 | 0.53 | 0.12 | 0.01 | 2.007 | 9.2 | 9.1 | 9 | 8.8 | 119.6 | 119.9 | 0.35 | 0.37 | 0.14 | -0.22 | 0.7 | 0.14 | -0.21 | 2.006 | 11.4 | 11.5 | 11.5 | 11.4 | 119.4 | 119.6 | LS |  | 2019-11-19 | 2.003 | 1.002 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191122 | 0.29 | 0.24 | 0.15 | 0.03 | 0.5 | 0.07 | 0.05 | 2.004 | 9.4 | 8.7 | 9.5 | 8.3 | 119.6 | 119.8 | 0.19 | 0.42 | 0.05 | -0.22 | 0.55 | 0.02 | -0.18 | 2.006 | 11.7 | 11.9 | 11.4 | 11.4 | 119.4 | 119.5 | LS |  | 2019-11-22 | 2.003 | 1.0005 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191128 | 0.37 | 0.1 | 0.19 | -0.01 | 0.53 | 0.13 | 0.06 | 1.993 | 9.4 | 8.6 | 8.9 | 8.8 | 119.7 | 120 | 0.38 | 0.29 | 0.16 | -0.21 | 0.65 | 0.12 | -0.18 | 1.99 | 11.4 | 11.4 | 11.4 | 11.4 | 119.5 | 119.7 | LS |  | 2019-11-28 | 2.003 | 0.99501 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191204 | 0.47 | 0.19 | 0.23 | 0.02 | 0.63 | 0.16 | 0.02 | 2.002 | 9.3 | 8.8 | 9.3 | 8.4 | 119.6 | 119.9 | 0.39 | 0.27 | 0.18 | -0.16 | 0.59 | 0.13 | -0.15 | 2 | 11.5 | 11.4 | 11.7 | 11.9 | 119.5 | 119.6 | LS |  | 2019-12-04 | 2.003 | 0.9995 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191211 | 0.44 | 0.14 | 0.23 | -0.06 | 0.59 | 0.16 | 0.02 | 1.987 | 9.6 | 8.4 | 9.4 | 8.2 | 119.6 | 119.9 | 0.29 | 0.24 | 0.11 | -0.14 | 0.52 | 0.07 | -0.14 | 1.987 | 11.6 | 11.7 | 11.7 | 11.9 | 119.4 | 119.6 | LS |  | 2019-12-11 | 2.003 | 0.99201 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191219 | 0.37 | 0.19 | 0.15 | 0.04 | 0.58 | 0.11 | 0.08 | 2.004 | 9.8 | 7.9 | 9.1 | 8.6 | 119.6 | 119.9 | 0.31 | 0.27 | 0.14 | -0.17 | 0.55 | 0.11 | -0.15 | 2.003 | 11.6 | 11.8 | 11.7 | 11.9 | 119.5 | 119.7 | LS |  | 2019-12-19 | 2.003 | 1.0005 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191228 | 0.35 | 0.16 | 0.2 | -0.01 | 0.58 | 0.12 | 0.02 | 2.008 | 9.3 | 8.9 | 9.4 | 7.9 | 119.7 | 119.9 | 0.3 | 0.27 | 0.11 | -0.12 | 0.55 | 0.08 | -0.12 | 2.009 | 11.4 | 11.4 | 11.8 | 12 | 119.5 | 119.6 | LS |  | 2019-12-28 | 2.003 | 1.0025 | 2 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191231 | 0.35 | 0.16 | 0.2 | -0.01 | 0.58 | 0.12 | 0.02 | 2.008 | 9.3 | 8.9 | 9.4 | 7.9 | 119.7 | 119.9 | 0.3 | 0.27 | 0.11 | -0.12 | 0.55 | 0.08 | -0.12 | 2.009 | 11.4 | 11.4 | 11.8 | 12 | 119.5 | 119.6 | LS |  | 2019-12-31 | 2.003 | 1.0025 | 2 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200107 | 0.32 | 0.17 | 0.13 | 0.03 | 0.57 | 0.1 | 0.07 | 2.004 | 10.1 | 7.2 | 9.1 | 8.7 | 119.3 | 119.9 | 0.33 | 0.28 | 0.12 | -0.17 | 0.56 | 0.1 | -0.14 | 1.998 | 11.8 | 12.1 | 11.7 | 12 | 119.4 | 119.7 | LS |  | 2020-01-07 | 2.003 | 1.0005 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200107 | 0.31 | 0.34 | 0.03 | 0.23 | 0.63 | 0.04 | 0.21 | 1.999 | 9.2 | 8.8 | 9 | 8.6 | 119.7 | 119.8 | 0.21 | 0.32 | 0.04 | -0.19 | 0.58 | 0.05 | -0.15 | 1.991 | 11.8 | 12.1 | 11.7 | 12 | 119.3 | 119.6 | ML |  | 2020-01-07 | 2.003 | 0.998 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200110 | 0.36 | 0.19 | 0.15 | -0.01 | 0.57 | 0.1 | 0.03 | 1.995 | 10.1 | 7.2 | 9.1 | 8.7 | 119.4 | 119.9 | 0.21 | 0.32 | 0.04 | -0.19 | 0.58 | 0.05 | -0.15 | 1.991 | 11.8 | 12.1 | 11.7 | 12 | 119.3 | 119.6 | LS |  | 2020-01-10 | 2.003 | 0.99601 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200124 | 0.39 | 1 | 0.21 | 0.07 | 1.17 | 0.11 | 0.08 | 2.009 | 8.9 | 9.5 | 8.4 | 9.5 | 119.6 | 119.7 | 0.5 | 0.7 | 0.26 | -0.22 | 0.76 | 0.08 | -0.19 | 2.001 | 11.2 | 11.4 | 11.3 | 11.2 | 119.4 | 119.6 | LS |  | 2020-01-24 | 2.003 | 1.003 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200130 | 0.82 | 1.05 | 0.44 | -0.06 | 1.11 | 0.14 | -0.01 | 2.011 | 9.2 | 9.2 | 8.8 | 9.2 | 119.6 | 119.8 | 0.47 | 0.59 | 0.24 | -0.2 | 0.72 | 0.13 | -0.19 | 2.002 | 11.4 | 11.5 | 11.6 | 11.6 | 119.4 | 119.6 | LS |  | 2020-01-30 | 2.003 | 1.004 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200206 | 0.77 | 0.95 | 0.41 | 0.01 | 1.01 | 0.12 | 0.06 | 2.005 | 8.8 | 9.4 | 8.9 | 8.9 | 119.6 | 119.8 | 0.45 | 0.53 | 0.25 | -0.17 | 0.66 | 0.11 | -0.16 | 1.998 | 11.3 | 11.4 | 11.5 | 11.5 | 119.4 | 119.5 | LS |  | 2020-02-06 | 2.003 | 1.001 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200213 | 0.75 | 1.01 | 0.4 | 0.01 | 1.01 | 0.07 | 0.04 | 2.003 | 9.3 | 8.9 | 8.9 | 8.9 | 119.6 | 119.8 | 0.39 | 0.67 | 0.16 | -0.13 | 0.72 | 0.06 | -0.12 | 1.995 | 11.6 | 11.6 | 11.7 | 11.8 | 119.4 | 119.5 | LS |  | 2020-02-13 | 2.003 | 1 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200228 | 0.68 | 1.09 | 0.34 | -0.03 | 1.09 | 0.06 | 0.02 | 2.01 | 9.2 | 9.2 | 9.2 | 8.9 | 119.5 | 119.7 | 0.48 | 0.61 | 0.22 | -0.17 | 0.74 | 0.07 | -0.17 | 2 | 11.5 | 11.4 | 11.6 | 11.6 | 119.3 | 119.5 | LS |  | 2020-02-28 | 2.003 | 1.0035 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200303 | 0.8 | 1.11 | 0.4 | -0.01 | 1.11 | 0.1 | 0.04 | 2.009 | 9.6 | 8.7 | 8.6 | 9.3 | 119.6 | 119.8 | 0.41 | 0.58 | 0.19 | -0.13 | 0.62 | 0.08 | -0.11 | 1.998 | 11.6 | 11.6 | 11.6 | 11.3 | 119.4 | 119.7 | LS |  | 2020-03-03 | 2.003 | 1.003 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200306 | 0.24 | 0.26 | 0.13 | 0.05 | 0.48 | 0.09 | 0.09 | 2.001 | 9 | 9 | 8.8 | 8.9 | 119.7 | 119.9 | 0.26 | 0.32 | 0.07 | -0.18 | 0.6 | 0.1 | -0.15 | 1.992 | 11.3 | 11.4 | 11.4 | 11.3 | 119.5 | 119.5 | LS |  | 2020-03-06 | 2.003 | 0.999 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200312 | 0.18 | 0.28 | 0.09 | 0.05 | 0.45 | 0.06 | 0.03 | 1.999 | 9 | 9 | 9.2 | 8.3 | 119.7 | 119.9 | 0.28 | 0.21 | 0.11 | -0.07 | 0.44 | 0.07 | -0.07 | 1.996 | 11.4 | 11.5 | 11.6 | 11.7 | 119.5 | 119.6 | LS |  | 2020-03-12 | 2.003 | 0.998 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200319 | 0.38 | 0.18 | 0.24 | -0.01 | 0.56 | 0.15 | 0.03 | 2.004 | 8.9 | 9.2 | 9.2 | 8.3 | 119.7 | 119.9 | 0.33 | 0.26 | 0.12 | -0.14 | 0.52 | 0.12 | -0.11 | 2.002 | 11.4 | 11.4 | 11.7 | 11.8 | 119.4 | 119.4 | LS |  | 2020-03-19 | 2.003 | 1.0005 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200325 | 0.31 | 0.21 | 0.18 | 0 | 0.51 | 0.1 | 0.03 | 2.003 | 9.3 | 8.7 | 9.1 | 8.5 | 119.7 | 119.9 | 0.32 | 0.25 | 0.15 | -0.16 | 0.47 | 0.11 | -0.11 | 2 | 11.5 | 11.5 | 11.7 | 11.8 | 119.4 | 119.6 | LS |  | 2020-03-25 | 2.003 | 1 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200402 | 0.24 | 0.25 | 0.13 | 0.06 | 0.45 | 0.08 | 0.06 | 2.002 | 9.1 | 9 | 9 | 8.8 | 119.6 | 119.8 | 0.31 | 0.21 | 0.13 | -0.11 | 0.5 | 0.1 | -0.13 | 1.999 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.6 | LS |  | 2020-04-02 | 2.003 | 0.9995 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200409 | 0.25 | 0.23 | 0.13 | -0.03 | 0.46 | 0.08 | 0 | 1.997 | 9.2 | 8.9 | 9.2 | 8.5 | 119.6 | 119.8 | 0.32 | 0.26 | 0.13 | -0.14 | 0.56 | 0.1 | -0.13 | 1.987 | 11.5 | 11.5 | 11.6 | 11.8 | 119.4 | 119.5 | LS |  | 2020-04-09 | 2.003 | 0.997 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200416 | 0.25 | 0.16 | 0.1 | -0.06 | 0.44 | 0.08 | 0 | 1.996 | 9.4 | 8.6 | 9.5 | 7.8 | 119.6 | 119.9 | 0.21 | 0.23 | 0.03 | -0.11 | 0.49 | 0.04 | -0.1 | 1.992 | 11.6 | 11.7 | 11.7 | 11.8 | 119.4 | 119.4 | LS |  | 2020-04-16 | 2.003 | 0.99651 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200423 | 0.22 | 0.22 | 0.1 | 0.03 | 0.45 | 0.08 | 0.05 | 1.997 | 9.4 | 8.7 | 9 | 8.8 | 119.6 | 119.8 | 0.25 | 0.18 | 0.07 | -0.03 | 0.48 | 0.05 | -0.09 | 1.992 | 11.7 | 11.8 | 11.7 | 11.9 | 119.3 | 119.6 | LS |  | 2020-04-23 | 2.003 | 0.997 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200507 | 0.26 | 0.19 | 0.09 | -0.01 | 0.53 | 0.09 | 0.05 | 2.004 | 9.3 | 8.9 | 9.1 | 8.7 | 119.6 | 119.8 | 0.27 | 0.28 | 0.05 | -0.14 | 0.54 | 0.04 | -0.14 | 1.994 | 11.5 | 11.4 | 11.7 | 11.8 | 119.5 | 119.4 | LS |  | 2020-05-07 | 2.003 | 1.0005 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200515 | 0.22 | 0.22 | 0.09 | 0.01 | 0.37 | 0.04 | 0.02 | 1.996 | 9.1 | 9.1 | 9.3 | 8.3 | 119.6 | 119.8 | 0.27 | 0.28 | 0.07 | -0.19 | 0.52 | 0.08 | -0.12 | 1.992 | 11.5 | 11.4 | 11.7 | 11.8 | 119.3 | 119.5 | AA |  | 2020-05-15 | 2.003 | 0.99651 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200528 | 0.26 | 0.17 | 0.11 | 0.02 | 0.44 | 0.06 | 0.04 | 2.002 | 9.2 | 9 | 9.3 | 8.2 | 119.6 | 119.9 | 0.28 | 0.28 | 0.07 | -0.15 | 0.53 | 0.07 | -0.12 | 1.996 | 11.5 | 11.4 | 11.7 | 11.9 | 119.4 | 119.7 | LS |  | 2020-05-28 | 2.003 | 0.9995 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200603 | 0.2 | 0.26 | 0.07 | 0.06 | 0.42 | 0.06 | 0.03 | 2.003 | 9.2 | 8.7 | 9 | 8.5 | 119.7 | 120 | 0.22 | 0.24 | 0.05 | -0.1 | 0.53 | 0.06 | -0.12 | 2.002 | 11.4 | 11.2 | 11.6 | 11.7 | 119.5 | 119.6 | LS |  | 2020-06-03 | 2.003 | 1 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200610 | 0.24 | 0.28 | 0.07 | 0.06 | 0.42 | 0.05 | 0.05 | 2.003 | 9.4 | 8.7 | 9.2 | 8.4 | 119.6 | 119.9 | 0.21 | 0.33 | 0.08 | -0.15 | 0.58 | 0.05 | -0.16 | 1.993 | 11.5 | 11.4 | 11.7 | 11.8 | 119.4 | 119.7 | AA |  | 2020-06-10 | 2.003 | 1 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200624 | 0.3 | 0.19 | 0.17 | 0.02 | 0.43 | 0.1 | 0.03 | 2 | 9.5 | 8.4 | 9.3 | 8.1 | 119.6 | 119.9 | 0.28 | 0.28 | 0.09 | -0.16 | 0.54 | 0.06 | -0.14 | 1.995 | 11.6 | 11.7 | 11.7 | 11.9 | 119.4 | 119.8 | LS after downtime (1) |  | 2020-06-24 | 2.003 | 0.9985 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200624 | 0.19 | 0.22 | 0.1 | 0.09 | 0.36 | 0.06 | 0.08 | 2.011 | 9.4 | 8.6 | 9.1 | 8.5 | 119.7 | 120 | 0.23 | 0.21 | 0.08 | -0.11 | 0.43 | 0.03 | -0.1 | 2.014 | 11.8 | 12 | 11.8 | 12.2 | 119.5 | 119.8 | LS after downtime (repetition 1) |  | 2020-06-24 | 2.003 | 1.004 | 2 | 1.007 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200624 | [14 empty] | 0.23 | 0.28 | 0.09 | -0.19 | 0.51 | 0.07 | -0.15 | 2.013 | 11.7 | 12 | 11.8 | 12.1 | 119.5 | 119.8 | LS after downtime (repetition 2) |  | 2020-06-24 | 2.003 | 0 | 2 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200624 | [14 empty] | 0.23 | 0.28 | 0.09 | -0.19 | 0.51 | 0.07 | -0.15 | 2.013 | 11.7 | 12 | 11.8 | 12.1 | 119.5 | 119.8 | LS after downtime (repetition 3, updated DM) |  | 2020-06-24 | 2.003 | 0 | 2 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200624 | [14 empty] | 0.24 | 0.22 | 0.1 | -0.15 | 0.48 | 0.09 | -0.13 | 2.019 | 11.7 | 11.9 | 11.8 | 12.2 | 119.5 | 119.7 | CVD repetition in the evening |  | 2020-06-24 | 2.003 | 0 | 2 | 1.0095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200624 | [14 empty] | 0.22 | 0.27 | 0.04 | -0.14 | 0.47 | 0.05 | -0.1 | 2 | 11.5 | 11.4 | 11.6 | 11.6 | 119.5 | 119.6 | LS |  | 2020-06-24 | 2.003 | 0 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200702 | 0.24 | 0.18 | 0.12 | 0 | 0.45 | 0.11 | 0.04 | 1.996 | 9.6 | 8.5 | 9.1 | 8.6 | 119.5 | 119.9 | 0.28 | 0.23 | 0.09 | -0.17 | 0.65 | 0.08 | -0.2 | 1.992 | 11.6 | 11.6 | 11.6 | 11.8 | 119.4 | 119.7 | AA |  | 2020-07-02 | 2.003 | 0.99651 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200705 | 0.39 | 0.25 | 0.19 | 0.02 | 0.5 | 0.09 | 0.05 | 2 | 8.1 | 9.6 | 8.8 | 8.6 | 119.6 | 120 | 0.28 | 0.28 | 0.15 | -0.16 | 0.55 | 0.13 | -0.12 | 2.021 | 12.1 | 11.8 | 11.7 | 11.9 | 119.3 | 119.8 | LS |  | 2020-07-05 | 2.003 | 0.9985 | 2 | 1.0105 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200707 | 0.25 | 0.3 | 0.13 | 0.07 | 0.42 | 0.07 | 0.08 | 2.003 | 8.8 | 9.3 | 9 | 8.5 | 119.6 | 119.9 | 0.27 | 0.26 | 0.11 | -0.15 | 0.54 | 0.1 | -0.11 | 1.997 | 11.3 | 11.4 | 11.5 | 11.4 | 119.5 | 119.7 | LS |  | 2020-07-07 | 2.003 | 1 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200708 | 0.19 | 0.26 | 0.1 | 0 | 0.37 | 0.08 | 0.05 | 2 | 9.3 | 8.7 | 9.2 | 8.2 | 119.6 | 119.9 | 0.21 | 0.29 | 0.08 | -0.14 | 0.44 | 0.08 | -0.08 | 1.995 | 11.4 | 11.2 | 11.6 | 11.8 | 119.4 | 119.7 | AA |  | 2020-07-08 | 2.003 | 0.9985 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200715 | 0.29 | 0.29 | 0.14 | 0.08 | 0.46 | 0.08 | 0.1 | 1.987 | 9.2 | 8.9 | 9 | 8.4 | 119.7 | 120 | 0.34 | 0.2 | 0.13 | -0.11 | 0.56 | 0.11 | -0.11 | 1.988 | 11.3 | 11.3 | 11.5 | 11.4 | 119.4 | 119.6 | AA |  | 2020-07-15 | 2.003 | 0.99201 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200727 | 0.26 | 0.31 | 0.12 | 0.07 | 0.5 | 0.09 | 0.05 | 1.999 | 8.9 | 9.1 | 9.1 | 8.2 | 119.7 | 120 | 0.32 | 0.18 | 0.15 | -0.08 | 0.48 | 0.11 | -0.08 | 1.994 | 11.4 | 11.4 | 11.7 | 11.7 | 119.5 | 119.5 | AA |  | 2020-07-27 | 2.003 | 0.998 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200805 | 0.36 | 0.14 | 0.21 | -0.02 | 0.5 | 0.14 | 0.04 | 1.992 | 9.6 | 8.5 | 9 | 8.4 | 119.6 | 120 | 0.27 | 0.22 | 0.09 | -0.16 | 0.52 | 0.08 | -0.11 | 1.995 | 11.5 | 11.3 | 11.4 | 11.3 | 119.3 | 119.6 | LS |  | 2020-08-05 | 2.003 | 0.99451 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200812 | 0.32 | 0.14 | 0.18 | -0.01 | 0.48 | 0.13 | 0.09 | 1.991 | 9.5 | 8.5 | 9 | 8.9 | 119.6 | 119.8 | 0.33 | 0.26 | 0.13 | -0.18 | 0.58 | 0.1 | -0.17 | 1.992 | 11.5 | 11.3 | 11.4 | 11.3 | 119.3 | 119.7 | LS |  | 2020-08-12 | 2.003 | 0.99401 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200819 | 0.29 | 0.15 | 0.11 | 0.03 | 0.43 | 0.09 | 0.07 | 1.992 | 9.5 | 8.2 | 9 | 8.3 | 119.7 | 120 | 0.31 | 0.28 | 0.15 | -0.18 | 0.64 | 0.13 | -0.18 | 1.993 | 11.5 | 11.5 | 11.6 | 11.7 | 119.4 | 119.7 | AD |  | 2020-08-19 | 2.003 | 0.99451 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200826 | 0.25 | 0.2 | 0.12 | 0.08 | 0.45 | 0.12 | 0.08 | 1.996 | 9.6 | 8.3 | 8.8 | 8.7 | 119.6 | 119.9 | 0.22 | 0.24 | 0.02 | -0.14 | 0.49 | 0.07 | -0.12 | 1.994 | 11.5 | 11.5 | 11.5 | 11.6 | 119.4 | 119.4 | AD |  | 2020-08-26 | 2.003 | 0.99651 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200903 | 0.28 | 0.33 | 0.13 | 0.13 | 0.53 | 0.09 | 0.13 | 1.998 | 9.2 | 8.9 | 9 | 8.6 | 119.6 | 119.9 | 0.23 | 0.15 | 0.1 | -0.05 | 0.45 | 0.1 | -0.08 | 1.998 | 11.5 | 11.4 | 11.7 | 11.7 | 119.3 | 119.6 | AD |  | 2020-09-03 | 2.003 | 0.9975 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200909 | 0.32 | 0.17 | 0.18 | 0.04 | 0.5 | 0.14 | 0.08 | 2.004 | 9.5 | 8.5 | 8.9 | 8.6 | 119.6 | 119.9 | 0.25 | 0.24 | 0.1 | -0.11 | 0.48 | 0.09 | -0.11 | 2.002 | 11.5 | 11.4 | 11.5 | 11.5 | 119.3 | 119.5 | AD |  | 2020-09-09 | 2.003 | 1.0005 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200916 | 0.28 | 0.44 | 0.15 | 0.11 | 0.49 | 0.06 | 0.12 | 2.006 | 9.4 | 8.8 | 9.2 | 8.4 | 119.6 | 119.9 | 0.25 | 0.21 | 0.09 | -0.12 | 0.49 | 0.08 | -0.07 | 2.004 | 11.4 | 11.1 | 11.6 | 11.5 | 119.3 | 119.6 | AD |  | 2020-09-16 | 2.003 | 1.0015 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200923 | 0.25 | 0.26 | 0.14 | 0.01 | 0.44 | 0.08 | 0.03 | 2.004 | 9.3 | 8.8 | 9.1 | 8.3 | 119.6 | 119.9 | 0.23 | 0.31 | 0.1 | -0.17 | 0.57 | 0.1 | -0.16 | 2.002 | 11.4 | 11.2 | 11.4 | 11.2 | 119.3 | 119.6 | AD |  | 2020-09-23 | 2.003 | 1.0005 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200930 | 0.31 | 0.27 | 0.17 | 0.02 | 0.46 | 0.12 | 0.03 | 2.005 | 9.2 | 9 | 9.3 | 8.1 | 119.6 | 119.9 | 0.26 | 0.27 | 0.13 | -0.14 | 0.57 | 0.11 | -0.12 | 2 | 11.3 | 11.2 | 11.6 | 11.6 | 119.4 | 119.5 | AD |  | 2020-09-30 | 2.003 | 1.001 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201007 | 0.3 | 0.19 | 0.15 | 0.04 | 0.51 | 0.14 | 0.05 | 1.999 | 9.6 | 8.5 | 8.7 | 8.9 | 119.5 | 119.9 | 0.28 | 0.24 | 0.1 | -0.13 | 0.51 | 0.08 | -0.13 | 1.997 | 11.5 | 11.3 | 11.5 | 11.3 | 119.2 | 119.6 | AD |  | 2020-10-07 | 2.003 | 0.998 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201012 | 0.26 | 0.38 | 0.12 | 0.14 | 0.54 | 0.08 | 0.09 | 1.998 | 9.3 | 8.7 | 8.7 | 9 | 119.7 | 120 | 0.28 | 0.21 | 0.1 | -0.08 | 0.45 | 0.07 | -0.09 | 1.992 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.4 | ML |  | 2020-10-12 | 2.003 | 0.9975 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201021 | 0.28 | 0.26 | 0.13 | 0.06 | 0.41 | 0.1 | 0.04 | 2 | 9.2 | 8.9 | 9.2 | 8.5 | 119.6 | 119.8 | 0.24 | 0.28 | 0.11 | -0.17 | 0.59 | 0.09 | -0.15 | 1.997 | 11.5 | 11.4 | 11.7 | 11.7 | 119.3 | 119.6 | ML |  | 2020-10-21 | 2.003 | 0.9985 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201030 | 0.26 | 0.26 | 0.13 | 0.01 | 0.41 | 0.08 | 0.03 | 1.999 | 8.7 | 9.2 | 9.1 | 8.4 | 119.7 | 119.9 | 0.24 | 0.24 | 0.11 | -0.08 | 0.55 | 0.07 | -0.11 | 1.994 | 11.3 | 11.4 | 11.7 | 11.8 | 119.5 | 119.3 | ML |  | 2020-10-30 | 2.003 | 0.998 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201105 | 0.18 | 0.16 | 0.01 | 0 | 0.46 | 0.04 | 0.03 | 1.998 | 9.5 | 8.6 | 9.5 | 8.2 | 119.6 | 119.7 | 0.19 | 0.2 | 0.01 | -0.11 | 0.49 | 0.03 | -0.12 | 1.998 | 11.6 | 11.4 | 11.7 | 11.8 | 119.3 | 119.6 | ML |  | 2020-11-05 | 2.003 | 0.9975 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201111 | 0.39 | 0.37 | 0.24 | 0.07 | 0.5 | 0.15 | 0.04 | 2.003 | 9.3 | 8.9 | 9.1 | 8.8 | 119.6 | 119.7 | 0.27 | 0.23 | 0.1 | -0.16 | 0.56 | 0.07 | -0.15 | 1.993 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.5 | ML |  | 2020-11-11 | 2.003 | 1 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201119 | 0.37 | 0.17 | 0.24 | 0.07 | 0.57 | 0.18 | 0.1 | 2.004 | 9.6 | 8.4 | 9.1 | 8.7 | 119.5 | 119.8 | 0.34 | 0.22 | 0.17 | -0.11 | 0.55 | 0.1 | -0.11 | 2 | 11.6 | 11.7 | 11.7 | 11.8 | 119.3 | 119.7 | ML, Matrix_1 |  | 2020-11-19 | 2.003 | 1.0005 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201127 | 0.29 | 0.33 | 0.01 | 0.17 | 0.7 | 0.02 | 0.17 | 2 | 9.3 | 8.5 | 8.8 | 8.6 | 119.7 | 119.9 | 0.16 | 0.22 | 0.04 | -0.08 | 0.35 | 0.03 | -0.04 | 1.996 | 11.4 | 11.4 | 11.5 | 11.7 | 119.4 | 119.7 | ML, Matrix_1 |  | 2020-11-27 | 2.003 | 0.9985 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201202 | 0.28 | 0.31 | 0 | 0.19 | 0.69 | 0.01 | 0.18 | 1.997 | 9.3 | 8.5 | 8.7 | 8.8 | 119.7 | 119.8 | 0.23 | 0.25 | -0.01 | -0.1 | 0.44 | -0.02 | -0.08 | 1.99 | 11.5 | 11.6 | 11.5 | 11.8 | 119.6 | 119.5 | ML, Matrix_1 |  | 2020-12-02 | 2.003 | 0.997 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201209 | 0.33 | 0.38 | -0.04 | 0.26 | 0.69 | 0 | 0.24 | 1.998 | 9.1 | 8.8 | 8.6 | 9 | 119.7 | 119.8 | 0.21 | 0.23 | 0.04 | -0.08 | 0.4 | 0.02 | -0.04 | 1.992 | 11.4 | 11.4 | 11.5 | 11.6 | 119.4 | 119.6 | ML, Matrix_1 |  | 2020-12-09 | 2.003 | 0.9975 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201216 | 0.56 | 0.26 | 0.19 | 0.15 | 0.72 | 0.12 | 0.15 | 1.999 | 9.4 | 8.6 | 8.6 | 9 | 119.6 | 119.8 | 0.2 | 0.14 | -0.02 | -0.04 | 0.39 | 0.02 | -0.02 | 1.99 | 11.4 | 11.5 | 11.4 | 11.5 | 119.5 | 119.5 | ML, Matrix_1 |  | 2020-12-16 | 2.003 | 0.998 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201227 | 0.26 | 0.39 | -0.01 | 0.23 | 0.72 | 0.05 | 0.22 | 2 | 9.3 | 8.5 | 8.9 | 8.8 | 119.7 | 119.8 | 0.26 | 0.17 | -0.07 | -0.04 | 0.38 | -0.02 | -0.03 | 1.997 | 11.4 | 11.4 | 11.5 | 11.6 | 119.4 | 119.5 | ML, Matrix_1 |  | 2020-12-27 | 2.003 | 0.9985 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201230 | 0.33 | 0.44 | -0.07 | 0.23 | 0.65 | -0.01 | 0.18 | 1.994 | 9.2 | 8.6 | 9.1 | 8.4 | 119.7 | 119.8 | 0.14 | 0.33 | 0 | -0.14 | 0.42 | 0.01 | -0.09 | 1.988 | 11.4 | 11.4 | 11.5 | 11.8 | 119.5 | 119.5 | ML, Matrix_1 |  | 2020-12-30 | 2.003 | 0.99551 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210107 | 0.31 | 0.34 | 0.03 | 0.23 | 0.63 | 0.04 | 0.21 | 1.999 | 9.2 | 8.8 | 9 | 8.6 | 119.7 | 119.8 | 0.19 | 0.23 | 0.08 | -0.09 | 0.41 | 0.08 | -0.07 | 1.995 | 11.4 | 11.4 | 11.5 | 11.7 | 119.4 | 119.5 | ML, Matrix_1 |  | 2021-01-07 | 2.003 | 0.998 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210113 | 0.42 | 0.4 | 0 | 0.26 | 0.64 | 0.02 | 0.2 | 1.986 | 9.1 | 8.8 | 8.5 | 8.8 | 119.7 | 120 | 0.15 | 0.25 | 0.01 | -0.11 | 0.46 | 0.06 | -0.07 | 1.979 | 11.4 | 11.3 | 11.4 | 11.7 | 119.4 | 119.8 | LL, Matrix_1 |  | 2021-01-13 | 2.003 | 0.99151 | 2 | 0.9895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210120 | 0.51 | 0.36 | -0.03 | 0.21 | 0.7 | 0 | 0.2 | 1.987 | 9.1 | 8.6 | 8.4 | 8.7 | 119.8 | 120.1 | 0.24 | 0.24 | -0.06 | -0.09 | 0.4 | -0.02 | -0.06 | 1.984 | 11.4 | 11.4 | 11.4 | 11.5 | 119.5 | 119.8 | ML, Matrix_1 |  | 2021-01-20 | 2.003 | 0.99201 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210128 | 0.43 | 0.39 | 0.04 | 0.24 | 0.6 | 0.03 | 0.19 | 1.997 | 9.2 | 8.7 | 8.6 | 8.6 | 119.7 | 120 | 0.18 | 0.25 | 0.05 | -0.07 | 0.37 | 0.04 | -0.06 | 1.989 | 11.4 | 11.3 | 11.3 | 11.4 | 119.5 | 119.8 | ML, Matrix_1 |  | 2021-01-28 | 2.003 | 0.997 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210129 | 0.38 | 0.3 | -0.07 | 0.2 | 0.68 | -0.02 | 0.17 | 1.993 | 8.9 | 9 | 8.7 | 8.5 | 119.7 | 120 | 0.22 | 0.17 | 0.04 | -0.05 | 0.32 | 0.04 | -0.05 | 1.988 | 11.4 | 11.3 | 11.5 | 11.7 | 119.4 | 119.7 | ML, Matrix_1 |  | 2021-01-29 | 2.003 | 0.99501 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210203 | 0.3 | 0.4 | 0.02 | 0.25 | 0.68 | 0.03 | 0.21 | 1.993 | 9 | 9 | 9.1 | 8.2 | 119.6 | 119.9 | 0.22 | 0.24 | 0.02 | -0.04 | 0.38 | 0.01 | -0.03 | 1.988 | 11.4 | 11.3 | 11.6 | 11.8 | 119.4 | 119.7 | ML, Matrix_1 |  | 2021-02-03 | 2.003 | 0.99501 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210211 | 0.34 | 0.29 | 0.02 | 0.17 | 0.7 | 0.01 | 0.19 | 1.996 | 9.4 | 8.5 | 8.8 | 8.5 | 119.6 | 120 | 0.19 | 0.62 | 0.05 | -0.3 | 0.73 | 0.07 | -0.27 | 1.981 | 11.5 | 11.5 | 11.3 | 11.6 | 119.5 | 119.7 | ML, Matrix_1 |  | 2021-02-11 | 2.003 | 0.99651 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210215 | 0.46 | 0.45 | -0.02 | 0.28 | 0.7 | 0.01 | 0.23 | 1.987 | 9.1 | 8.8 | 8.5 | 8.8 | 119.7 | 120 | 0.17 | 0.38 | 0.05 | -0.13 | 0.57 | 0.08 | -0.09 | 1.978 | 11.5 | 11.4 | 11.2 | 11.4 | 119.4 | 119.7 | ML, Matrix_1 |  | 2021-02-15 | 2.003 | 0.99201 | 2 | 0.989 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210223 | 0.26 | 0.41 | 0.13 | 0.23 | 0.88 | 0.13 | 0.25 | 2.012 | 9.4 | 8.6 | 9 | 8.5 | 119.6 | 120 | 0.27 | 0.19 | 0.05 | -0.05 | 0.63 | 0.12 | 0.02 | 2.007 | 11.5 | 11.4 | 11.5 | 11.4 | 119.4 | 119.8 | ML, MatriXX_2 |  | 2021-02-23 | 2.003 | 1.0045 | 2 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210224 | 0.33 | 0.38 | -0.01 | 0.23 | 0.68 | 0.05 | 0.2 | 2.006 | 9.6 | 7.9 | 8 | 9.1 | 119.6 | 119.9 | 0.18 | 0.2 | -0.03 | -0.04 | 0.4 | -0.01 | -0.04 | 1.994 | 11.5 | 11.7 | 11.2 | 11.2 | 119.4 | 119.8 | ML, MatriXX_1 |  | 2021-02-24 | 2.003 | 1.0015 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210303 | 0.36 | 0.36 | 0.02 | 0.18 | 0.61 | 0.01 | 0.15 | 1.999 | 9.3 | 8.6 | 9 | 8.2 | 119.7 | 120 | 0.25 | 0.2 | 0.06 | 0.01 | 0.36 | 0.06 | -0.05 | 1.989 | 11.5 | 11.5 | 11.5 | 11.8 | 119.5 | 119.9 | ML, MatriXX_1 |  | 2021-03-03 | 2.003 | 0.998 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210310 | 0.32 | 0.46 | -0.05 | 0.27 | 0.71 | -0.02 | 0.2 | 1.992 | 9.3 | 8.6 | 8.9 | 8.4 | 119.7 | 120 | 0.16 | 0.09 | 0.03 | -0.01 | 0.38 | 0.08 | 0.02 | 1.987 | 11.5 | 11.5 | 11.6 | 11.7 | 119.4 | 119.7 | ML, MatriXX_1 |  | 2021-03-10 | 2.003 | 0.99451 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210324 | 0.27 | 0.46 | 0.15 | 0.29 | 0.82 | 0.15 | 0.26 | 1.998 | 9.5 | 8.6 | 9.1 | 9 | 119.6 | 119.8 | 0.29 | 0.17 | 0.08 | 0.01 | 0.5 | 0.09 | 0 | 1.992 | 11.6 | 11.6 | 11.6 | 11.5 | 119.4 | 119.7 | ML, MatriXX_2 |  | 2021-03-24 | 2.003 | 0.9975 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210325 | 0.28 | 0.31 | -0.01 | 0.15 | 0.55 | 0.01 | 0.11 | 1.998 | 8.9 | 9 | 9.1 | 8.1 | 119.7 | 120 | 0.22 | 0.19 | 0.04 | 0.06 | 0.38 | 0.03 | 0.03 | 1.99 | 11.4 | 11.5 | 11.7 | 12 | 119.5 | 119.6 | ML, MatriXX_1 |  | 2021-03-25 | 2.003 | 0.9975 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210331 | 0.28 | 0.44 | 0.1 | 0.17 | 0.65 | 0.11 | 0.21 | 2.002 | 9.4 | 8.7 | 8.8 | 9 | 119.6 | 119.9 | 0.29 | 0.22 | 0.1 | 0.02 | 0.6 | 0.13 | 0.06 | 1.995 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | ML |  | 2021-03-31 | 2.003 | 0.9995 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210408 | 0.25 | 0.41 | 0.09 | 0.19 | 0.69 | 0.09 | 0.23 | 1.999 | 9.2 | 8.9 | 8.9 | 8.8 | 119.6 | 119.9 | 0.3 | 0.11 | 0.08 | 0 | 0.53 | 0.1 | -0.01 | 1.992 | 11.5 | 11.4 | 11.6 | 11.6 | 119.4 | 119.7 | ML, MatriXX_2 |  | 2021-04-08 | 2.003 | 0.998 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210414 | 0.33 | 0.32 | -0.06 | 0.19 | 0.68 | 0 | 0.18 | 1.997 | 8.8 | 9.2 | 9.1 | 8.2 | 119.7 | 120 | 0.17 | 0.23 | 0.02 | -0.05 | 0.35 | 0.02 | -0.02 | 1.988 | 11.3 | 11.5 | 11.6 | 11.9 | 119.5 | 119.8 | ML, Matrixx_1 |  | 2021-04-14 | 2.003 | 0.997 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210421 | 0.28 | 0.39 | 0.1 | 0.18 | 0.7 | 0.1 | 0.18 | 2.003 | 9.3 | 8.9 | 9.3 | 8.4 | 119.6 | 119.9 | 0.24 | 0.13 | 0.06 | -0.02 | 0.46 | 0.09 | -0.01 | 1.994 | 11.6 | 11.6 | 11.7 | 11.9 | 119.4 | 119.7 | ML, Matrixx_2 |  | 2021-04-21 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210428 | 0.24 | 0.3 | 0.01 | -0.1 | 0.41 | 0.01 | -0.09 | 2.006 | 11.4 | 11.4 | 11.6 | 11.8 | 119.4 | 119.7 | 0.24 | 0.3 | 0.01 | -0.1 | 0.41 | 0.01 | -0.09 | 2.006 | 11.4 | 11.4 | 11.6 | 11.8 | 119.4 | 119.7 | ML, MatriXX_1 |  | 2021-04-28 | 2.003 | 1.0015 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.33 | 0.29 | 0.01 | 0.14 | 0.63 | -0.03 | 0.13 | 2.004 | 9.2 | 8.8 | 8.6 | 8.6 | 119.7 | 120 | 0.25 | 0.42 | 0 | -0.18 | 0.57 | 0.01 | -0.19 | 1.993 | 11.4 | 11.3 | 11.3 | 11.4 | 119.5 | 119.8 | ML, MatriXX_1, run 1 |  | 2021-04-30 | 2.003 | 1.0005 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.51 | 0.27 | -0.03 | 0.16 | 0.63 | -0.01 | 0.13 | 2.006 | 9.2 | 8.7 | 8.5 | 8.6 | 119.7 | 120 | 0.19 | 0.46 | 0.03 | -0.23 | 0.55 | 0.03 | -0.19 | 1.992 | 11.4 | 11.3 | 11.3 | 11.5 | 119.5 | 119.8 | ML, MatriXX_1, run 2 |  | 2021-04-30 | 2.003 | 1.0015 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.5 | 0.29 | -0.03 | 0.16 | 0.58 | -0.02 | 0.13 | 2.005 | 9.2 | 8.8 | 8.5 | 8.6 | 119.7 | 120.1 | 0.21 | 0.43 | 0.06 | -0.2 | 0.59 | 0.05 | -0.2 | 1.991 | 11.4 | 11.3 | 11.3 | 11.5 | 119.5 | 119.8 | ML, MatriXX_1, run 3 |  | 2021-04-30 | 2.003 | 1.001 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210507 | 0.44 | 0.25 | 0 | 0.16 | 0.61 | 0.02 | 0.16 | 1.992 | 9.4 | 8.5 | 8.4 | 9 | 119.6 | 120 | 0.19 | 0.26 | 0.02 | -0.1 | 0.37 | 0.02 | -0.08 | 1.983 | 11.5 | 11.7 | 11.4 | 11.6 | 119.4 | 119.8 | ML, MatriXX_1 |  | 2021-05-07 | 2.003 | 0.99451 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210509 | 0.51 | 0.38 | -0.04 | 0.17 | 0.69 | -0.03 | 0.17 | 1.995 | 9.5 | 8.3 | 8.3 | 9 | 119.7 | 120 | 0.22 | 0.66 | -0.03 | -0.3 | 0.7 | 0 | -0.28 | 1.973 | 11.5 | 11.8 | 11.4 | 11.9 | 119.6 | 119.9 | ML, Matrixx 1 |  | 2021-05-09 | 2.003 | 0.99601 | 2 | 0.9865 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210512 | 0.24 | 0.42 | 0.09 | 0.2 | 0.73 | 0.11 | 0.22 | 2.01 | 9.4 | 8.7 | 9 | 8.5 | 119.7 | 120 | 0.24 | 0.16 | 0.03 | -0.01 | 0.51 | 0.07 | 0.01 | 2.004 | 11.6 | 11.7 | 11.7 | 11.8 | 119.4 | 119.7 | ML, MatriXX_2 |  | 2021-05-12 | 2.003 | 1.0035 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210520 | 0.55 | 0.44 | -0.12 | 0.24 | 0.68 | -0.01 | 0.2 | 1.997 | 8.7 | 9.2 | 8.4 | 8.9 | 119.7 | 120 | 0.29 | 0.24 | 0.01 | -0.07 | 0.44 | 0.02 | -0.08 | 1.987 | 11.5 | 11.5 | 11.5 | 11.7 | 119.4 | 119.8 | ML, MatriXX_1 |  | 2021-05-20 | 2.003 | 0.997 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210526 | 0.22 | 0.46 | 0.08 | 0.21 | 0.67 | 0.12 | 0.22 | 2.006 | 9 | 9.2 | 9.1 | 8.6 | 119.6 | 119.9 | 0.28 | 0.17 | 0.09 | 0.01 | 0.56 | 0.12 | 0.03 | 1.993 | 11.5 | 11.5 | 11.8 | 11.9 | 119.4 | 119.7 | ML, MatriXX_2 |  | 2021-05-26 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210603 | 0.3 | 0.49 | 0.11 | 0.24 | 0.69 | 0.13 | 0.24 | 2.02 | 9.5 | 8.6 | 9.1 | 8.6 | 119.6 | 119.9 | 0.3 | 0.15 | 0.09 | -0.02 | 0.5 | 0.11 | 0 | 2.012 | 11.6 | 11.7 | 11.8 | 11.9 | 119.3 | 119.6 | ML, MatriXX_2 |  | 2021-06-03 | 2.003 | 1.0085 | 2 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210607 | 0.32 | 0.41 | 0.09 | 0.19 | 0.65 | 0.08 | 0.22 | 2.021 | 9.4 | 8.7 | 9 | 8.7 | 119.6 | 119.9 | 0.26 | 0.09 | 0.05 | -0.01 | 0.5 | 0.08 | 0.01 | 2.011 | 11.6 | 11.7 | 11.6 | 11.6 | 119.4 | 119.7 | cvd, Matrix 2 |  | 2021-06-07 | 2.003 | 1.009 | 2 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210608 | 0.25 | 0.41 | 0.01 | 0.2 | 0.66 | 0.07 | 0.22 | 2.021 | 9.4 | 8.7 | 9.4 | 8.1 | 119.6 | 119.9 | 0.28 | 0.15 | 0.04 | -0.01 | 0.47 | 0.05 | -0.01 | 2.019 | 11.8 | 11.9 | 11.9 | 12.2 | 119.3 | 119.5 | GB, Matrix 2 |  | 2021-06-08 | 2.003 | 1.009 | 2 | 1.0095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210609 | 0.32 | 0.45 | 0.09 | 0.22 | 0.63 | 0.09 | 0.22 | 2.017 | 9.4 | 8.6 | 9 | 8.7 | 119.6 | 119.9 | 0.3 | 0.19 | 0.11 | -0.08 | 0.49 | 0.12 | -0.05 | 2.008 | 11.6 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | CVD, Matrix 2 |  | 2021-06-09 | 2.003 | 1.007 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210610 | 0.23 | 0.56 | 0.06 | 0.28 | 0.69 | 0.09 | 0.28 | 2.023 | 9.3 | 8.8 | 9.2 | 8.7 | 119.7 | 119.8 | 0.28 | 0.15 | 0.08 | -0.01 | 0.46 | 0.1 | -0.01 | 2.008 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.6 | CVD, Matrix 2 |  | 2021-06-10 | 2.003 | 1.01 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210611 | 0.3 | 0.67 | 0.17 | 0.31 | 0.87 | 0.11 | 0.29 | 2.016 | 9.5 | 8.5 | 8.2 | 9.4 | 119.6 | 119.7 | 0.4 | 0.17 | 0.17 | 0 | 0.55 | 0.16 | -0.01 | 2.005 | 11.6 | 11.7 | 11.8 | 11.4 | 119.4 | 119.7 | CVD, Matrix 2 |  | 2021-06-11 | 2.003 | 1.0065 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210615 | 0.25 | 0.42 | 0.09 | 0.21 | 0.7 | 0.09 | 0.26 | 2.02 | 8.9 | 9.3 | 9 | 8.8 | 119.6 | 119.9 | 0.29 | 0.19 | 0.09 | -0.07 | 0.54 | 0.1 | -0.04 | 2.017 | 11.5 | 11.6 | 11.6 | 11.5 | 119.4 | 119.7 | CVD, Matrix 2 |  | 2021-06-15 | 2.003 | 1.0085 | 2 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210616 | 0.31 | 0.47 | 0.15 | 0.28 | 0.79 | 0.13 | 0.28 | 2.021 | 9.5 | 8.5 | 8.8 | 8.9 | 119.7 | 119.9 | 0.27 | 0.15 | 0.1 | -0.03 | 0.51 | 0.13 | -0.01 | 2.008 | 11.6 | 11.8 | 11.5 | 11.5 | 119.5 | 119.8 | CVD, Matrix 2 |  | 2021-06-16 | 2.003 | 1.009 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210617 | 0.26 | 0.36 | 0.1 | 0.2 | 0.66 | 0.1 | 0.22 | 2.023 | 8.7 | 9.4 | 9 | 8.7 | 119.6 | 119.9 | 0.26 | 0.15 | 0.05 | 0 | 0.46 | 0.09 | -0.01 | 2.015 | 11.5 | 11.6 | 11.6 | 11.6 | 119.4 | 119.7 | CVD, Matrix 2 |  | 2021-06-17 | 2.003 | 1.01 | 2 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210618 | 0.3 | 0.47 | 0.14 | 0.3 | 0.83 | 0.17 | 0.3 | 2.016 | 9.4 | 8.6 | 9.1 | 8.6 | 119.6 | 119.9 | 0.32 | 0.16 | 0.11 | -0.02 | 0.53 | 0.12 | 0.02 | 2.006 | 11.6 | 11.7 | 11.7 | 11.8 | 119.5 | 119.8 | GB |  | 2021-06-18 | 2.003 | 1.0065 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210622 | 0.26 | 0.46 | 0.07 | 0.25 | 0.67 | 0.09 | 0.26 | 2.018 | 9.5 | 8.6 | 9 | 8.6 | 119.6 | 119.9 | 0.23 | 0.14 | 0.06 | -0.01 | 0.48 | 0.08 | 0.03 | 2.009 | 11.6 | 11.5 | 11.6 | 11.5 | 119.5 | 119.8 | GB |  | 2021-06-22 | 2.003 | 1.0075 | 2 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210623 | 0.29 | 0.46 | 0.13 | 0.19 | 0.72 | 0.14 | 0.27 | 2.018 | 9.3 | 8.8 | 8.7 | 8.9 | 119.7 | 119.9 | 0.32 | 0.15 | 0.12 | -0.04 | 0.49 | 0.12 | 0.01 | 2.01 | 11.5 | 11.6 | 11.5 | 11.4 | 119.5 | 119.8 | GB |  | 2021-06-23 | 2.003 | 1.0075 | 2 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210624 | 0.21 | 0.44 | 0.1 | 0.28 | 0.65 | 0.11 | 0.26 | 2.014 | 9.5 | 8.5 | 9 | 8.7 | 119.6 | 119.9 | 0.25 | 0.16 | 0.03 | -0.02 | 0.49 | 0.08 | 0.01 | 2.005 | 11.6 | 11.6 | 11.6 | 11.5 | 119.4 | 119.7 | GB |  | 2021-06-24 | 2.003 | 1.0055 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210629 | 0.25 | 0.4 | 0.08 | 0.2 | 0.7 | 0.1 | 0.27 | 2.009 | 9.4 | 8.6 | 9 | 8.7 | 119.6 | 119.8 | 0.24 | 0.13 | 0.04 | -0.02 | 0.51 | 0.09 | 0.01 | 2.001 | 11.6 | 11.7 | 11.6 | 11.6 | 119.4 | 119.7 | GB |  | 2021-06-29 | 2.003 | 1.003 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210630 | 0.33 | 0.39 | 0.16 | 0.16 | 0.78 | 0.15 | 0.19 | 2.01 | 9.7 | 8.1 | 9.1 | 8.6 | 119.5 | 119.9 | 0.25 | 0.17 | 0.08 | 0 | 0.47 | 0.08 | 0 | 2.002 | 11.7 | 11.9 | 11.6 | 11.6 | 119.4 | 119.7 | GB |  | 2021-06-30 | 2.003 | 1.0035 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210701 | 0.35 | 0.38 | 0.17 | 0.23 | 0.65 | 0.11 | 0.22 | 2.019 | 9.5 | 8.6 | 8.8 | 8.9 | 119.6 | 119.8 | 0.25 | 0.15 | 0.08 | -0.01 | 0.51 | 0.11 | -0.01 | 2.012 | 11.6 | 11.5 | 11.5 | 11.1 | 119.3 | 119.6 | ML |  | 2021-07-01 | 2.003 | 1.008 | 2 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210706 | 0.27 | 0.55 | 0.09 | 0.22 | 0.72 | 0.12 | 0.27 | 2.001 | 9.1 | 9 | 8.7 | 9 | 119.7 | 119.9 | 0.31 | 0.18 | 0.11 | -0.01 | 0.63 | 0.17 | 0 | 1.998 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.6 | ML |  | 2021-07-06 | 2.003 | 0.999 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210707 | 0.28 | 0.47 | 0.13 | 0.27 | 0.8 | 0.13 | 0.27 | 2.005 | 9.3 | 8.8 | 8.9 | 8.6 | 119.7 | 119.9 | 0.25 | 0.12 | 0.09 | -0.01 | 0.53 | 0.12 | 0.01 | 2.004 | 11.6 | 11.6 | 11.5 | 11.3 | 119.3 | 119.6 | ML |  | 2021-07-07 | 2.003 | 1.001 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210708 | 0.3 | 0.54 | 0.17 | 0.23 | 0.86 | 0.14 | 0.27 | 2.005 | 9.2 | 8.9 | 9.2 | 8.4 | 119.6 | 119.9 | 0.28 | 0.12 | 0.11 | 0.01 | 0.54 | 0.13 | 0.04 | 1.996 | 11.5 | 11.4 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-07-08 | 2.003 | 1.001 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210709 | 0.23 | 0.52 | 0.08 | 0.26 | 0.83 | 0.18 | 0.25 | 2.008 | 9.6 | 8.4 | 9 | 8.6 | 119.6 | 119.9 | 0.23 | 0.22 | 0.08 | -0.01 | 0.58 | 0.1 | 0.02 | 2.002 | 11.6 | 11.7 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-07-09 | 2.003 | 1.0025 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210714 | 0.33 | 0.58 | 0.13 | 0.21 | 0.85 | 0.14 | 0.25 | 2.002 | 9.3 | 8.8 | 8.9 | 8.7 | 119.6 | 119.9 | 0.31 | 0.18 | 0.11 | -0.01 | 0.63 | 0.14 | 0.03 | 1.996 | 11.5 | 11.4 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-07-14 | 2.003 | 0.9995 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210715 | 0.27 | 0.53 | 0.14 | 0.25 | 0.82 | 0.15 | 0.29 | 2.016 | 9 | 9.4 | 9.3 | 8.8 | 119.3 | 119.6 | 0.3 | 0.22 | 0.1 | -0.07 | 0.51 | 0.13 | -0.03 | 2.004 | 11.5 | 11.5 | 11.6 | 11.4 | 119.1 | 119.4 | CVD |  | 2021-07-15 | 2.003 | 1.0065 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210716 | 0.3 | 0.44 | 0.11 | 0.22 | 0.76 | 0.12 | 0.24 | 1.998 | 9.2 | 8.8 | 9.1 | 8.4 | 119.7 | 120 | 0.26 | 0.21 | 0.05 | -0.06 | 0.54 | 0.12 | 0 | 1.991 | 11.5 | 11.5 | 11.6 | 11.6 | 119.5 | 119.7 | ML |  | 2021-07-16 | 2.003 | 0.9975 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210719 | 0.39 | 0.36 | 0.17 | 0.19 | 0.78 | 0.2 | 0.23 | 1.998 | 9.6 | 8.3 | 8.9 | 8.6 | 119.6 | 120 | 0.35 | 0.26 | 0.13 | -0.08 | 0.57 | 0.15 | -0.03 | 1.987 | 11.6 | 11.9 | 11.7 | 11.8 | 119.5 | 119.8 | ML |  | 2021-07-19 | 2.003 | 0.9975 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210721 | 0.3 | 0.39 | 0.09 | 0.17 | 0.68 | 0.12 | 0.21 | 2.001 | 9.2 | 9 | 8.9 | 8.7 | 119.6 | 119.9 | 0.27 | 0.13 | 0.08 | -0.03 | 0.51 | 0.1 | 0.01 | 1.992 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.7 | ML |  | 2021-07-21 | 2.003 | 0.999 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210722 | 0.35 | 0.44 | 0.11 | 0.22 | 0.67 | 0.1 | 0.21 | 2.002 | 9.1 | 9.1 | 9.5 | 7.7 | 119.6 | 119.9 | 0.27 | 0.13 | 0.08 | -0.03 | 0.51 | 0.1 | 0.01 | 1.992 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.7 | ML |  | 2021-07-22 | 2.003 | 0.9995 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210723 | 0.35 | 0.49 | 0.13 | 0.22 | 0.74 | 0.13 | 0.26 | 2.003 | 9.3 | 8.8 | 8.9 | 8.9 | 119.6 | 119.9 | 0.23 | 0.21 | 0.05 | -0.09 | 0.56 | 0.1 | -0.03 | 1.99 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.7 | ML |  | 2021-07-23 | 2.003 | 1 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210727 | 0.31 | 0.43 | 0.15 | 0.24 | 0.8 | 0.18 | 0.24 | 1.998 | 9.5 | 8.5 | 8.9 | 8.6 | 119.7 | 120 | 0.24 | 0.17 | 0.06 | -0.03 | 0.57 | 0.12 | 0 | 1.992 | 11.6 | 11.6 | 11.6 | 11.5 | 119.5 | 119.8 | ML |  | 2021-07-27 | 2.003 | 0.9975 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210728 | 0.36 | 0.47 | 0.17 | 0.29 | 0.89 | 0.18 | 0.3 | 2.007 | 9.8 | 8.1 | 8.7 | 9 | 119.5 | 119.9 | 0.3 | 0.15 | 0.12 | 0.02 | 0.63 | 0.13 | 0.08 | 2 | 11.6 | 11.8 | 11.5 | 11.2 | 119.3 | 119.7 | ML |  | 2021-07-28 | 2.003 | 1.002 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210729 | 0.27 | 0.47 | 0.13 | 0.23 | 0.73 | 0.14 | 0.26 | 2.003 | 9.3 | 8.8 | 9 | 8.7 | 119.6 | 119.9 | 0.21 | 0.19 | 0.05 | -0.05 | 0.64 | 0.13 | 0.02 | 1.998 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-07-29 | 2.003 | 1 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210730 | 0.36 | 0.47 | 0.18 | 0.22 | 0.83 | 0.16 | 0.25 | 1.998 | 8.8 | 9.4 | 9 | 8.7 | 119.6 | 119.9 | 0.29 | 0.12 | 0.11 | -0.01 | 0.58 | 0.16 | 0.03 | 1.991 | 11.7 | 11.6 | 11.6 | 11.4 | 119.3 | 119.7 | ML |  | 2021-07-30 | 2.003 | 0.9975 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210803 | 0.29 | 0.42 | 0.12 | 0.18 | 0.77 | 0.14 | 0.24 | 2.002 | 9.1 | 9 | 9.1 | 8.5 | 119.7 | 120 | 0.25 | 0.22 | 0.1 | 0.03 | 0.61 | 0.13 | 0.04 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-08-03 | 2.003 | 0.9995 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210804 | 0.36 | 0.44 | 0.16 | 0.15 | 0.79 | 0.17 | 0.19 | 2.006 | 9.2 | 9.1 | 9.1 | 8.5 | 119.6 | 119.9 | 0.27 | 0.17 | 0.08 | 0.01 | 0.6 | 0.14 | 0.02 | 1.996 | 11.4 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-08-04 | 2.003 | 1.0015 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210805 | 0.34 | 0.62 | 0.16 | 0.33 | 1.06 | 0.16 | 0.34 | 2 | 9.3 | 8.9 | 9 | 8.8 | 119.6 | 119.9 | 0.25 | 0.14 | 0.06 | -0.02 | 0.56 | 0.11 | 0 | 1.994 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | ML |  | 2021-08-05 | 2.003 | 0.9985 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210806 | 0.23 | 0.5 | 0.11 | 0.24 | 0.8 | 0.13 | 0.25 | 2.002 | 9.1 | 9 | 9 | 8.5 | 119.6 | 119.9 | 0.27 | 0.26 | 0.1 | -0.06 | 0.65 | 0.14 | -0.02 | 1.993 | 11.4 | 11.4 | 11.6 | 11.4 | 119.4 | 119.7 | ML |  | 2021-08-06 | 2.003 | 0.9995 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210810 | 0.37 | 0.49 | 0.18 | 0.2 | 0.77 | 0.16 | 0.26 | 2.005 | 9.4 | 8.8 | 8.9 | 8.8 | 119.6 | 119.9 | 0.24 | 0.16 | 0.09 | -0.01 | 0.58 | 0.11 | 0 | 1.994 | 11.6 | 11.6 | 11.6 | 11.6 | 119.4 | 119.7 | ML |  | 2021-08-10 | 2.003 | 1.001 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210811 | 0.41 | 0.5 | 0.18 | 0.26 | 0.84 | 0.19 | 0.24 | 2.004 | 9.3 | 8.9 | 8.9 | 8.7 | 119.6 | 119.9 | 0.33 | 0.16 | 0.13 | 0.02 | 0.62 | 0.17 | 0.03 | 1.991 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | CVD |  | 2021-08-11 | 2.003 | 1.0005 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210812 | 0.35 | 0.48 | 0.17 | 0.26 | 0.82 | 0.18 | 0.27 | 2.001 | 8.9 | 9.3 | 9.1 | 8.4 | 119.6 | 119.9 | 0.4 | 0.2 | 0.16 | -0.04 | 0.66 | 0.19 | 0.02 | 2 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ML |  | 2021-08-12 | 2.003 | 0.999 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210813 | 0.44 | 0.42 | 0.2 | 0.25 | 0.93 | 0.22 | 0.26 | 1.998 | 9.5 | 8.7 | 8.6 | 8.9 | 119.6 | 119.9 | 0.33 | 0.2 | 0.12 | 0.02 | 0.63 | 0.16 | 0.03 | 1.988 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.8 | ML |  | 2021-08-13 | 2.003 | 0.9975 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210817 | 0.34 | 0.51 | 0.12 | 0.23 | 0.77 | 0.16 | 0.21 | 2 | 9.2 | 9 | 9 | 8.6 | 119.6 | 119.9 | 0.38 | 0.15 | 0.18 | -0.01 | 0.62 | 0.12 | 0.01 | 1.995 | 11.4 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | ML |  | 2021-08-17 | 2.003 | 0.9985 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210818 | 0.27 | 0.4 | 0.08 | 0.2 | 0.77 | 0.13 | 0.24 | 1.999 | 9.4 | 8.7 | 8.8 | 8.9 | 119.6 | 119.9 | 0.27 | 0.19 | 0.08 | 0 | 0.55 | 0.12 | 0.02 | 1.987 | 11.5 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | CVD |  | 2021-08-18 | 2.003 | 0.998 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210819 | 0.29 | 0.47 | 0.13 | 0.2 | 0.83 | 0.15 | 0.23 | 1.989 | 9.2 | 9 | 8.8 | 8.8 | 119.6 | 119.9 | 0.3 | 0.2 | 0.1 | -0.02 | 0.58 | 0.15 | 0.03 | 1.979 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | cvd |  | 2021-08-19 | 2.003 | 0.99301 | 2 | 0.9895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210820 | 0.29 | 0.44 | 0.08 | 0.21 | 0.71 | 0.12 | 0.26 | 2 | 9.4 | 8.8 | 8.9 | 8.8 | 119.6 | 119.9 | 0.29 | 0.15 | 0.09 | 0 | 0.6 | 0.13 | 0.04 | 1.982 | 11.6 | 11.6 | 11.4 | 11.2 | 119.3 | 119.7 | ML |  | 2021-08-20 | 2.003 | 0.9985 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210824 | 0.28 | 0.51 | 0.09 | 0.24 | 0.71 | 0.12 | 0.26 | 1.996 | 9.3 | 8.9 | 9 | 8.7 | 119.6 | 119.9 | 0.23 | 0.16 | 0.05 | -0.01 | 0.54 | 0.11 | 0.01 | 1.992 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.6 | CVD |  | 2021-08-24 | 2.003 | 0.99651 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210825 | 0.24 | 0.38 | 0.03 | 0.25 | 0.67 | 0.11 | 0.25 | 2.015 | 9.3 | 8.8 | 8.9 | 8.7 | 119.6 | 119.9 | 0.24 | 0.18 | 0.02 | -0.04 | 0.49 | 0.07 | -0.02 | 2.006 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | GB |  | 2021-08-25 | 2.003 | 1.006 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210826 | 0.28 | 0.48 | 0.11 | 0.22 | 0.65 | 0.1 | 0.21 | 1.999 | 9.2 | 9 | 9 | 8.5 | 119.6 | 119.9 | 0.23 | 0.18 | 0.05 | 0.03 | 0.58 | 0.11 | 0.03 | 1.991 | 11.4 | 11.3 | 11.6 | 11.5 | 119.4 | 119.7 | GB |  | 2021-08-26 | 2.003 | 0.998 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210831 | 0.38 | 0.53 | 0.2 | 0.29 | 0.91 | 0.21 | 0.28 | 1.984 | 9.1 | 9 | 9.1 | 8.5 | 119.6 | 119.9 | 0.45 | 0.2 | 0.22 | -0.01 | 0.67 | 0.22 | 0.02 | 1.982 | 11.3 | 11.4 | 11.5 | 11.4 | 119.4 | 119.7 | CVD, Matrix 2 |  | 2021-08-31 | 2.003 | 0.99051 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210901 | 0.36 | 0.44 | 0.14 | 0.23 | 0.82 | 0.18 | 0.24 | 1.986 | 9.2 | 8.9 | 8.9 | 8.6 | 119.6 | 120 | 0.32 | 0.15 | 0.12 | -0.04 | 0.59 | 0.17 | 0 | 1.979 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | GB |  | 2021-09-01 | 2.003 | 0.99151 | 2 | 0.9895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210902 | 0.18 | 0.49 | 0.04 | 0.2 | 0.64 | 0.07 | 0.19 | 1.997 | 9.1 | 9.1 | 9.2 | 8.4 | 119.6 | 119.9 | 0.26 | 0.21 | 0 | 0.03 | 0.57 | 0.06 | 0.03 | 1.99 | 11.4 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | GB |  | 2021-09-02 | 2.003 | 0.997 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210903 | 0.3 | 0.44 | 0.09 | 0.17 | 0.73 | 0.11 | 0.2 | 1.998 | 9.4 | 8.8 | 8.7 | 8.9 | 119.6 | 119.9 | 0.18 | 0.22 | 0.06 | 0.06 | 0.56 | 0.09 | 0.05 | 1.986 | 11.5 | 11.4 | 11.5 | 11.5 | 119.4 | 119.8 | GB |  | 2021-09-03 | 2.003 | 0.9975 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210906 | 0.45 | 0.41 | 0.05 | 0.23 | 0.69 | 0.04 | 0.19 | 1.99 | 9.2 | 8.8 | 8.5 | 8.7 | 119.7 | 120 | 0.23 | 0.2 | 0.1 | -0.03 | 0.31 | 0.06 | -0.02 | 1.983 | 11.4 | 11.4 | 11.3 | 11.3 | 119.4 | 119.8 | CVD, Matrixx 1 |  | 2021-09-06 | 2.003 | 0.99351 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210908 | 0.33 | 0.44 | 0.14 | 0.21 | 0.86 | 0.18 | 0.24 | 1.985 | 9.1 | 9.1 | 8.8 | 8.8 | 119.6 | 119.9 | 0.3 | 0.14 | 0.08 | -0.02 | 0.47 | 0.1 | -0.01 | 1.982 | 11.4 | 11.4 | 11.4 | 11.1 | 119.4 | 119.7 | ML |  | 2021-09-08 | 2.003 | 0.99101 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210909 | 0.29 | 0.41 | 0.12 | 0.26 | 0.85 | 0.12 | 0.26 | 1.986 | 9 | 9.1 | 9.3 | 8.2 | 119.7 | 119.9 | 0.31 | 0.19 | 0.1 | 0.05 | 0.55 | 0.1 | 0.01 | 1.975 | 11.5 | 11.5 | 11.7 | 11.7 | 119.5 | 119.7 | GB |  | 2021-09-09 | 2.003 | 0.99151 | 2 | 0.9875 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210910 | 0.23 | 0.46 | 0.04 | 0.24 | 0.81 | 0.12 | 0.22 | 1.988 | 9.6 | 8.3 | 8.9 | 8.8 | 119.6 | 119.9 | 0.24 | 0.17 | 0.05 | -0.05 | 0.57 | 0.09 | -0.05 | 1.98 | 11.5 | 11.5 | 11.7 | 11.3 | 119.3 | 119.6 | GB |  | 2021-09-10 | 2.003 | 0.99251 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210914 | 0.43 | 0.5 | 0.22 | 0.27 | 0.86 | 0.2 | 0.31 | 1.988 | 9.2 | 9 | 9 | 8.7 | 119.6 | 119.9 | 0.29 | 0.17 | 0.11 | -0.01 | 0.61 | 0.15 | 0.04 | 1.982 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.7 | ML |  | 2021-09-14 | 2.003 | 0.99251 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210916 | 0.3 | 0.55 | 0.12 | 0.25 | 0.78 | 0.13 | 0.24 | 1.993 | 9.2 | 8.9 | 9.3 | 8.4 | 119.7 | 119.9 | 0.2 | 0.12 | 0.07 | -0.01 | 0.57 | 0.13 | 0.03 | 1.986 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.8 | CVD |  | 2021-09-16 | 2.003 | 0.99501 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210917 | 0.34 | 0.55 | 0.13 | 0.23 | 0.8 | 0.13 | 0.24 | 1.996 | 9.2 | 9 | 9.1 | 8.6 | 119.6 | 119.9 | 0.28 | 0.2 | 0.08 | 0.05 | 0.58 | 0.09 | 0.07 | 1.985 | 11.6 | 11.7 | 11.8 | 12 | 119.5 | 119.8 | GB |  | 2021-09-17 | 2.003 | 0.99651 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210921 | 0.32 | 0.53 | 0.14 | 0.2 | 0.77 | 0.13 | 0.23 | 1.99 | 9.3 | 8.9 | 9.3 | 8.6 | 119.6 | 119.8 | 0.29 | 0.21 | 0.1 | 0.05 | 0.58 | 0.12 | 0.06 | 1.981 | 11.5 | 11.5 | 11.7 | 11.7 | 119.4 | 119.6 | CVD |  | 2021-09-21 | 2.003 | 0.99351 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210922 | 0.15 | 0.55 | 0.02 | 0.24 | 0.62 | 0.06 | 0.22 | 1.992 | 9.3 | 8.9 | 9.4 | 8.7 | 119.6 | 119.6 | 0.2 | 0.16 | 0.03 | -0.01 | 0.54 | 0.08 | 0.01 | 1.983 | 11.6 | 11.5 | 11.8 | 11.8 | 119.4 | 119.4 | GB |  | 2021-09-22 | 2.003 | 0.99451 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210924 | 0.24 | 0.55 | 0.04 | 0.22 | 0.71 | 0.08 | 0.23 | 1.992 | 9.3 | 8.9 | 9.4 | 8.4 | 119.6 | 119.8 | 0.27 | 0.23 | -0.02 | 0.05 | 0.57 | 0.04 | 0.05 | 1.979 | 11.6 | 11.6 | 11.9 | 12 | 119.5 | 119.6 | ML |  | 2021-09-24 | 2.003 | 0.99451 | 2 | 0.9895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210928 | 0.27 | 0.44 | 0.1 | 0.21 | 0.67 | 0.11 | 0.22 | 1.987 | 8.8 | 9.3 | 9.4 | 8.2 | 119.6 | 119.8 | 0.24 | 0.15 | 0.06 | -0.01 | 0.53 | 0.1 | 0.03 | 1.981 | 11.6 | 11.5 | 11.7 | 11.7 | 119.4 | 119.6 | GB |  | 2021-09-28 | 2.003 | 0.99201 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210929 | 0.25 | 0.39 | 0.06 | 0.19 | 0.66 | 0.09 | 0.23 | 2.016 | 9 | 9.1 | 9 | 8.7 | 119.6 | 119.8 | 0.27 | 0.17 | 0.04 | 0.01 | 0.49 | 0.08 | 0.01 | 2.007 | 11.6 | 11.5 | 11.8 | 11.7 | 119.5 | 119.6 | GB |  | 2021-09-29 | 2.003 | 1.0065 | 2 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210930 | 0.26 | 0.51 | 0.13 | 0.25 | 0.75 | 0.13 | 0.3 | 2.013 | 8.9 | 9.1 | 9.1 | 8.6 | 119.7 | 119.8 | 0.31 | 0.16 | 0.11 | -0.01 | 0.57 | 0.13 | 0.02 | 2.005 | 11.6 | 11.5 | 11.7 | 11.6 | 119.4 | 119.6 | GB |  | 2021-09-30 | 2.003 | 1.005 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211001 | 0.19 | 0.51 | -0.02 | 0.24 | 0.63 | 0 | 0.23 | 2.017 | 8.9 | 9.2 | 9.2 | 8.4 | 119.6 | 119.8 | 0.29 | 0.22 | -0.01 | 0.04 | 0.56 | 0.02 | 0.02 | 2.011 | 11.4 | 11.5 | 11.6 | 11.6 | 119.4 | 119.6 | GB |  | 2021-10-01 | 2.003 | 1.007 | 2 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211004 | 0.23 | 0.42 | 0.12 | 0.27 | 0.71 | 0.15 | 0.24 | 2.018 | 9.5 | 8.5 | 8.9 | 8.8 | 119.6 | 119.9 | 0.24 | 0.2 | 0.04 | -0.1 | 0.53 | 0.09 | -0.05 | 2.007 | 11.6 | 11.7 | 11.7 | 11.8 | 119.5 | 119.8 | GB |  | 2021-10-04 | 2.003 | 1.0075 | 2 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211005 | 0.41 | 0.43 | 0.19 | 0.22 | 0.82 | 0.14 | 0.25 | 2.012 | 9.8 | 7.8 | 9.1 | 8.5 | 119.6 | 120 | 0.37 | 0.14 | 0.13 | -0.01 | 0.63 | 0.16 | 0.02 | 1.998 | 11.7 | 12.1 | 11.8 | 11.9 | 119.5 | 119.8 | GB |  | 2021-10-05 | 2.003 | 1.0045 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211006 | 0.32 | 0.37 | 0.12 | 0.17 | 0.69 | 0.13 | 0.2 | 2.013 | 9.4 | 8.8 | 8.9 | 8.9 | 119.6 | 119.9 | 0.26 | 0.19 | 0.06 | -0.05 | 0.57 | 0.12 | 0 | 2 | 11.6 | 11.6 | 11.7 | 11.8 | 119.4 | 119.7 | GB |  | 2021-10-06 | 2.003 | 1.005 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211007 | 0.36 | 0.37 | 0.15 | 0.18 | 0.62 | 0.14 | 0.14 | 2.012 | 9.4 | 8.7 | 8.7 | 8.9 | 119.6 | 119.9 | 0.29 | 0.12 | 0.07 | -0.02 | 0.55 | 0.11 | -0.05 | 2.005 | 11.5 | 11.4 | 11.5 | 11.5 | 119.5 | 119.7 | GB |  | 2021-10-07 | 2.003 | 1.0045 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211008 | 0.36 | 0.41 | 0.14 | 0.21 | 0.66 | 0.12 | 0.23 | 2.016 | 9.3 | 8.9 | 9 | 8.7 | 119.6 | 119.9 | 0.28 | 0.13 | 0.1 | -0.03 | 0.46 | 0.12 | -0.01 | 2.004 | 11.5 | 11.4 | 11.6 | 11.6 | 119.4 | 119.7 | GB |  | 2021-10-08 | 2.003 | 1.0065 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211012 | 0.21 | 0.42 | 0.09 | 0.23 | 0.67 | 0.11 | 0.25 | 2.008 | 9.5 | 8.6 | 9 | 8.8 | 119.6 | 119.8 | 0.21 | 0.15 | 0.07 | -0.04 | 0.52 | 0.1 | -0.01 | 1.998 | 11.6 | 11.6 | 11.6 | 11.5 | 119.4 | 119.7 | CVD |  | 2021-10-12 | 2.003 | 1.0025 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211013 | 0.33 | 0.54 | 0.16 | 0.24 | 0.74 | 0.16 | 0.23 | 2.012 | 9.3 | 8.9 | 9 | 8.8 | 119.6 | 119.9 | 0.27 | 0.14 | 0.09 | 0 | 0.54 | 0.13 | 0.01 | 2.002 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.7 | CVD |  | 2021-10-13 | 2.003 | 1.0045 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211014 | 0.35 | 0.46 | 0.18 | 0.22 | 0.74 | 0.11 | 0.23 | 2.013 | 8.5 | 9.5 | 9 | 8.6 | 119.6 | 119.9 | 0.25 | 0.16 | 0.08 | -0.01 | 0.52 | 0.12 | 0.01 | 2.002 | 11.8 | 11.6 | 11.7 | 11.8 | 119.5 | 119.8 | CVD |  | 2021-10-14 | 2.003 | 1.005 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211017 | 0.38 | 0.45 | 0.21 | 0.23 | 0.88 | 0.24 | 0.29 | 2.018 | 9 | 9.3 | 8.9 | 9 | 119.6 | 119.8 | 0.38 | 0.18 | 0.16 | -0.02 | 0.67 | 0.18 | 0.01 | 2.011 | 11.5 | 11.6 | 11.6 | 11.5 | 119.3 | 119.5 | EF |  | 2021-10-17 | 2.003 | 1.0075 | 2 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211019 | 0.24 | 0.53 | 0.09 | 0.24 | 0.74 | 0.11 | 0.25 | 2.016 | 8.6 | 9.4 | 9 | 8.7 | 119.6 | 119.8 | 0.28 | 0.19 | 0.09 | 0 | 0.57 | 0.14 | 0.01 | 2.009 | 11.6 | 11.5 | 11.6 | 11.3 | 119.3 | 119.6 | GB |  | 2021-10-19 | 2.003 | 1.0065 | 2 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211020 | 0.35 | 0.43 | 0.16 | 0.22 | 0.77 | 0.17 | 0.26 | 2.017 | 9.3 | 8.9 | 8.9 | 8.8 | 119.6 | 119.8 | 0.28 | 0.17 | 0.09 | -0.05 | 0.6 | 0.13 | 0 | 2.009 | 11.4 | 11.3 | 11.6 | 11.5 | 119.4 | 119.6 | ML |  | 2021-10-20 | 2.003 | 1.007 | 2 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211021 | 0.29 | 0.46 | 0.11 | 0.2 | 0.74 | 0.12 | 0.21 | 2.006 | 9.2 | 8.9 | 9.1 | 8.5 | 119.6 | 119.9 | 0.31 | 0.2 | 0.11 | -0.02 | 0.56 | 0.13 | 0 | 1.995 | 11.4 | 11.3 | 11.6 | 11.5 | 119.5 | 119.7 | GB |  | 2021-10-21 | 2.003 | 1.0015 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211022 | 0.36 | 0.43 | 0.15 | 0.2 | 0.65 | 0.15 | 0.23 | 2.009 | 9.1 | 9.1 | 8.9 | 8.8 | 119.6 | 119.8 | 0.31 | 0.24 | 0.09 | -0.11 | 0.53 | 0.15 | -0.06 | 1.998 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.6 | GB |  | 2021-10-22 | 2.003 | 1.003 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211027 | 0.38 | 0.4 | 0.18 | 0.21 | 0.77 | 0.16 | 0.22 | 2.021 | 9.1 | 9.1 | 9 | 9 | 119.6 | 119.6 | 0.34 | 0.13 | 0.11 | -0.02 | 0.58 | 0.15 | -0.02 | 2.008 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.4 | GB |  | 2021-10-27 | 2.003 | 1.009 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211028 | 0.28 | 0.58 | 0.12 | 0.26 | 0.82 | 0.17 | 0.29 | 2.01 | 8.9 | 9.2 | 9.2 | 8.6 | 119.6 | 119.7 | 0.33 | 0.12 | 0.15 | 0.02 | 0.64 | 0.15 | 0.05 | 2.004 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.4 | GB |  | 2021-10-28 | 2.003 | 1.0035 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211102 | 0.36 | 0.54 | 0.15 | 0.24 | 0.76 | 0.15 | 0.28 | 2.009 | 9.4 | 8.7 | 8.9 | 8.9 | 119.6 | 119.8 | 0.38 | 0.21 | 0.17 | 0.02 | 0.7 | 0.18 | 0.06 | 1.999 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.6 | GB |  | 2021-11-02 | 2.003 | 1.003 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211103 | 0.33 | 0.47 | 0.15 | 0.27 | 0.81 | 0.13 | 0.25 | 2.023 | 9.2 | 8.9 | 9.1 | 8.6 | 119.6 | 119.8 | 0.3 | 0.21 | 0.07 | 0.04 | 0.56 | 0.11 | 0.02 | 2.008 | 11.4 | 11.4 | 11.7 | 11.6 | 119.4 | 119.5 | GB |  | 2021-11-03 | 2.003 | 1.01 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211104 | 0.36 | 0.45 | 0.18 | 0.21 | 0.85 | 0.18 | 0.24 | 1.989 | 9.2 | 9 | 8.8 | 8.9 | 119.6 | 119.8 | 0.28 | 0.22 | 0.09 | 0.02 | 0.57 | 0.12 | 0.03 | 1.978 | 11.4 | 11.5 | 11.5 | 11.2 | 119.4 | 119.6 | ML/ wrong calibration factor was applied |  | 2021-11-04 | 2.003 | 0.99301 | 2 | 0.989 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211110 | 0.42 | 0.48 | 0.23 | 0.28 | 0.8 | 0.21 | 0.29 | 2.011 | 9.8 | 7.8 | 8.9 | 8.8 | 119.5 | 119.8 | 0.39 | 0.21 | 0.16 | 0 | 0.63 | 0.17 | 0.03 | 1.999 | 11.6 | 11.8 | 11.5 | 11.2 | 119.3 | 119.6 | GB |  | 2021-11-10 | 2.003 | 1.004 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211111 | 0.29 | 0.49 | 0.07 | 0.2 | 0.7 | 0.08 | 0.21 | 2.018 | 8.8 | 9.3 | 9.2 | 8.6 | 119.6 | 119.8 | 0.3 | 0.18 | 0.11 | 0 | 0.62 | 0.14 | 0.03 | 2.008 | 11.7 | 11.6 | 11.6 | 11.4 | 119.4 | 119.5 | ML |  | 2021-11-11 | 2.003 | 1.0075 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211112 | 0.29 | 0.5 | 0.16 | 0.24 | 0.84 | 0.17 | 0.27 | 2.01 | 8.7 | 9.4 | 9.2 | 8.5 | 119.6 | 119.8 | 0.37 | 0.15 | 0.14 | -0.01 | 0.56 | 0.16 | 0.02 | 2.001 | 11.7 | 11.6 | 11.6 | 11.5 | 119.3 | 119.6 | ML |  | 2021-11-12 | 2.003 | 1.0035 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211116 | 0.31 | 0.46 | 0.16 | 0.23 | 0.85 | 0.19 | 0.29 | 1.992 | 9.1 | 9.1 | 9.1 | 8.7 | 119.6 | 119.8 | 0.29 | 0.19 | 0.1 | -0.02 | 0.59 | 0.15 | 0.04 | 1.981 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.6 | GB |  | 2021-11-16 | 2.003 | 0.99451 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211117 | 0.34 | 0.46 | 0.14 | 0.27 | 0.69 | 0.14 | 0.24 | 2.001 | 9.3 | 8.8 | 9.1 | 8.6 | 119.6 | 119.8 | 0.27 | 0.15 | 0.07 | -0.01 | 0.6 | 0.12 | 0.03 | 1.99 | 11.5 | 11.5 | 11.7 | 11.7 | 119.4 | 119.6 | GB |  | 2021-11-17 | 2.003 | 0.999 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211118 | 0.33 | 0.44 | 0.14 | 0.2 | 0.83 | 0.16 | 0.23 | 1.997 | 9.1 | 9 | 9.3 | 8.4 | 119.6 | 119.8 | 0.36 | 0.13 | 0.15 | -0.03 | 0.64 | 0.17 | 0.02 | 1.984 | 11.5 | 11.5 | 11.7 | 11.8 | 119.4 | 119.6 | ML |  | 2021-11-18 | 2.003 | 0.997 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211119 | 0.39 | 0.45 | 0.17 | 0.22 | 0.83 | 0.18 | 0.23 | 2.009 | 9.4 | 8.7 | 8.9 | 8.8 | 119.6 | 119.8 | 0.27 | 0.16 | 0.11 | -0.01 | 0.63 | 0.16 | 0.01 | 1.994 | 11.5 | 11.3 | 11.5 | 11.3 | 119.4 | 119.6 | ML |  | 2021-11-19 | 2.003 | 1.003 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211123 | 0.38 | 0.45 | 0.17 | 0.27 | 0.77 | 0.15 | 0.27 | 2.002 | 9.5 | 8.5 | 9 | 8.9 | 119.6 | 119.8 | 0.29 | 0.18 | 0.09 | 0.03 | 0.64 | 0.13 | 0.03 | 1.992 | 11.5 | 11.5 | 11.5 | 11.2 | 119.4 | 119.5 | ML |  | 2021-11-23 | 2.003 | 0.9995 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211124 | 0.36 | 0.47 | 0.13 | 0.26 | 0.8 | 0.16 | 0.26 | 2.003 | 9.3 | 8.9 | 9 | 8.8 | 119.6 | 119.8 | 0.31 | 0.12 | 0.12 | -0.05 | 0.55 | 0.15 | -0.02 | 1.991 | 11.5 | 11.5 | 11.5 | 11.2 | 119.4 | 119.6 | ML |  | 2021-11-24 | 2.003 | 1 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211125 | 0.37 | 0.44 | 0.2 | 0.22 | 0.87 | 0.2 | 0.25 | 2.01 | 9.6 | 8.3 | 8.9 | 8.8 | 119.6 | 119.8 | 0.26 | 0.2 | 0.12 | 0.05 | 0.59 | 0.13 | 0.02 | 1.998 | 11.6 | 11.8 | 11.7 | 11.6 | 119.4 | 119.6 | CVD |  | 2021-11-25 | 2.003 | 1.0035 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211130 | 0.37 | 0.49 | 0.06 | 0.3 | 0.8 | 0.06 | 0.24 | 2.011 | 9.2 | 8.7 | 8.6 | 8.9 | 119.6 | 119.8 | 0.23 | 0.18 | 0.08 | 0.08 | 0.34 | 0.06 | 0.02 | 1.997 | 11.4 | 11.3 | 11.3 | 11.2 | 119.4 | 119.6 | GB |  | 2021-11-30 | 2.003 | 1.004 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211201 | 0.49 | 0.38 | 0.22 | 0.17 | 0.76 | 0.2 | 0.19 | 2.016 | 9.5 | 8.5 | 8.8 | 8.8 | 119.6 | 119.8 | 0.26 | 0.12 | 0.06 | -0.03 | 0.49 | 0.11 | -0.01 | 2.005 | 11.5 | 11.4 | 11.5 | 11.1 | 119.4 | 119.6 | GB |  | 2021-12-01 | 2.003 | 1.0065 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211202 | 0.37 | 0.4 | 0.12 | 0.17 | 0.76 | 0.16 | 0.19 | 2.002 | 9.4 | 8.7 | 8.9 | 8.9 | 119.6 | 119.8 | 0.25 | 0.13 | 0.07 | -0.04 | 0.49 | 0.1 | -0.01 | 1.988 | 11.6 | 11.5 | 11.6 | 11.5 | 119.4 | 119.5 | GB |  | 2021-12-02 | 2.003 | 0.9995 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211203 | 0.22 | 0.49 | 0.08 | 0.26 | 0.67 | 0.12 | 0.25 | 2.005 | 8.8 | 9.3 | 9.1 | 8.6 | 119.6 | 119.7 | 0.3 | 0.19 | 0.1 | 0.03 | 0.58 | 0.14 | 0.02 | 1.993 | 11.7 | 11.6 | 11.6 | 11.4 | 119.3 | 119.5 | KA |  | 2021-12-03 | 2.003 | 1.001 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211208 | 0.28 | 0.47 | 0.11 | 0.23 | 0.74 | 0.13 | 0.25 | 2.001 | 9.3 | 8.8 | 9 | 8.9 | 119.6 | 119.7 | 0.27 | 0.15 | 0.07 | 0.01 | 0.52 | 0.1 | 0 | 1.991 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.5 | GB |  | 2021-12-08 | 2.003 | 0.999 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211209 | 0.33 | 0.46 | 0.12 | 0.22 | 0.71 | 0.11 | 0.22 | 2.005 | 9.5 | 8.6 | 8.8 | 9 | 119.6 | 119.7 | 0.22 | 0.17 | 0.07 | 0.02 | 0.57 | 0.12 | 0.02 | 1.993 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.5 | ML |  | 2021-12-09 | 2.003 | 1.001 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211210 | 0.3 | 0.47 | 0.13 | 0.23 | 0.71 | 0.12 | 0.23 | 2.005 | 9.2 | 9 | 9.2 | 8.5 | 119.6 | 119.7 | 0.28 | 0.18 | 0.09 | -0.06 | 0.57 | 0.12 | -0.01 | 1.996 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.5 | ML |  | 2021-12-10 | 2.003 | 1.001 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211213 | 0.27 | 0.49 | 0.12 | 0.25 | 0.74 | 0.15 | 0.28 | 2.005 | 9.4 | 8.7 | 9 | 8.8 | 119.6 | 119.7 | 0.34 | 0.23 | 0.13 | 0 | 0.59 | 0.16 | 0 | 1.992 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.5 | ML |  | 2021-12-13 | 2.003 | 1.001 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211215 | 0.35 | 0.53 | 0.13 | 0.3 | 0.79 | 0.17 | 0.28 | 2 | 9.5 | 8.5 | 9.2 | 8.5 | 119.6 | 119.8 | 0.29 | 0.18 | 0.11 | 0 | 0.57 | 0.15 | 0.03 | 1.993 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.5 | ML |  | 2021-12-15 | 2.003 | 0.9985 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211217 | 0.32 | 0.49 | 0.14 | 0.25 | 0.78 | 0.14 | 0.24 | 1.999 | 9.2 | 9 | 9.2 | 8.6 | 119.6 | 119.8 | 0.33 | 0.18 | 0.11 | 0.02 | 0.62 | 0.15 | 0.04 | 1.988 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.5 | ML |  | 2021-12-17 | 2.003 | 0.998 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211221 | 0.4 | 0.44 | 0.17 | 0.25 | 0.75 | 0.17 | 0.25 | 2.004 | 9.2 | 9 | 9.1 | 8.6 | 119.6 | 119.8 | 0.3 | 0.11 | 0.1 | -0.01 | 0.57 | 0.12 | 0.01 | 1.993 | 11.5 | 11.4 | 11.6 | 11.3 | 119.4 | 119.5 | CVD |  | 2021-12-21 | 2.003 | 1.0005 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211223 | 0.26 | 0.45 | 0.07 | 0.26 | 0.74 | 0.1 | 0.24 | 2.011 | 9.1 | 9.1 | 9.1 | 8.7 | 119.6 | 119.7 | 0.27 | 0.17 | 0.07 | 0.03 | 0.56 | 0.11 | 0.01 | 2.001 | 11.4 | 11.4 | 11.6 | 11.3 | 119.4 | 119.5 | LL |  | 2021-12-23 | 2.003 | 1.004 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220104 | 0.44 | 0.78 | 0.22 | 0.53 | 1.04 | 0.23 | 0.43 | 2.028 | 9.4 | 9 | 9.2 | 8.9 | 119.7 | 119.8 | 0.35 | 0.31 | 0.15 | 0.17 | 0.71 | 0.16 | 0.15 | 2.007 | 11.6 | 11.6 | 11.7 | 11.7 | 119.5 | 119.6 | GB |  | 2022-01-04 | 2.003 | 1.0125 | 2 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220105 | 0.32 | 0.43 | 0.1 | 0.2 | 0.78 | 0.11 | 0.23 | 1.998 | 9.3 | 8.8 | 9.2 | 8.6 | 119.6 | 119.7 | 0.3 | 0.13 | 0.1 | -0.02 | 0.53 | 0.13 | 0 | 1.984 | 11.5 | 11.5 | 11.7 | 11.5 | 119.4 | 119.5 | GB, kört med 0,8 grader lägre temperatur i dosimetry manager. |  | 2022-01-05 | 2.003 | 0.9975 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220106 | 0.25 | 0.55 | 0 | 0.25 | 0.65 | 0.06 | 0.23 | 2.002 | 9.2 | 8.9 | 9.2 | 8.5 | 119.7 | 119.8 | 0.26 | 0.17 | 0.02 | -0.01 | 0.59 | 0.1 | 0.02 | 1.987 | 11.6 | 11.5 | 11.8 | 11.9 | 119.6 | 119.6 | GB |  | 2022-01-06 | 2.003 | 0.9995 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220107 | 0.31 | 0.52 | 0.1 | 0.28 | 0.8 | 0.13 | 0.25 | 2.006 | 9.3 | 9 | 9 | 8.9 | 119.6 | 119.7 | 0.37 | 0.21 | 0.13 | 0.06 | 0.59 | 0.14 | 0.05 | 1.992 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.5 | LL |  | 2022-01-07 | 2.003 | 1.0015 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220111 | 0.32 | 0.36 | 0.13 | 0.21 | 0.7 | 0.16 | 0.22 | 2.014 | 9.3 | 9 | 9.2 | 8.7 | 119.6 | 119.7 | 0.28 | 0.13 | 0.08 | -0.01 | 0.54 | 0.11 | 0 | 2.002 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.5 | ML |  | 2022-01-11 | 2.003 | 1.0055 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220112 | 0.19 | 0.52 | 0.04 | 0.19 | 0.83 | 0.1 | 0.2 | 2.013 | 9.4 | 8.7 | 9.2 | 8.6 | 119.6 | 119.8 | 0.28 | 0.27 | 0.08 | 0.04 | 0.62 | 0.11 | 0.03 | 1.998 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.5 | GB |  | 2022-01-12 | 2.003 | 1.005 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220113 | 0.3 | 0.49 | 0.1 | 0.23 | 0.77 | 0.14 | 0.26 | 2.006 | 9.3 | 8.8 | 9.1 | 8.8 | 119.6 | 119.7 | 0.34 | 0.2 | 0.13 | -0.06 | 0.56 | 0.16 | 0 | 1.993 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.5 | ML |  | 2022-01-13 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220114 | 0.35 | 0.45 | 0.12 | 0.25 | 0.84 | 0.15 | 0.22 | 2.01 | 9.1 | 9.1 | 9.2 | 8.6 | 119.6 | 119.8 | 0.31 | 0.16 | 0.11 | 0.02 | 0.58 | 0.14 | 0 | 1.998 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | GB |  | 2022-01-14 | 2.003 | 1.0035 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220118 | 0.37 | 0.43 | 0.12 | 0.23 | 0.72 | 0.14 | 0.2 | 2.01 | 9.1 | 9.1 | 9.2 | 8.7 | 119.6 | 119.8 | 0.23 | 0.15 | 0.06 | 0.01 | 0.57 | 0.11 | -0.01 | 1.995 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.6 | ML |  | 2022-01-18 | 2.003 | 1.0035 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220119 | 0.32 | 0.46 | 0.14 | 0.2 | 0.73 | 0.15 | 0.23 | 2.008 | 9.1 | 9.1 | 9.1 | 8.7 | 119.6 | 119.8 | 0.32 | 0.24 | 0.1 | -0.08 | 0.58 | 0.14 | -0.03 | 1.998 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | GB |  | 2022-01-19 | 2.003 | 1.0025 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220120 | 0.35 | 0.52 | 0.15 | 0.28 | 0.97 | 0.21 | 0.3 | 2.001 | 9.3 | 8.8 | 9.1 | 8.7 | 119.6 | 119.8 | 0.32 | 0.15 | 0.13 | 0 | 0.68 | 0.16 | 0.04 | 1.985 | 11.6 | 11.7 | 11.7 | 11.8 | 119.5 | 119.6 | ML |  | 2022-01-20 | 2.003 | 0.999 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220121 | 0.33 | 0.46 | 0.09 | 0.21 | 0.77 | 0.15 | 0.2 | 2.006 | 9.3 | 9 | 8.9 | 9 | 119.6 | 119.7 | 0.29 | 0.17 | 0.1 | 0.02 | 0.54 | 0.11 | 0.01 | 1.991 | 11.6 | 11.5 | 11.6 | 11.3 | 119.4 | 119.5 | ML |  | 2022-01-21 | 2.003 | 1.0015 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220125 | 0.26 | 0.51 | 0.09 | 0.27 | 0.71 | 0.1 | 0.25 | 1.998 | 9.2 | 9 | 9.2 | 8.5 | 119.6 | 119.8 | 0.29 | 0.2 | 0.06 | 0.01 | 0.58 | 0.12 | 0.03 | 1.986 | 11.7 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | CVD |  | 2022-01-25 | 2.003 | 0.9975 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220126 | 0.26 | 0.46 | 0.06 | 0.19 | 0.7 | 0.1 | 0.24 | 2.003 | 9.1 | 9.1 | 9 | 8.9 | 119.6 | 119.7 | 0.26 | 0.16 | 0.07 | -0.06 | 0.53 | 0.14 | -0.01 | 1.986 | 11.5 | 11.5 | 11.7 | 11.7 | 119.5 | 119.6 | ML |  | 2022-01-26 | 2.003 | 1 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220128 | 0.4 | 0.5 | 0.21 | 0.24 | 0.97 | 0.25 | 0.3 | 2.005 | 9.1 | 9.2 | 9.1 | 8.8 | 119.6 | 119.8 | 0.36 | 0.17 | 0.15 | 0 | 0.69 | 0.18 | 0.04 | 1.987 | 11.6 | 11.6 | 11.7 | 11.8 | 119.5 | 119.6 | ML |  | 2022-01-28 | 2.003 | 1.001 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220201 | 0.39 | 0.53 | 0.17 | 0.2 | 0.93 | 0.19 | 0.25 | 1.998 | 9.3 | 8.9 | 9 | 8.9 | 119.6 | 119.7 | 0.3 | 0.16 | 0.12 | -0.01 | 0.57 | 0.14 | 0 | 1.983 | 11.5 | 11.6 | 11.6 | 11.4 | 119.4 | 119.5 | CVD |  | 2022-02-01 | 2.003 | 0.9975 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220202 | 0.37 | 0.44 | 0.1 | 0.17 | 0.84 | 0.14 | 0.2 | 2.007 | 9.4 | 8.7 | 9 | 8.9 | 119.6 | 119.7 | 0.28 | 0.14 | 0.08 | -0.05 | 0.53 | 0.12 | -0.01 | 1.99 | 11.6 | 11.7 | 11.7 | 11.8 | 119.4 | 119.5 | ML |  | 2022-02-02 | 2.003 | 1.002 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220203 | 0.42 | 0.54 | 0.2 | 0.25 | 0.94 | 0.2 | 0.25 | 2.005 | 9.4 | 8.8 | 9.2 | 8.8 | 119.6 | 119.7 | 0.36 | 0.19 | 0.12 | 0.03 | 0.57 | 0.12 | 0.02 | 1.991 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.5 | ML |  | 2022-02-03 | 2.003 | 1.001 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220204 | 0.34 | 0.51 | 0.13 | 0.25 | 0.83 | 0.16 | 0.29 | 2.008 | 9.2 | 9 | 9.1 | 8.7 | 119.7 | 119.8 | 0.26 | 0.15 | 0.06 | -0.04 | 0.55 | 0.12 | 0.02 | 1.993 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.6 | GB |  | 2022-02-04 | 2.003 | 1.0025 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220208 | 0.38 | 0.73 | 0.22 | 0.31 | 1.03 | 0.18 | 0.3 | 2.007 | 9.6 | 8.4 | 8.5 | 9.3 | 119.6 | 119.7 | 0.24 | 0.15 | 0.03 | 0 | 0.54 | 0.08 | 0.02 | 1.993 | 11.7 | 11.9 | 11.6 | 11.3 | 119.4 | 119.6 | LL |  | 2022-02-08 | 2.003 | 1.002 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220209 | 0.43 | 0.49 | 0.17 | 0.28 | 0.88 | 0.19 | 0.24 | 2.02 | 9.6 | 8.5 | 9.1 | 8.7 | 119.6 | 119.8 | 0.3 | 0.13 | 0.1 | -0.05 | 0.51 | 0.14 | -0.01 | 2.005 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.6 | ML |  | 2022-02-09 | 2.003 | 1.0085 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220215 | 0.39 | 0.5 | 0.15 | 0.23 | 0.85 | 0.19 | 0.25 | 1.995 | 9.4 | 8.8 | 9.2 | 8.7 | 119.6 | 119.8 | 0.32 | 0.17 | 0.1 | 0.03 | 0.56 | 0.12 | 0.04 | 1.984 | 11.6 | 11.5 | 11.7 | 11.6 | 119.4 | 119.5 | ML |  | 2022-02-15 | 2.003 | 0.99601 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220216 | 0.51 | 0.43 | 0.21 | 0.2 | 0.95 | 0.21 | 0.19 | 1.996 | 9.4 | 8.9 | 9.2 | 8.5 | 119.6 | 119.8 | 0.29 | 0.2 | 0.06 | 0.05 | 0.58 | 0.11 | 0.02 | 1.983 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.5 | GB |  | 2022-02-16 | 2.003 | 0.99651 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220217 | 0.33 | 0.43 | 0.09 | 0.22 | 0.75 | 0.16 | 0.23 | 2.003 | 8.9 | 9.2 | 9 | 8.7 | 119.7 | 119.8 | 0.29 | 0.24 | -0.02 | -0.05 | 0.5 | 0.07 | -0.04 | 1.989 | 11.5 | 11.5 | 11.5 | 11.2 | 119.5 | 119.6 | ML |  | 2022-02-17 | 2.003 | 1 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220218 | 0.37 | 0.49 | 0.16 | 0.21 | 0.85 | 0.18 | 0.2 | 2.002 | 9 | 9.3 | 9 | 8.9 | 119.6 | 119.7 | 0.27 | 0.14 | 0.07 | -0.01 | 0.56 | 0.12 | 0.01 | 1.988 | 11.8 | 11.6 | 11.6 | 11.4 | 119.4 | 119.5 | GB |  | 2022-02-18 | 2.003 | 0.9995 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220222 | 0.39 | 0.44 | 0.11 | 0.24 | 0.88 | 0.17 | 0.19 | 2.007 | 9.3 | 8.8 | 9.2 | 8.5 | 119.6 | 119.8 | 0.31 | 0.19 | 0.09 | 0.03 | 0.59 | 0.13 | 0.02 | 1.993 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.6 | CVD |  | 2022-02-22 | 2.003 | 1.002 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220223 | 0.45 | 0.44 | 0.22 | 0.25 | 0.84 | 0.21 | 0.22 | 2 | 9.1 | 9 | 9.2 | 8.4 | 119.7 | 119.9 | 0.31 | 0.16 | 0.1 | 0 | 0.61 | 0.15 | 0.03 | 1.992 | 11.9 | 11.7 | 11.7 | 11.8 | 119.5 | 119.6 | ML |  | 2022-02-23 | 2.003 | 0.9985 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220225 | 0.45 | 0.48 | 0.17 | 0.19 | 0.88 | 0.21 | 0.21 | 2.007 | 9.3 | 9 | 8.9 | 9.1 | 119.6 | 119.7 | 0.3 | 0.15 | 0.08 | -0.02 | 0.51 | 0.12 | 0.03 | 1.992 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | KA |  | 2022-02-25 | 2.003 | 1.002 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220301 | 0.4 | 0.61 | 0.15 | 0.28 | 0.93 | 0.15 | 0.34 | 2.007 | 9.1 | 9.1 | 9.3 | 8.7 | 119.6 | 119.7 | 0.33 | 0.16 | 0.12 | -0.03 | 0.6 | 0.15 | 0.04 | 1.995 | 11.6 | 11.6 | 11.7 | 11.7 | 119.4 | 119.5 | ML |  | 2022-03-01 | 2.003 | 1.002 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220302 | 0.46 | 0.4 | 0.2 | 0.17 | 0.75 | 0.17 | 0.17 | 2.008 | 9.3 | 9 | 9.2 | 8.6 | 119.6 | 119.8 | 0.29 | 0.12 | 0.09 | -0.03 | 0.57 | 0.12 | 0.02 | 1.993 | 11.5 | 11.5 | 11.8 | 12 | 119.5 | 119.5 | ML |  | 2022-03-02 | 2.003 | 1.0025 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220303 | 0.54 | 0.43 | 0.27 | 0.24 | 1.03 | 0.27 | 0.31 | 2.004 | 9 | 9.3 | 9.3 | 8.6 | 119.5 | 119.7 | 0.48 | 0.24 | 0.18 | 0.07 | 0.88 | 0.22 | 0.1 | 1.994 | 11.7 | 11.6 | 11.7 | 11.6 | 119.3 | 119.4 | ML |  | 2022-03-03 | 2.003 | 1.0005 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220304 | 0.49 | 0.46 | 0.2 | 0.21 | 0.88 | 0.21 | 0.21 | 2.013 | 9.3 | 8.9 | 9.1 | 8.7 | 119.6 | 119.8 | 0.29 | 0.17 | 0.06 | 0.03 | 0.58 | 0.09 | 0.01 | 2 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.5 | GB |  | 2022-03-04 | 2.003 | 1.005 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220308 | 0.53 | 0.41 | 0.21 | 0.19 | 0.88 | 0.22 | 0.22 | 2.011 | 9.8 | 8.1 | 8.8 | 9.1 | 119.6 | 119.8 | 0.27 | 0.12 | 0.06 | -0.01 | 0.52 | 0.12 | 0.01 | 1.999 | 11.9 | 11.7 | 11.7 | 11.5 | 119.4 | 119.5 | CVD |  | 2022-03-08 | 2.003 | 1.004 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220309 | 0.38 | 0.47 | 0.14 | 0.25 | 0.81 | 0.17 | 0.24 | 2.013 | 9.5 | 8.7 | 9.1 | 8.7 | 119.6 | 119.8 | 0.25 | 0.16 | 0.07 | 0.01 | 0.51 | 0.1 | 0.03 | 2 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | ML |  | 2022-03-09 | 2.003 | 1.005 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220310 | 0.41 | 0.49 | 0.18 | 0.26 | 0.84 | 0.19 | 0.24 | 2.007 | 9.2 | 9 | 9.1 | 8.6 | 119.7 | 119.8 | 0.33 | 0.16 | 0.1 | 0.01 | 0.6 | 0.13 | 0.03 | 1.995 | 11.5 | 11.6 | 11.7 | 11.6 | 119.4 | 119.6 | CVD |  | 2022-03-10 | 2.003 | 1.002 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220314 | 0.32 | 0.52 | 0.12 | 0.23 | 0.75 | 0.17 | 0.23 | 2.016 | 9.3 | 8.9 | 9.2 | 8.6 | 119.6 | 119.8 | 0.24 | 0.18 | 0.07 | -0.01 | 0.52 | 0.11 | 0 | 2.001 | 11.5 | 11.5 | 11.7 | 11.5 | 119.5 | 119.6 | GB |  | 2022-03-14 | 2.003 | 1.0065 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220316 | 0.5 | 0.41 | 0.2 | 0.23 | 0.92 | 0.22 | 0.24 | 2.008 | 9.6 | 8.5 | 9 | 8.9 | 119.6 | 119.8 | 0.35 | 0.16 | 0.08 | 0.01 | 0.58 | 0.13 | 0.02 | 1.997 | 11.6 | 11.7 | 11.7 | 11.5 | 119.4 | 119.6 | GB |  | 2022-03-16 | 2.003 | 1.0025 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220317 | 0.46 | 0.44 | 0.25 | 0.26 | 0.9 | 0.22 | 0.25 | 2.014 | 9.7 | 8.4 | 9.1 | 8.9 | 119.6 | 119.7 | 0.27 | 0.16 | 0.05 | -0.04 | 0.5 | 0.11 | 0.01 | 2.001 | 11.7 | 11.8 | 11.7 | 11.6 | 119.4 | 119.5 | ML |  | 2022-03-17 | 2.003 | 1.0055 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220318 | 0.42 | 0.49 | 0.17 | 0.23 | 0.93 | 0.19 | 0.2 | 2.016 | 9 | 9.2 | 9.4 | 8.3 | 119.6 | 119.8 | 0.28 | 0.18 | 0.08 | 0.03 | 0.64 | 0.13 | 0.06 | 2.004 | 11.7 | 11.6 | 11.7 | 11.7 | 119.4 | 119.5 | GB |  | 2022-03-18 | 2.003 | 1.0065 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220323 | 0.46 | 0.43 | 0.2 | 0.23 | 0.85 | 0.19 | 0.23 | 2.017 | 9.5 | 8.6 | 8.9 | 9 | 119.6 | 119.8 | 0.28 | 0.18 | 0.06 | -0.02 | 0.61 | 0.12 | 0.02 | 2.004 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.6 | ML |  | 2022-03-23 | 2.003 | 1.007 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220324 | 0.45 | 0.45 | 0.19 | 0.23 | 0.89 | 0.21 | 0.25 | 2.008 | 9.5 | 8.7 | 9.1 | 8.8 | 119.6 | 119.7 | 0.31 | 0.1 | 0.11 | -0.03 | 0.54 | 0.13 | 0 | 1.994 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | CVD |  | 2022-03-24 | 2.003 | 1.0025 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220325 | 0.43 | 0.47 | 0.2 | 0.2 | 0.99 | 0.25 | 0.27 | 2.008 | 9.3 | 8.9 | 9.1 | 8.8 | 119.6 | 119.7 | 0.38 | 0.17 | 0.15 | -0.01 | 0.73 | 0.2 | 0.05 | 1.994 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | ML |  | 2022-03-25 | 2.003 | 1.0025 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220329 | 0.52 | 0.49 | 0.26 | 0.22 | 1.01 | 0.22 | 0.21 | 2.005 | 9.1 | 9.1 | 9.3 | 8.5 | 119.6 | 119.7 | 0.28 | 0.15 | 0.08 | -0.02 | 0.59 | 0.11 | 0.02 | 1.995 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.5 | GB |  | 2022-03-29 | 2.003 | 1.001 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220330 | 0.45 | 0.46 | 0.2 | 0.24 | 0.93 | 0.22 | 0.23 | 2.014 | 9.5 | 8.6 | 9 | 8.8 | 119.6 | 119.8 | 0.31 | 0.14 | 0.1 | -0.04 | 0.57 | 0.12 | 0 | 2.001 | 11.5 | 11.6 | 11.6 | 11.3 | 119.4 | 119.6 | GB |  | 2022-03-30 | 2.003 | 1.0055 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220401 | 0.49 | 0.39 | 0.19 | 0.24 | 0.91 | 0.22 | 0.24 | 2.008 | 9.5 | 8.6 | 9 | 8.9 | 119.6 | 119.8 | 0.33 | 0.13 | 0.1 | -0.05 | 0.51 | 0.11 | -0.01 | 1.995 | 11.6 | 11.6 | 11.6 | 11.3 | 119.4 | 119.6 | ML |  | 2022-04-01 | 2.003 | 1.0025 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220405 | 0.46 | 0.42 | 0.2 | 0.23 | 0.86 | 0.21 | 0.24 | 2.013 | 9.6 | 8.5 | 9 | 8.8 | 119.6 | 119.8 | 0.32 | 0.18 | 0.09 | 0.02 | 0.56 | 0.12 | 0.02 | 1.998 | 11.6 | 11.5 | 11.6 | 11.4 | 119.4 | 119.5 | ML |  | 2022-04-05 | 2.003 | 1.005 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220406 | 0.36 | 0.53 | 0.15 | 0.23 | 0.88 | 0.19 | 0.26 | 2 | 9.1 | 8.9 | 9 | 8.6 | 119.8 | 119.9 | 0.31 | 0.17 | 0.09 | -0.05 | 0.58 | 0.14 | 0.01 | 1.987 | 11.6 | 11.5 | 11.6 | 11.5 | 119.6 | 119.7 | ML |  | 2022-04-06 | 2.003 | 0.9985 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220407 | 0.42 | 0.42 | 0.15 | 0.21 | 0.84 | 0.17 | 0.21 | 2.009 | 9.2 | 9 | 9.1 | 8.8 | 119.6 | 119.7 | 0.32 | 0.12 | 0.12 | 0.01 | 0.55 | 0.13 | 0.02 | 1.996 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.5 | ML |  | 2022-04-07 | 2.003 | 1.003 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220408 | 0.46 | 0.4 | 0.25 | 0.18 | 0.9 | 0.23 | 0.23 | 2.002 | 9.2 | 9.1 | 8.9 | 9 | 119.6 | 119.7 | 0.31 | 0.16 | 0.1 | 0.01 | 0.6 | 0.13 | 0.05 | 1.989 | 11.6 | 11.6 | 11.9 | 11.9 | 119.5 | 119.5 | GB |  | 2022-04-08 | 2.003 | 0.9995 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220412 | 0.46 | 0.44 | 0.16 | 0.21 | 0.85 | 0.19 | 0.24 | 2.005 | 9.4 | 8.8 | 8.9 | 8.9 | 119.6 | 119.8 | 0.34 | 0.13 | 0.1 | -0.02 | 0.6 | 0.14 | 0.02 | 1.993 | 11.5 | 11.5 | 11.7 | 11.5 | 119.4 | 119.6 | ML |  | 2022-04-12 | 2.003 | 1.001 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220413 | 0.44 | 0.45 | 0.17 | 0.21 | 0.93 | 0.18 | 0.25 | 2.008 | 9.3 | 9 | 9.2 | 8.8 | 119.6 | 119.7 | 0.25 | 0.1 | 0.05 | -0.04 | 0.53 | 0.1 | 0.01 | 1.992 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.5 | ML |  | 2022-04-13 | 2.003 | 1.0025 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220414 | 0.42 | 0.48 | 0.16 | 0.22 | 0.94 | 0.2 | 0.26 | 2.008 | 9.1 | 9.1 | 9 | 8.9 | 119.6 | 119.7 | 0.26 | 0.14 | 0.09 | 0.01 | 0.63 | 0.14 | 0.04 | 1.994 | 11.7 | 11.6 | 11.8 | 11.7 | 119.4 | 119.5 | LL |  | 2022-04-14 | 2.003 | 1.0025 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220420 | 0.37 | 0.42 | 0.15 | 0.21 | 0.91 | 0.18 | 0.21 | 2.007 | 9.1 | 9.2 | 9.3 | 8.6 | 119.6 | 119.8 | 0.29 | 0.17 | 0.08 | -0.01 | 0.61 | 0.12 | 0.02 | 1.995 | 11.7 | 11.6 | 11.7 | 11.7 | 119.4 | 119.5 | AD |  | 2022-04-20 | 2.003 | 1.002 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220421 | 0.4 | 0.56 | 0.2 | 0.23 | 0.98 | 0.21 | 0.22 | 2.017 | 8.8 | 9.3 | 9.2 | 8.5 | 119.7 | 119.8 | 0.3 | 0.11 | 0.1 | -0.02 | 0.58 | 0.14 | 0.02 | 2.004 | 11.8 | 11.6 | 11.7 | 11.6 | 119.5 | 119.6 | AD |  | 2022-04-21 | 2.003 | 1.007 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220422 | 0.41 | 0.44 | 0.17 | 0.25 | 0.8 | 0.17 | 0.22 | 2.006 | 9.3 | 9 | 9 | 8.9 | 119.6 | 119.7 | 0.26 | 0.2 | 0.05 | 0.03 | 0.54 | 0.11 | 0.04 | 1.993 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.5 | GB |  | 2022-04-22 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220426 | 0.41 | 0.46 | 0.19 | 0.22 | 0.91 | 0.21 | 0.22 | 2.003 | 9.2 | 9.1 | 9.1 | 8.8 | 119.6 | 119.8 | 0.32 | 0.11 | 0.11 | -0.02 | 0.58 | 0.13 | 0.01 | 1.99 | 11.5 | 11.6 | 11.6 | 11.5 | 119.4 | 119.6 | EF |  | 2022-04-26 | 2.003 | 1 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220503 | 0.45 | 0.5 | 0.19 | 0.29 | 0.78 | 0.19 | 0.23 | 2.017 | 9.7 | 8.3 | 9 | 8.9 | 119.6 | 119.8 | 0.3 | 0.13 | 0.02 | -0.02 | 0.45 | 0.09 | 0.01 | 2.005 | 11.7 | 11.8 | 11.6 | 11.3 | 119.4 | 119.6 | GB |  | 2022-05-03 | 2.003 | 1.007 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220504 | 0.38 | 0.51 | 0.17 | 0.27 | 0.82 | 0.17 | 0.21 | 2.01 | 9.2 | 9 | 9.3 | 8.5 | 119.7 | 119.8 | 0.3 | 0.19 | 0.08 | 0.03 | 0.55 | 0.13 | 0.05 | 1.999 | 11.6 | 11.6 | 11.7 | 11.6 | 119.4 | 119.6 | GB |  | 2022-05-04 | 2.003 | 1.0035 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220506 | 0.41 | 0.39 | 0.14 | 0.22 | 0.82 | 0.18 | 0.25 | 2.008 | 9.1 | 9.1 | 9 | 8.6 | 119.7 | 119.9 | 0.29 | 0.19 | 0.07 | -0.01 | 0.54 | 0.12 | 0.01 | 1.994 | 11.5 | 11.5 | 11.7 | 11.8 | 119.6 | 119.7 | LL |  | 2022-05-06 | 2.003 | 1.0025 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220510 | 0.42 | 0.49 | 0.18 | 0.25 | 0.85 | 0.17 | 0.25 | 2.008 | 9.2 | 9 | 9.1 | 8.6 | 119.7 | 119.9 | 0.28 | 0.17 | 0.06 | -0.03 | 0.57 | 0.12 | 0.02 | 1.996 | 11.6 | 11.6 | 11.7 | 11.7 | 119.4 | 119.7 | LL |  | 2022-05-10 | 2.003 | 1.0025 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220511 | 0.41 | 0.48 | 0.16 | 0.21 | 0.79 | 0.18 | 0.23 | 1.994 | 9.2 | 9 | 9.1 | 8.7 | 119.6 | 119.9 | 0.27 | 0.14 | 0.07 | -0.04 | 0.49 | 0.11 | -0.01 | 1.98 | 11.5 | 11.5 | 11.8 | 11.9 | 119.5 | 119.7 | LL |  | 2022-05-11 | 2.003 | 0.99551 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220512 | 0.35 | 0.51 | 0.13 | 0.26 | 0.85 | 0.18 | 0.21 | 1.992 | 8.9 | 9.2 | 9.2 | 8.5 | 119.7 | 119.8 | 0.22 | 0.15 | 0.03 | 0 | 0.52 | 0.11 | 0.02 | 1.98 | 11.6 | 11.5 | 11.8 | 11.8 | 119.5 | 119.6 | ML |  | 2022-05-12 | 2.003 | 0.99451 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220513 | 0.55 | 0.43 | 0.27 | 0.24 | 1.01 | 0.24 | 0.23 | 2.006 | 9.6 | 8.6 | 9 | 8.8 | 119.6 | 119.8 | 0.34 | 0.15 | 0.13 | 0 | 0.61 | 0.13 | 0.03 | 1.993 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.6 | AD |  | 2022-05-13 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220517 | 0.39 | 0.46 | 0.17 | 0.22 | 0.85 | 0.18 | 0.23 | 1.991 | 9.3 | 9 | 9.1 | 8.8 | 119.6 | 119.7 | 0.32 | 0.13 | 0.12 | -0.01 | 0.56 | 0.14 | 0 | 1.98 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.5 | ML |  | 2022-05-17 | 2.003 | 0.99401 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220518 | 0.36 | 0.48 | 0.17 | 0.22 | 0.84 | 0.2 | 0.24 | 1.997 | 9.3 | 8.9 | 9 | 8.9 | 119.6 | 119.8 | 0.32 | 0.14 | 0.11 | -0.01 | 0.53 | 0.13 | 0.02 | 1.986 | 11.5 | 11.5 | 11.7 | 11.5 | 119.4 | 119.6 | CVD |  | 2022-05-18 | 2.003 | 0.997 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220519 | 0.44 | 0.43 | 0.23 | 0.19 | 0.9 | 0.23 | 0.21 | 1.997 | 9.6 | 8.4 | 8.9 | 8.7 | 119.6 | 119.9 | 0.33 | 0.24 | 0.11 | 0.07 | 0.61 | 0.14 | 0.07 | 1.985 | 11.6 | 11.8 | 11.6 | 11.7 | 119.5 | 119.7 | ML |  | 2022-05-19 | 2.003 | 0.997 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220520 | 0.41 | 0.42 | 0.17 | 0.19 | 0.78 | 0.18 | 0.2 | 1.994 | 9.2 | 9.1 | 9.1 | 8.6 | 119.6 | 119.8 | 0.3 | 0.12 | 0.12 | -0.02 | 0.52 | 0.14 | -0.01 | 1.983 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.6 | GB |  | 2022-05-20 | 2.003 | 0.99551 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220524 | 0.46 | 0.43 | 0.16 | 0.23 | 0.88 | 0.16 | 0.25 | 2 | 9.5 | 8.7 | 9 | 8.6 | 119.6 | 119.9 | 0.33 | 0.14 | 0.11 | 0 | 0.57 | 0.12 | 0.05 | 1.987 | 11.5 | 11.4 | 11.6 | 11.5 | 119.4 | 119.7 | EF YM |  | 2022-05-24 | 2.003 | 0.9985 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220525 | 0.38 | 0.43 | 0.15 | 0.17 | 0.84 | 0.18 | 0.21 | 2.013 | 9.2 | 9 | 9 | 8.7 | 119.6 | 119.8 | 0.34 | 0.15 | 0.14 | -0.05 | 0.5 | 0.12 | -0.03 | 2.001 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.6 | ML |  | 2022-05-25 | 2.003 | 1.005 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220601 | 0.3 | 0.61 | 0.18 | 0.25 | 0.95 | 0.2 | 0.28 | 2.013 | 9.2 | 9 | 8.5 | 9.1 | 119.7 | 119.8 | 0.26 | 0.15 | 0.09 | 0.03 | 0.52 | 0.12 | 0.02 | 2 | 11.4 | 11.4 | 11.5 | 11.2 | 119.5 | 119.7 | ML |  | 2022-06-01 | 2.003 | 1.005 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220602 | 0.36 | 0.49 | 0.15 | 0.24 | 0.82 | 0.18 | 0.26 | 2.011 | 9.1 | 9.1 | 8.9 | 8.9 | 119.6 | 119.8 | 0.28 | 0.13 | 0.09 | -0.03 | 0.55 | 0.13 | 0 | 1.997 | 11.7 | 11.6 | 11.6 | 11.3 | 119.4 | 119.6 | GB |  | 2022-06-02 | 2.003 | 1.004 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220608 | 0.39 | 0.5 | 0.17 | 0.23 | 0.8 | 0.2 | 0.25 | 2.008 | 9.4 | 8.8 | 9.2 | 8.7 | 119.6 | 119.8 | 0.31 | 0.1 | 0.12 | 0 | 0.53 | 0.14 | 0.03 | 1.995 | 11.6 | 11.5 | 11.7 | 11.5 | 119.4 | 119.5 | YM |  | 2022-06-08 | 2.003 | 1.0025 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220609 | 0.39 | 0.43 | 0.2 | 0.21 | 0.74 | 0.19 | 0.24 | 2.005 | 9.2 | 8.9 | 9.1 | 8.7 | 119.7 | 119.8 | 0.31 | 0.13 | 0.14 | -0.02 | 0.57 | 0.15 | 0.01 | 1.995 | 11.5 | 11.5 | 11.6 | 11.5 | 119.5 | 119.6 | YM |  | 2022-06-09 | 2.003 | 1.001 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220610 | 0.38 | 0.48 | 0.17 | 0.23 | 0.86 | 0.2 | 0.26 | 2.004 | 9.2 | 9 | 9.1 | 8.8 | 119.6 | 119.8 | 0.28 | 0.17 | 0.04 | -0.01 | 0.48 | 0.1 | 0.01 | 1.992 | 11.6 | 11.5 | 11.6 | 11.4 | 119.4 | 119.5 | GB |  | 2022-06-10 | 2.003 | 1.0005 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220615 | 0.38 | 0.52 | 0.19 | 0.21 | 0.76 | 0.23 | 0.22 | 2.003 | 9.5 | 8.6 | 9 | 8.9 | 119.6 | 119.8 | 0.35 | 0.16 | 0.09 | -0.01 | 0.54 | 0.15 | 0.01 | 1.99 | 11.6 | 11.7 | 11.7 | 11.6 | 119.5 | 119.7 | GB |  | 2022-06-15 | 2.003 | 1 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220616 | 0.44 | 0.5 | 0.2 | 0.22 | 0.82 | 0.16 | 0.2 | 2.005 | 9.3 | 8.9 | 9.1 | 8.6 | 119.6 | 119.9 | 0.26 | 0.17 | 0.08 | 0.02 | 0.55 | 0.11 | 0.03 | 1.993 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.6 | GB |  | 2022-06-16 | 2.003 | 1.001 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220617 | 0.41 | 0.52 | 0.24 | 0.25 | 0.84 | 0.24 | 0.27 | 2.01 | 9.2 | 9.1 | 9.1 | 8.7 | 119.6 | 119.8 | 0.27 | 0.11 | 0.09 | 0.01 | 0.55 | 0.17 | -0.01 | 1.997 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.6 | GB |  | 2022-06-17 | 2.003 | 1.0035 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220621 | 0.42 | 0.38 | 0.23 | 0.23 | 0.79 | 0.21 | 0.24 | 2.008 | 8.8 | 9.4 | 9 | 8.6 | 119.7 | 119.9 | 0.31 | 0.16 | 0.09 | 0.03 | 0.58 | 0.12 | 0.03 | 1.997 | 11.8 | 11.6 | 11.7 | 11.7 | 119.5 | 119.7 | ML |  | 2022-06-21 | 2.003 | 1.0025 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220622 | 0.44 | 0.53 | 0.25 | 0.25 | 0.89 | 0.26 | 0.3 | 2.005 | 9.1 | 9.1 | 9 | 8.8 | 119.7 | 119.9 | 0.33 | 0.15 | 0.15 | -0.03 | 0.64 | 0.19 | 0.06 | 1.991 | 11.5 | 11.5 | 11.8 | 11.9 | 119.6 | 119.8 | YM |  | 2022-06-22 | 2.003 | 1.001 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220623 | 0.57 | 0.48 | 0.34 | 0.2 | 0.93 | 0.32 | 0.26 | 2.007 | 9.1 | 9.2 | 8.9 | 8.9 | 119.6 | 119.9 | 0.5 | 0.13 | 0.24 | -0.03 | 0.7 | 0.24 | 0.01 | 1.994 | 11.5 | 11.5 | 11.7 | 11.9 | 119.5 | 119.8 | ML |  | 2022-06-23 | 2.003 | 1.002 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220627 | 0.36 | 0.48 | 0.15 | 0.2 | 0.77 | 0.2 | 0.23 | 2.012 | 9 | 9.1 | 8.8 | 8.7 | 119.7 | 119.9 | 0.38 | 0.17 | 0.14 | -0.01 | 0.54 | 0.15 | -0.01 | 2.002 | 11.5 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | ML |  | 2022-06-27 | 2.003 | 1.0045 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220629 | 0.46 | 0.46 | 0.19 | 0.21 | 0.9 | 0.18 | 0.25 | 2.018 | 9.4 | 8.8 | 9 | 8.9 | 119.6 | 119.7 | 0.22 | 0.26 | 0 | -0.03 | 0.51 | 0.08 | 0.05 | 2.004 | 11.5 | 11.4 | 11.5 | 11.2 | 119.5 | 119.5 | ML |  | 2022-06-29 | 2.003 | 1.0075 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220630 | 0.4 | 0.52 | 0.14 | 0.27 | 0.86 | 0.2 | 0.23 | 2.013 | 9.7 | 8.1 | 9 | 8.9 | 119.6 | 119.7 | 0.25 | 0.25 | 0.01 | -0.06 | 0.47 | 0.1 | 0.02 | 2.002 | 11.6 | 11.7 | 11.5 | 11.2 | 119.5 | 119.4 | AD |  | 2022-06-30 | 2.003 | 1.005 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220701 | 0.48 | 0.58 | 0.21 | 0.22 | 0.94 | 0.22 | 0.3 | 2.02 | 9.4 | 8.7 | 9.1 | 9 | 119.6 | 119.5 | 0.35 | 0.28 | 0.11 | -0.02 | 0.53 | 0.15 | 0.01 | 2.008 | 11.4 | 11.4 | 11.6 | 11.2 | 119.5 | 119.3 | ML |  | 2022-07-01 | 2.003 | 1.0085 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220705 | 0.49 | 0.48 | 0.24 | 0.22 | 0.89 | 0.21 | 0.23 | 2.005 | 9 | 9.3 | 8.9 | 8.5 | 119.6 | 120 | 0.35 | 0.16 | 0.1 | -0.02 | 0.61 | 0.14 | 0.03 | 1.996 | 11.7 | 11.6 | 11.5 | 11.5 | 119.4 | 119.8 | ML |  | 2022-07-05 | 2.003 | 1.001 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220706 | 0.41 | 0.47 | 0.17 | 0.17 | 0.88 | 0.18 | 0.24 | 2.005 | 9.3 | 8.9 | 8.8 | 8.7 | 119.6 | 120 | 0.24 | 0.2 | 0.03 | -0.02 | 0.55 | 0.09 | 0.05 | 1.998 | 11.4 | 11.3 | 11.4 | 11.2 | 119.5 | 119.8 | ML |  | 2022-07-06 | 2.003 | 1.001 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220708 | 0.45 | 0.44 | 0.18 | 0.22 | 0.85 | 0.19 | 0.24 | 2.006 | 8.9 | 9.2 | 9.1 | 8.1 | 119.7 | 120 | 0.34 | 0.14 | 0.13 | 0 | 0.62 | 0.15 | 0.05 | 1.999 | 11.4 | 11.5 | 11.5 | 11.5 | 119.5 | 119.8 | ML |  | 2022-07-08 | 2.003 | 1.0015 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220711 | 0.48 | 0.43 | 0.26 | 0.18 | 0.82 | 0.21 | 0.18 | 1.995 | 9.4 | 9 | 9.1 | 8.6 | 119.6 | 120 | 0.32 | 0.14 | 0.12 | -0.05 | 0.55 | 0.18 | 0 | 1.98 | 11.9 | 11.8 | 11.8 | 12.1 | 119.5 | 119.8 | ML |  | 2022-07-11 | 2.003 | 0.99601 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220713 | 0.45 | 0.54 | 0.16 | 0.27 | 0.83 | 0.18 | 0.23 | 2.003 | 9.9 | 8.1 | 9 | 8.8 | 119.5 | 119.9 | 0.26 | 0.16 | 0.09 | -0.03 | 0.59 | 0.12 | 0.02 | 1.987 | 11.9 | 12.4 | 11.8 | 11.9 | 119.4 | 119.8 | YM |  | 2022-07-13 | 2.003 | 1 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220714 | 0.41 | 0.43 | 0.16 | 0.19 | 0.86 | 0.21 | 0.25 | 1.994 | 9.4 | 8.9 | 9 | 8.6 | 119.6 | 120 | 0.39 | 0.14 | 0.18 | -0.05 | 0.66 | 0.2 | -0.01 | 1.982 | 11.7 | 11.8 | 11.8 | 11.8 | 119.5 | 119.8 | ML |  | 2022-07-14 | 2.003 | 0.99551 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220719 | 0.48 | 0.39 | 0.29 | 0.23 | 0.92 | 0.25 | 0.28 | 1.989 | 9.9 | 8.1 | 8.8 | 9 | 119.5 | 119.9 | 0.33 | 0.15 | 0.14 | 0.01 | 0.57 | 0.13 | 0.01 | 1.974 | 11.9 | 12.3 | 11.8 | 11.8 | 119.4 | 119.8 | YM |  | 2022-07-19 | 2.003 | 0.99301 | 2 | 0.987 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220720 | 0.55 | 0.46 | 0.32 | 0.2 | 0.89 | 0.25 | 0.19 | 1.98 | 9.3 | 8.9 | 9.1 | 8.3 | 119.8 | 120.1 | 0.59 | 0.15 | 0.26 | -0.04 | 0.68 | 0.24 | -0.02 | 1.969 | 11.9 | 11.9 | 11.8 | 12.1 | 119.6 | 119.9 | YM |  | 2022-07-20 | 2.003 | 0.98852 | 2 | 0.9845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220721 | 0.49 | 0.51 | 0.21 | 0.27 | 0.98 | 0.21 | 0.3 | 1.988 | 9.3 | 9.1 | 9.2 | 8.6 | 119.6 | 119.9 | 0.33 | 0.2 | 0.11 | 0.03 | 0.68 | 0.16 | 0.06 | 1.973 | 11.8 | 11.9 | 11.8 | 12 | 119.5 | 119.8 | ML |  | 2022-07-21 | 2.003 | 0.99251 | 2 | 0.9865 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220722 | 0.41 | 0.46 | 0.18 | 0.23 | 0.78 | 0.18 | 0.26 | 1.991 | 9.6 | 8.7 | 9.1 | 8.7 | 119.6 | 119.9 | 0.27 | 0.12 | 0.07 | -0.03 | 0.52 | 0.12 | 0.01 | 1.976 | 11.8 | 12.1 | 11.8 | 12 | 119.5 | 119.8 | ML |  | 2022-07-22 | 2.003 | 0.99401 | 2 | 0.988 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220726 | 0.47 | 0.44 | 0.18 | 0.25 | 0.97 | 0.2 | 0.32 | 1.991 | 9.5 | 8.8 | 9 | 8.7 | 119.6 | 120 | 0.4 | 0.16 | 0.17 | 0 | 0.69 | 0.2 | 0.05 | 1.977 | 11.8 | 11.7 | 11.7 | 11.8 | 119.5 | 119.8 | ML |  | 2022-07-26 | 2.003 | 0.99401 | 2 | 0.9885 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220728 | 0.44 | 0.35 | 0.24 | 0.17 | 0.84 | 0.26 | 0.24 | 1.988 | 9.7 | 8.4 | 9 | 8.7 | 119.6 | 120 | 0.37 | 0.17 | 0.15 | -0.01 | 0.58 | 0.16 | -0.02 | 1.974 | 11.8 | 12.1 | 11.8 | 11.9 | 119.5 | 119.8 | ML |  | 2022-07-28 | 2.003 | 0.99251 | 2 | 0.987 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220729 | 0.5 | 0.42 | 0.26 | 0.22 | 0.91 | 0.27 | 0.26 | 1.988 | 9.4 | 9 | 9 | 8.7 | 119.6 | 120 | 0.3 | 0.17 | 0.09 | -0.03 | 0.58 | 0.14 | 0.03 | 1.971 | 11.8 | 11.8 | 11.9 | 12.1 | 119.5 | 119.8 | AC |  | 2022-07-29 | 2.003 | 0.99251 | 2 | 0.9855 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220802 | 0.52 | 0.57 | 0.26 | 0.27 | 1.02 | 0.26 | 0.27 | 1.994 | 9.7 | 8.6 | 9.2 | 8.6 | 119.6 | 120 | 0.35 | 0.15 | 0.12 | 0.02 | 0.57 | 0.14 | 0.02 | 1.976 | 11.8 | 12.1 | 11.8 | 12 | 119.5 | 119.8 | KA |  | 2022-08-02 | 2.003 | 0.99551 | 2 | 0.988 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220803 | 0.58 | 0.66 | 0.25 | 0.31 | 1.29 | 0.34 | 0.34 | 1.998 | 9.7 | 8.6 | 8.8 | 9.2 | 119.6 | 119.8 | 0.52 | 0.15 | 0.28 | 0.01 | 0.67 | 0.24 | -0.02 | 1.978 | 11.8 | 12 | 11.9 | 11.7 | 119.5 | 119.8 | AC |  | 2022-08-03 | 2.003 | 0.9975 | 2 | 0.989 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220804 | 0.52 | 0.38 | -0.35 | 0.22 | 1.02 | -0.34 | 0.17 | 1.998 | 9 | 8.8 | 7.6 | 10 | 120 | 119.5 | 0.12 | 0.31 | 0 | 0.15 | 0.52 | -0.03 | 0.11 | 1.983 | 11.9 | 11.6 | 12.3 | 11.9 | 119.8 | 119.4 | YM |  | 2022-08-04 | 2.003 | 0.9975 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220809 | 0.38 | 0.78 | 0.17 | 0.26 | 1.13 | 0.26 | 0.29 | 2.004 | 9.1 | 9.3 | 9.1 | 8.9 | 119.6 | 119.9 | 0.35 | 0.11 | 0.17 | 0.01 | 0.61 | 0.18 | 0.05 | 1.982 | 12.1 | 11.8 | 11.8 | 11.8 | 119.5 | 119.8 | GB |  | 2022-08-09 | 2.003 | 1.0005 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220810 | 0.44 | 0.52 | 0.18 | 0.23 | 0.92 | 0.18 | 0.26 | 1.998 | 9.4 | 8.9 | 9.1 | 8.5 | 119.7 | 120 | 0.29 | 0.2 | 0.09 | 0.03 | 0.55 | 0.11 | 0.07 | 1.984 | 11.7 | 11.8 | 11.7 | 11.9 | 119.5 | 119.9 | GB |  | 2022-08-10 | 2.003 | 0.9975 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220812 | 0.44 | 0.43 | 0.21 | 0.21 | 0.92 | 0.2 | 0.25 | 1.992 | 9.2 | 9.2 | 9 | 8.7 | 119.7 | 120 | 0.41 | 0.38 | 0.21 | 0.24 | 0.82 | 0.23 | 0.2 | 1.995 | 12.2 | 11.8 | 11.9 | 12.3 | 119.6 | 119.9 | KA |  | 2022-08-12 | 2.003 | 0.99451 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220816 | 0.55 | 0.84 | 0.26 | 0.37 | 1.19 | 0.2 | 0.37 | 2.009 | 9.4 | 9.1 | 9.2 | 8.9 | 119.6 | 119.9 | 0.38 | 0.32 | 0.13 | 0.18 | 0.71 | 0.17 | 0.15 | 1.984 | 11.9 | 11.7 | 11.9 | 12.2 | 119.5 | 119.8 | YM |  | 2022-08-16 | 2.003 | 1.003 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220817 | 0.67 | 0.41 | 0.41 | 0.19 | 1.04 | 0.38 | 0.22 | 1.994 | 9.8 | 8.4 | 9 | 8.7 | 119.5 | 120 | 0.49 | 0.13 | 0.25 | 0.01 | 0.63 | 0.22 | -0.01 | 1.978 | 11.8 | 12.2 | 11.8 | 11.9 | 119.5 | 119.8 | YM |  | 2022-08-17 | 2.003 | 0.99551 | 2 | 0.989 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220818 | 0.55 | 0.54 | 0.27 | 0.26 | 0.98 | 0.24 | 0.26 | 1.993 | 9.6 | 8.8 | 9.1 | 8.7 | 119.6 | 119.9 | 0.4 | 0.15 | 0.15 | -0.01 | 0.61 | 0.16 | 0.01 | 1.978 | 11.9 | 11.8 | 11.8 | 12 | 119.5 | 119.8 | GB |  | 2022-08-18 | 2.003 | 0.99501 | 2 | 0.989 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220819 | 0.46 | 0.5 | 0.19 | 0.2 | 0.88 | 0.23 | 0.25 | 1.999 | 9.6 | 8.7 | 9 | 8.9 | 119.6 | 119.9 | 0.37 | 0.15 | 0.11 | -0.05 | 0.66 | 0.16 | 0.01 | 1.985 | 11.8 | 11.9 | 11.8 | 11.9 | 119.5 | 119.8 | AC |  | 2022-08-19 | 2.003 | 0.998 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220823 | 0.63 | 0.83 | 0.19 | 0.36 | 0.98 | 0.23 | 0.41 | 2.004 | 9.4 | 9 | 9 | 9 | 119.7 | 119.9 | 0.44 | 0.21 | 0.17 | -0.02 | 0.61 | 0.18 | 0.06 | 1.98 | 11.9 | 11.7 | 11.7 | 11.7 | 119.5 | 119.9 | YM |  | 2022-08-23 | 2.003 | 1.0005 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220824 | 0.44 | 0.56 | 0.2 | 0.25 | 0.92 | 0.19 | 0.24 | 2.001 | 9.5 | 8.9 | 8.8 | 9.1 | 119.6 | 119.9 | 0.35 | 0.22 | 0.14 | 0.04 | 0.62 | 0.14 | 0.05 | 1.984 | 11.8 | 12.1 | 11.8 | 11.9 | 119.5 | 119.9 | GB |  | 2022-08-24 | 2.003 | 0.999 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220825 | 0.63 | 0.58 | 0.22 | 0.28 | 1.05 | 0.24 | 0.27 | 2.009 | 9.8 | 8.4 | 9 | 8.9 | 119.6 | 119.9 | 0.38 | 0.16 | 0.16 | 0.02 | 0.64 | 0.14 | 0.03 | 1.991 | 11.8 | 12.2 | 11.7 | 11.7 | 119.5 | 119.8 | AC |  | 2022-08-25 | 2.003 | 1.003 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220826 | 0.61 | 0.81 | 0.26 | 0.38 | 1.25 | 0.29 | 0.37 | 2.009 | 9.7 | 8.4 | 9 | 9 | 119.6 | 119.9 | 0.47 | 0.43 | 0.25 | 0.26 | 0.91 | 0.25 | 0.19 | 2.004 | 11.9 | 12.3 | 12 | 12 | 119.6 | 119.9 | AC |  | 2022-08-26 | 2.003 | 1.003 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220830 | 0.41 | 0.54 | 0.16 | 0.27 | 0.89 | 0.21 | 0.24 | 1.998 | 9.7 | 8.5 | 8.8 | 8.9 | 119.6 | 120 | 0.31 | 0.16 | 0.11 | -0.03 | 0.55 | 0.13 | 0.01 | 1.98 | 11.8 | 12.1 | 11.8 | 11.9 | 119.5 | 119.8 | CVD |  | 2022-08-30 | 2.003 | 0.9975 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220901 | 0.61 | 0.64 | 0.2 | 0.25 | 1.03 | 0.21 | 0.3 | 1.999 | 9.5 | 9 | 8.9 | 9.1 | 119.6 | 119.9 | 0.33 | 0.14 | 0.11 | 0 | 0.59 | 0.16 | 0.03 | 1.979 | 11.8 | 11.9 | 11.7 | 11.9 | 119.5 | 119.8 | ML |  | 2022-09-01 | 2.003 | 0.998 | 2 | 0.9895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220902 | 0.59 | 0.77 | 0.3 | 0.32 | 1.27 | 0.24 | 0.36 | 1.999 | 9.3 | 9.1 | 8.9 | 8.9 | 119.7 | 120 | 0.36 | 0.19 | 0.14 | -0.04 | 0.63 | 0.17 | 0.01 | 1.98 | 12 | 11.9 | 11.9 | 12.1 | 119.6 | 120 | GB |  | 2022-09-02 | 2.003 | 0.998 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220906 | 0.51 | 0.54 | 0.21 | 0.4 | 0.98 | 0.21 | 0.33 | 1.998 | 9.7 | 8.6 | 8.8 | 9.1 | 119.6 | 119.9 | 0.37 | 0.25 | 0.12 | -0.05 | 0.64 | 0.17 | 0.02 | 1.979 | 11.8 | 12.1 | 11.7 | 11.8 | 119.5 | 119.8 | ML |  | 2022-09-06 | 2.003 | 0.9975 | 2 | 0.9895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220907 | 0.54 | 0.64 | 0.28 | 0.22 | 1 | 0.27 | 0.28 | 1.999 | 9.6 | 8.8 | 9 | 9 | 119.6 | 119.9 | 0.44 | 0.17 | 0.17 | -0.01 | 0.7 | 0.19 | 0.04 | 1.982 | 11.8 | 12 | 11.8 | 11.9 | 119.4 | 119.8 | CVD KW |  | 2022-09-07 | 2.003 | 0.998 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220908 | 0.41 | 0.49 | 0.2 | 0.18 | 0.98 | 0.24 | 0.23 | 1.997 | 9.3 | 9.2 | 9 | 8.9 | 119.6 | 119.9 | 0.45 | 0.17 | 0.2 | -0.02 | 0.62 | 0.2 | -0.01 | 1.981 | 12.1 | 11.8 | 11.8 | 11.9 | 119.5 | 119.8 | ML |  | 2022-09-08 | 2.003 | 0.997 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220909 | 0.51 | 0.57 | 0.23 | 0.29 | 0.97 | 0.24 | 0.31 | 1.988 | 9.1 | 9.2 | 8.9 | 8.8 | 119.8 | 120.1 | 0.38 | 0.16 | 0.2 | -0.03 | 0.64 | 0.2 | -0.01 | 1.974 | 12.2 | 11.8 | 11.8 | 12 | 119.6 | 120 | CVD |  | 2022-09-09 | 2.003 | 0.99251 | 2 | 0.987 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220913 | 0.44 | 0.51 | 0.18 | 0.21 | 0.83 | 0.2 | 0.28 | 1.995 | 9.5 | 8.8 | 8.9 | 8.8 | 119.6 | 120 | 0.38 | 0.2 | 0.12 | 0.01 | 0.68 | 0.17 | 0.07 | 1.98 | 11.7 | 11.8 | 11.7 | 11.8 | 119.5 | 119.8 | ML |  | 2022-09-13 | 2.003 | 0.99601 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220914 | 0.5 | 0.5 | 0.26 | 0.22 | 0.89 | 0.24 | 0.22 | 1.996 | 9.6 | 8.7 | 8.8 | 8.9 | 119.6 | 120 | 0.39 | 0.13 | 0.16 | 0 | 0.64 | 0.19 | 0.01 | 1.982 | 11.8 | 12 | 11.7 | 11.8 | 119.4 | 119.8 | ML |  | 2022-09-14 | 2.003 | 0.99651 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220915 | 0.46 | 0.37 | 0.2 | 0.22 | 0.86 | 0.2 | 0.22 | 1.992 | 9.9 | 8.1 | 8.9 | 8.8 | 119.5 | 120 | [14 empty] | GB, R31 not measured because of downtime in the morning |  | 2022-09-15 | 2.003 | 0.99451 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220916 | 0.43 | 0.5 | 0.23 | 0.29 | 0.86 | 0.22 | 0.32 | 1.989 | 9.6 | 8.6 | 8.9 | 8.9 | 119.6 | 119.9 | 0.47 | 0.25 | 0.2 | -0.04 | 0.66 | 0.22 | 0 | 1.975 | 11.8 | 12 | 11.8 | 12 | 119.5 | 119.8 | GB |  | 2022-09-16 | 2.003 | 0.99301 | 2 | 0.9875 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220920 | 0.47 | 0.57 | 0.24 | 0.27 | 0.78 | 0.17 | 0.3 | 1.992 | 9.4 | 9 | 8.7 | 8.9 | 119.7 | 120 | 0.6 | 0.24 | 0.3 | -0.05 | 0.63 | 0.24 | 0.01 | 1.975 | 11.8 | 11.7 | 11.7 | 11.7 | 119.5 | 119.8 | YM |  | 2022-09-20 | 2.003 | 0.99451 | 2 | 0.9875 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220921 | 0.53 | 0.61 | 0.25 | 0.32 | 0.96 | 0.24 | 0.29 | 1.996 | 9.1 | 9.4 | 9 | 8.7 | 119.6 | 120 | 0.44 | 0.2 | 0.22 | 0.02 | 0.66 | 0.19 | 0.02 | 1.982 | 12 | 11.8 | 11.7 | 11.8 | 119.5 | 119.9 | ML |  | 2022-09-21 | 2.003 | 0.99651 | 2 | 0.991 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220922 | 0.53 | 0.45 | 0.28 | 0.19 | 0.95 | 0.28 | 0.2 | 1.993 | 9.5 | 8.9 | 8.7 | 8.9 | 119.7 | 120 | 0.42 | 0.17 | 0.2 | -0.01 | 0.63 | 0.22 | 0.01 | 1.984 | 12 | 11.8 | 11.8 | 11.9 | 119.5 | 119.8 | GB |  | 2022-09-22 | 2.003 | 0.99501 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220923 | 0.47 | 0.4 | 0.21 | 0.16 | 0.84 | 0.21 | 0.2 | 1.99 | 9.5 | 8.9 | 8.8 | 8.9 | 119.6 | 119.9 | 0.3 | 0.18 | 0.13 | 0.02 | 0.58 | 0.14 | 0.01 | 1.977 | 11.8 | 11.8 | 11.8 | 12 | 119.5 | 119.8 | GB |  | 2022-09-23 | 2.003 | 0.99351 | 2 | 0.9885 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220927 | 0.57 | 0.39 | 0.32 | 0.12 | 0.92 | 0.25 | 0.15 | 1.998 | 9.2 | 9.2 | 9.1 | 8.5 | 119.7 | 120 | 0.41 | 0.19 | 0.17 | -0.02 | 0.61 | 0.17 | -0.01 | 1.986 | 11.9 | 11.7 | 11.7 | 11.7 | 119.5 | 119.8 | GB |  | 2022-09-27 | 2.003 | 0.9975 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220928 | 0.6 | 0.49 | 0.27 | 0.17 | 0.94 | 0.27 | 0.21 | 2.004 | 9.1 | 9.4 | 9.1 | 8.5 | 119.6 | 120 | 0.44 | 0.19 | 0.21 | -0.04 | 0.71 | 0.23 | 0.04 | 1.991 | 12.1 | 11.7 | 11.7 | 11.9 | 119.5 | 119.8 | YM |  | 2022-09-28 | 2.003 | 1.0005 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220930 | 0.66 | 0.47 | 0.36 | 0.21 | 0.95 | 0.31 | 0.23 | 2.005 | 9.6 | 8.7 | 9 | 8.8 | 119.6 | 119.9 | 0.59 | 0.13 | 0.27 | -0.05 | 0.71 | 0.24 | -0.01 | 1.992 | 11.7 | 11.9 | 11.7 | 11.8 | 119.5 | 119.8 | ML |  | 2022-09-30 | 2.003 | 1.001 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221009 | 0.47 | 0.52 | 0.23 | 0.24 | 0.87 | 0.2 | 0.27 | 2.004 | 9.4 | 8.8 | 8.9 | 8.5 | 119.6 | 119.9 | 0.39 | 0.22 | 0.15 | 0.02 | 0.6 | 0.14 | 0.04 | 1.993 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.8 | CVD |  | 2022-10-09 | 2.003 | 1.0005 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221011 | 0.46 | 0.46 | 0.21 | 0.2 | 0.78 | 0.2 | 0.25 | 2.001 | 8.9 | 9.3 | 8.8 | 9 | 119.7 | 119.8 | 0.32 | 0.17 | 0.09 | -0.05 | 0.53 | 0.13 | -0.01 | 1.997 | 11.5 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | AC, MatriXX2, aftrernoon measurements |  | 2022-10-11 | 2.003 | 0.999 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221012 | 0.37 | 0.53 | 0.14 | 0.31 | 0.81 | 0.13 | 0.31 | 2.004 | 9.1 | 9.1 | 8.8 | 8.5 | 119.7 | 120 | 0.29 | 0.22 | 0.06 | -0.01 | 0.59 | 0.09 | 0.06 | 1.995 | 11.4 | 11.4 | 11.4 | 11.3 | 119.4 | 119.8 | AC |  | 2022-10-12 | 2.003 | 1.0005 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221013 | 0.39 | 0.49 | 0.17 | 0.22 | 0.76 | 0.14 | 0.24 | 2.009 | 9.1 | 9 | 8.9 | 8.5 | 119.6 | 120 | 0.27 | 0.18 | 0.1 | -0.03 | 0.51 | 0.11 | 0.02 | 1.997 | 11.4 | 11.4 | 11.6 | 11.6 | 119.5 | 119.7 | ML |  | 2022-10-13 | 2.003 | 1.003 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221014 | 0.34 | 0.67 | 0.21 | 0.34 | 0.88 | 0.17 | 0.3 | 2.005 | 9.5 | 8.4 | 8.6 | 8.9 | 119.6 | 119.9 | 0.26 | 0.23 | 0.04 | -0.04 | 0.49 | 0.08 | 0 | 1.995 | 11.6 | 11.6 | 11.4 | 11.1 | 119.4 | 119.7 | YM |  | 2022-10-14 | 2.003 | 1.001 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221019 | 0.46 | 0.45 | 0.14 | 0.22 | 0.85 | 0.18 | 0.26 | 2 | 9.3 | 8.8 | 8.7 | 8.7 | 119.6 | 120 | 0.29 | 0.14 | 0.06 | -0.05 | 0.55 | 0.1 | 0.01 | 1.989 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.8 | ML |  | 2022-10-19 | 2.003 | 0.9985 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221020 | 0.47 | 0.45 | 0.16 | 0.21 | 0.82 | 0.16 | 0.22 | 2.005 | 9.5 | 8.7 | 9 | 8.4 | 119.6 | 119.9 | 0.23 | 0.15 | 0.03 | 0 | 0.49 | 0.08 | 0.04 | 1.993 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.7 | ML |  | 2022-10-20 | 2.003 | 1.001 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221021 | 0.42 | 0.63 | 0.26 | 0.3 | 0.93 | 0.22 | 0.32 | 2 | 8.9 | 9.3 | 8.5 | 9 | 119.6 | 119.9 | 0.32 | 0.16 | 0.1 | -0.04 | 0.56 | 0.12 | 0.02 | 1.988 | 11.4 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | CVD Matrixx2 |  | 2022-10-21 | 2.003 | 0.9985 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221025 | 0.4 | 0.53 | 0.19 | 0.25 | 0.82 | 0.19 | 0.26 | 2.002 | 9 | 9.1 | 8.8 | 8.6 | 119.6 | 119.9 | 0.31 | 0.19 | 0.1 | 0 | 0.58 | 0.14 | -0.01 | 1.991 | 11.4 | 11.4 | 11.4 | 11.1 | 119.4 | 119.7 | on |  | 2022-10-25 | 2.003 | 0.9995 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221101 | 0.54 | 0.55 | 0.25 | 0.29 | 1.06 | 0.3 | 0.3 | 2.019 | 9.6 | 8.5 | 8.8 | 8.7 | 119.6 | 120 | 0.32 | 0.17 | 0.1 | -0.02 | 0.66 | 0.17 | 0.03 | 2.013 | 11.5 | 11.6 | 11.5 | 11.4 | 119.4 | 119.8 | YM |  | 2022-11-01 | 2.003 | 1.008 | 2 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221102 | 0.49 | 0.51 | 0.27 | 0.21 | 0.96 | 0.26 | 0.26 | 2.012 | 9.3 | 8.9 | 8.6 | 8.9 | 119.6 | 119.9 | 0.45 | 0.24 | 0.18 | 0.04 | 0.65 | 0.17 | 0.04 | 2.001 | 11.5 | 11.5 | 11.6 | 11.6 | 119.5 | 119.8 | ML |  | 2022-11-02 | 2.003 | 1.0045 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221104 | 0.49 | 0.68 | 0.18 | 0.44 | 1.01 | 0.21 | 0.38 | 2.016 | 8.8 | 9.2 | 8.2 | 8.9 | 119.7 | 120 | 0.18 | 0.28 | 0.02 | 0.17 | 0.46 | 0.06 | 0.15 | 2.004 | 11.5 | 11.4 | 11.3 | 11.1 | 119.5 | 119.9 | MatriXX_1 |  | 2022-11-04 | 2.003 | 1.0065 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221108 | 0.45 | 0.48 | 0.26 | 0.21 | 0.93 | 0.25 | 0.26 | 1.998 | 9.5 | 8.7 | 8.9 | 8.6 | 119.6 | 120 | 0.32 | 0.15 | 0.15 | 0.01 | 0.65 | 0.18 | 0.07 | 1.987 | 11.5 | 11.6 | 11.5 | 11.2 | 119.4 | 119.7 | YM |  | 2022-11-08 | 2.003 | 0.9975 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221109 | 0.53 | 0.53 | 0.31 | 0.29 | 0.98 | 0.27 | 0.29 | 2.003 | 9.5 | 8.6 | 8.8 | 8.5 | 119.6 | 120 | 0.37 | 0.2 | 0.15 | -0.02 | 0.63 | 0.19 | 0.01 | 1.994 | 11.5 | 11.4 | 11.4 | 11 | 119.4 | 119.8 | ML |  | 2022-11-09 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221110 | 0.42 | 0.45 | 0.22 | 0.25 | 0.95 | 0.22 | 0.26 | 1.998 | 8.7 | 9.4 | 9 | 8.4 | 119.6 | 119.9 | 0.33 | 0.16 | 0.14 | 0 | 0.61 | 0.16 | 0.04 | 1.991 | 11.6 | 11.6 | 11.5 | 11.3 | 119.3 | 119.7 | ML |  | 2022-11-10 | 2.003 | 0.9975 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221111 | 0.58 | 0.52 | 0.37 | 0.24 | 0.95 | 0.29 | 0.24 | 1.996 | 8.7 | 9.4 | 8.8 | 8.6 | 119.6 | 120 | 0.45 | 0.16 | 0.23 | 0.01 | 0.63 | 0.19 | 0.02 | 1.987 | 11.4 | 11.5 | 11.4 | 11.1 | 119.4 | 119.7 | ML, From DailyQA, Matrixx 2, REFERENCE images | MatriXX_2 | 2022-11-11 | 2.003 | 0.99651 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 2
20221111 | 0.43 | 0.32 | 0.2 | -0.04 | 0.81 | 0.18 | 0.02 | 1.994 | 9.3 | 9.1 | 8.6 | 8.8 | 119.6 | 119.9 | 0.2 | 0.41 | 0.04 | -0.24 | 0.67 | 0.07 | -0.23 | 1.986 | 11.4 | 11.6 | 11.4 | 11.3 | 119.2 | 119.6 | ML, From DailyQA, Matrixx 3, REFERENCE images | MatriXX_3 | 2022-11-11 | 2.003 | 0.99551 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 3
20221111 | 0.37 | 0.32 | 0.13 | -0.05 | 0.8 | 0.16 | 0.01 | 1.996 | 9 | 9.3 | 8.7 | 8.8 | 119.5 | 119.9 | 0.22 | 0.4 | 0.05 | -0.22 | 0.66 | 0.08 | -0.21 | 1.984 | 11.5 | 11.5 | 11.3 | 11.3 | 119.3 | 119.7 | ML, From DailyQA, Matrixx 3 | MatriXX_3 | 2022-11-11 | 2.003 | 0.99651 | 2 | 0.992 | 1.01 | 0.99 | 1.02 | 0.98 | 4
20221111 | 0.51 | 0.54 | 0.27 | 0.27 | 0.93 | 0.23 | 0.26 | 1.996 | 9.3 | 8.9 | 8.9 | 8.4 | 119.6 | 120 | 0.44 | 0.14 | 0.26 | 0.01 | 0.64 | 0.22 | 0.01 | 1.985 | 11.4 | 11.4 | 11.4 | 11.1 | 119.4 | 119.8 | ML |  | 2022-11-11 | 2.003 | 0.99651 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221116 | 0.55 | 0.5 | 0.33 | 0.28 | 1.01 | 0.3 | 0.28 | 2.006 | 9.7 | 8.3 | 8.8 | 8.5 | 119.5 | 120 | 0.35 | 0.12 | 0.13 | 0.03 | 0.59 | 0.15 | 0.03 | 1.997 | 11.6 | 11.7 | 11.5 | 11.1 | 119.4 | 119.7 | YM |  | 2022-11-16 | 2.003 | 1.0015 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221117 | 0.44 | 0.55 | 0.17 | 0.26 | 1.01 | 0.21 | 0.27 | 1.999 | 9.3 | 8.8 | 8.9 | 8.4 | 119.6 | 120 | 0.48 | 0.44 | 0.09 | 0.23 | 0.88 | 0.11 | 0.22 | 1.998 | 11.4 | 11.3 | 11.5 | 11.2 | 119.4 | 119.7 | YM |  | 2022-11-17 | 2.003 | 0.998 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.5 | 0.48 | 0.26 | 0.27 | 0.92 | 0.21 | 0.3 | 1.988 | 7.6 | 9.7 | 8.7 | 8 | 119.8 | 120.3 | 0.29 | 0.22 | 0.1 | 0.04 | 0.66 | 0.13 | 0.08 | 1.98 | 12 | 11.6 | 11.6 | 11.6 | 119.7 | 120.1 | CVD, students informed me about deviating gamma |  | 2022-11-18 | 2.003 | 0.99251 | 2 | 0.99 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221123 | 0.45 | 0.46 | 0.22 | 0.24 | 0.89 | 0.24 | 0.29 | 1.997 | 8.7 | 9.4 | 8.7 | 8.8 | 119.6 | 119.9 | 0.3 | 0.21 | 0.11 | 0.05 | 0.7 | 0.15 | 0.07 | 1.989 | 11.6 | 11.6 | 11.4 | 11.1 | 119.4 | 119.7 | ML |  | 2022-11-23 | 2.003 | 0.997 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221125 | 0.55 | 0.52 | 0.31 | 0.26 | 0.93 | 0.26 | 0.26 | 2.004 | 9.1 | 9.2 | 8.9 | 8.5 | 119.6 | 120 | 0.4 | 0.21 | 0.18 | 0.02 | 0.66 | 0.18 | -0.01 | 1.997 | 11.4 | 11.5 | 11.4 | 11.1 | 119.4 | 119.7 | CVD |  | 2022-11-25 | 2.003 | 1.0005 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221129 | 0.42 | 0.56 | 0.22 | 0.32 | 0.97 | 0.21 | 0.26 | 2.011 | 9.1 | 9.1 | 8.7 | 8.5 | 119.6 | 120 | 0.3 | 0.27 | 0.13 | 0.08 | 0.63 | 0.13 | 0.07 | 2.003 | 11.3 | 11.3 | 11.5 | 11.2 | 119.5 | 119.8 | CVD |  | 2022-11-29 | 2.003 | 1.004 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221201 | 0.38 | 0.77 | 0.28 | 0.39 | 1.13 | 0.27 | 0.37 | 2.008 | 9.6 | 8.4 | 8.5 | 8.9 | 119.6 | 119.9 | 0.27 | 0.2 | 0.08 | 0.05 | 0.64 | 0.13 | 0.06 | 1.997 | 11.5 | 11.5 | 11.3 | 11.1 | 119.4 | 119.8 | CVD |  | 2022-12-01 | 2.003 | 1.0025 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221202 | 0.6 | 0.5 | 0.3 | 0.27 | 1 | 0.26 | 0.31 | 2.005 | 9.5 | 8.7 | 8.9 | 8.6 | 119.6 | 119.9 | 0.4 | 0.17 | 0.17 | 0.01 | 0.67 | 0.19 | 0 | 1.997 | 11.5 | 11.4 | 11.4 | 11.1 | 119.4 | 119.7 | CVD, Matlab R2021b in new Citrix portal |  | 2022-12-02 | 2.003 | 1.001 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221206 | 0.61 | 0.52 | 0.34 | 0.29 | 1.08 | 0.3 | 0.31 | 2.003 | 9.6 | 8.4 | 8.8 | 8.7 | 119.6 | 119.9 | 0.42 | 0.26 | 0.2 | 0.09 | 0.71 | 0.17 | 0.05 | 1.991 | 11.7 | 11.9 | 11.7 | 11.6 | 119.4 | 119.7 | ML |  | 2022-12-06 | 2.003 | 1 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221207 | 0.47 | 0.52 | 0.2 | 0.25 | 1.03 | 0.25 | 0.29 | 2.021 | 9.4 | 8.7 | 8.6 | 8.6 | 119.6 | 120 | 0.35 | 0.2 | 0.09 | 0.04 | 0.59 | 0.13 | 0.02 | 2.012 | 11.3 | 11.2 | 11.3 | 11 | 119.4 | 119.8 | ML |  | 2022-12-07 | 2.003 | 1.009 | 2 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221208 | 0.57 | 0.58 | 0.3 | 0.28 | 1.09 | 0.33 | 0.32 | 2.003 | 9.1 | 9.3 | 9 | 8.6 | 119.6 | 119.9 | 0.51 | 0.24 | 0.21 | -0.01 | 0.68 | 0.22 | 0.03 | 1.992 | 11.4 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | ML |  | 2022-12-08 | 2.003 | 1 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221209 | 0.48 | 0.53 | 0.23 | 0.31 | 1.02 | 0.25 | 0.31 | 2.011 | 9 | 9.2 | 9 | 8.4 | 119.6 | 120 | 0.31 | 0.2 | 0.11 | 0.01 | 0.67 | 0.13 | 0.02 | 2.007 | 11.3 | 11.5 | 11.4 | 11 | 119.4 | 119.7 | ML, From DailyQA, Matrixx 2 | MatriXX_2 | 2022-12-09 | 2.003 | 1.004 | 2 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221209 | 0.53 | 0.47 | 0.28 | 0.25 | 0.94 | 0.25 | 0.3 | 2.005 | 9 | 9.2 | 8.7 | 8.7 | 119.6 | 119.9 | 0.31 | 0.2 | 0.13 | -0.03 | 0.61 | 0.16 | 0.02 | 1.999 | 11.3 | 11.4 | 11.3 | 11.1 | 119.4 | 119.7 | ML |  | 2022-12-09 | 2.003 | 1.001 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 2
20221220 | 0.61 | 0.57 | 0.26 | 0.09 | 0.94 | 0.23 | 0.14 | 2.011 | 9.8 | 8.2 | 8.6 | 8.9 | 119.5 | 119.9 | 0.26 | 0.4 | 0.1 | -0.2 | 0.67 | 0.12 | -0.2 | 2.002 | 11.6 | 11.7 | 11.4 | 11.4 | 119.3 | 119.7 | AC, Matrixx3 |  | 2022-12-20 | 2.003 | 1.004 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221222 | 0.54 | 0.44 | 0.22 | 0.14 | 0.87 | 0.2 | 0.12 | 2.016 | 9.3 | 9 | 8.5 | 8.9 | 119.6 | 119.9 | 0.26 | 0.31 | 0.14 | -0.17 | 0.63 | 0.14 | -0.2 | 2.002 | 11.5 | 11.5 | 11.4 | 11.3 | 119.3 | 119.7 | AC, Matrixx3 |  | 2022-12-22 | 2.003 | 1.0065 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221223 | 0.53 | 0.45 | 0.23 | 0.16 | 0.88 | 0.19 | 0.14 | 2.01 | 9.5 | 8.6 | 8.2 | 9 | 119.6 | 120 | 0.17 | 0.39 | 0.08 | -0.2 | 0.64 | 0.09 | -0.21 | 1.993 | 11.5 | 11.5 | 11.5 | 11.5 | 119.4 | 119.8 | AC, Matrixx3 |  | 2022-12-23 | 2.003 | 1.0035 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221229 | 0.44 | 0.41 | 0.21 | 0.05 | 0.91 | 0.22 | 0.08 | 2.022 | 9.1 | 9.3 | 8.6 | 8.7 | 119.6 | 120 | 0.26 | 0.39 | 0.11 | -0.18 | 0.7 | 0.12 | -0.2 | 2.008 | 11.7 | 11.6 | 11.6 | 11.7 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2022-12-29 | 2.003 | 1.0095 | 2 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221230 | 0.53 | 0.3 | 0.29 | 0.06 | 0.85 | 0.21 | 0.1 | 2.013 | 8.7 | 9.5 | 8.6 | 8.7 | 119.6 | 120 | 0.23 | 0.35 | 0.12 | -0.21 | 0.65 | 0.12 | -0.22 | 2.002 | 11.7 | 11.6 | 11.4 | 11.2 | 119.3 | 119.7 | ML, MatriXX_3 |  | 2022-12-30 | 2.003 | 1.005 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230103 | 0.63 | 0.47 | 0.21 | 0.07 | 0.92 | 0.18 | 0.11 | 2.013 | 9.4 | 9 | 8.5 | 8.9 | 119.6 | 120 | 0.21 | 0.41 | 0.07 | -0.22 | 0.72 | 0.12 | -0.23 | 2.001 | 11.5 | 11.5 | 11.5 | 11.5 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-01-03 | 2.003 | 1.005 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230111 | 0.59 | 0.49 | 0.21 | 0.07 | 0.93 | 0.2 | 0.13 | 2.019 | 9.4 | 8.9 | 8.5 | 9 | 119.6 | 119.9 | 0.17 | 0.43 | 0.01 | -0.22 | 0.59 | 0.05 | -0.18 | 2.009 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-01-11 | 2.003 | 1.008 | 2 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230113 | 0.69 | 0.6 | 0.29 | 0.12 | 0.93 | 0.22 | 0.11 | 2.018 | 9.8 | 8.4 | 8.5 | 9 | 119.5 | 119.9 | 0.23 | 0.42 | 0.05 | -0.19 | 0.59 | 0.07 | -0.19 | 2.006 | 11.7 | 11.8 | 11.6 | 11.5 | 119.3 | 119.7 | ML, MatriXX_3 |  | 2023-01-13 | 2.003 | 1.0075 | 2 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230117 | 0.65 | 0.49 | 0.31 | 0.06 | 0.77 | 0.21 | 0.1 | 2.004 | 9.6 | 8.5 | 8.5 | 8.9 | 119.6 | 120 | 0.23 | 0.43 | 0.1 | -0.21 | 0.71 | 0.08 | -0.25 | 1.994 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-01-17 | 2.003 | 1.0005 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230119 | 0.48 | 0.46 | 0.22 | 0.08 | 0.72 | 0.16 | 0.08 | 2.006 | 9.3 | 9 | 8.4 | 9.1 | 119.6 | 119.9 | 0.21 | 0.44 | 0.07 | -0.24 | 0.72 | 0.08 | -0.25 | 1.993 | 11.6 | 11.6 | 11.6 | 11.5 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-01-19 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230127 | 0.37 | 0.29 | 0.19 | 0.05 | 0.72 | 0.16 | 0.01 | 2.018 | 9.2 | 9.1 | 8.7 | 8.7 | 119.6 | 119.9 | 0.12 | 0.33 | 0.01 | -0.19 | 0.6 | 0.03 | -0.23 | 2.003 | 11.5 | 11.5 | 11.6 | 11.9 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-01-27 | 2.003 | 1.0075 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230131 | 0.47 | 0.34 | 0.16 | -0.05 | 0.76 | 0.11 | 0.01 | 2.005 | 9.4 | 8.9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.16 | 0.43 | 0.03 | -0.24 | 0.64 | 0.02 | -0.27 | 1.994 | 11.5 | 11.5 | 11.5 | 11.6 | 119.4 | 119.7 | YM |  | 2023-01-31 | 2.003 | 1.001 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230202 | 0.52 | 0.38 | 0.22 | 0.08 | 0.74 | 0.16 | 0.07 | 2.007 | 9.2 | 9.1 | 8.5 | 9 | 119.6 | 119.9 | 0.14 | 0.41 | 0.05 | -0.2 | 0.68 | 0.06 | -0.23 | 1.994 | 11.5 | 11.5 | 11.4 | 11.3 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-02-02 | 2.003 | 1.002 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230203 | 0.46 | 0.29 | 0.17 | 0.06 | 0.64 | 0.15 | 0.06 | 2.007 | 9.1 | 9.2 | 8.5 | 8.9 | 119.6 | 119.9 | 0.17 | 0.37 | 0.03 | -0.21 | 0.64 | 0.07 | -0.24 | 1.995 | 11.7 | 11.6 | 11.5 | 11.5 | 119.3 | 119.7 | ML, MatriXX_3 |  | 2023-02-03 | 2.003 | 1.002 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230207 | 0.56 | 0.35 | 0.19 | 0.03 | 0.8 | 0.16 | 0.07 | 2.011 | 9.1 | 9.3 | 8.5 | 8.9 | 119.6 | 119.9 | 0.2 | 0.31 | 0.03 | -0.14 | 0.55 | 0.05 | -0.19 | 1.999 | 11.5 | 11.5 | 11.3 | 11.3 | 119.4 | 119.7 | YM |  | 2023-02-07 | 2.003 | 1.004 | 2 | 0.9995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230208 | 0.52 | 0.26 | 0.11 | -0.08 | 0.84 | 0.14 | -0.08 | 2.009 | 9.4 | 9 | 8.6 | 9 | 119.6 | 119.9 | 0.16 | 0.48 | 0.03 | -0.27 | 0.76 | 0.04 | -0.27 | 1.994 | 11.5 | 11.5 | 11.5 | 11.6 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-02-08 | 2.003 | 1.003 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230209 | 0.48 | 0.39 | 0.19 | 0.07 | 0.73 | 0.16 | 0.07 | 2.006 | 9 | 9.3 | 8.5 | 9 | 119.6 | 119.9 | 0.24 | 0.47 | 0.08 | -0.27 | 0.75 | 0.08 | -0.28 | 1.993 | 11.8 | 11.6 | 11.5 | 11.6 | 119.4 | 119.7 | ML, MatriXX_3 |  | 2023-02-09 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230210 | 0.53 | 0.36 | 0.28 | 0.01 | 0.7 | 0.21 | 0.05 | 2.006 | 9.9 | 7.9 | 8.7 | 8.8 | 119.5 | 119.9 | 0.19 | 0.45 | 0.08 | -0.28 | 0.71 | 0.09 | -0.27 | 1.989 | 11.6 | 11.5 | 11.7 | 12 | 119.5 | 119.7 | ML, MatriXX_3 |  | 2023-02-10 | 2.003 | 1.0015 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230214 | 0.35 | 0.58 | 0.23 | 0.22 | 1.02 | 0.22 | 0.27 | 2.011 | 9.5 | 8.7 | 8.6 | 8.9 | 119.6 | 120 | 0.26 | 0.16 | 0.07 | 0 | 0.54 | 0.11 | 0 | 2.005 | 11.5 | 11.4 | 11.6 | 11.3 | 119.4 | 119.7 | AJ, MatriXX_2 | MatriXX_2 | 2023-02-14 | 2.003 | 1.004 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 2
20230214 | 0.35 | 0.58 | 0.23 | 0.22 | 1.02 | 0.22 | 0.27 | 2.011 | 9.5 | 8.7 | 8.6 | 8.9 | 119.6 | 120 | 0.26 | 0.16 | 0.07 | 0 | 0.54 | 0.11 | 0 | 2.005 | 11.5 | 11.4 | 11.6 | 11.3 | 119.4 | 119.7 | AE, MatriXX_2 | MatriXX_2 | 2023-02-14 | 2.003 | 1.004 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 3
20230214 | 0.35 | 0.58 | 0.23 | 0.22 | 1.02 | 0.22 | 0.27 | 2.011 | 9.5 | 8.7 | 8.6 | 8.9 | 119.6 | 120 | 0.26 | 0.16 | 0.07 | 0 | 0.54 | 0.11 | 0 | 2.005 | 11.5 | 11.4 | 11.6 | 11.3 | 119.4 | 119.7 | ML, MatriXX_2 |  | 2023-02-14 | 2.003 | 1.004 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 4
20230215 | 0.55 | 0.35 | 0.13 | 0.19 | 0.83 | 0.16 | 0.2 | 2.008 | 8.9 | 9.2 | 8.3 | 8.9 | 119.7 | 120 | 0.22 | 0.24 | 0.11 | -0.09 | 0.43 | 0.1 | -0.05 | 1.997 | 11.5 | 11.5 | 11.2 | 11.2 | 119.4 | 119.8 | SM, SR, JR, AR, Only Matrixx 1 on G1 from DailyQA | MatriXX_1 | 2023-02-15 | 2.003 | 1.0025 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 5
20230215 | 0.55 | 0.35 | 0.13 | 0.19 | 0.83 | 0.16 | 0.2 | 2.008 | 8.9 | 9.2 | 8.3 | 8.9 | 119.7 | 120 | 0.22 | 0.24 | 0.11 | -0.09 | 0.43 | 0.1 | -0.05 | 1.997 | 11.5 | 11.5 | 11.2 | 11.2 | 119.4 | 119.8 | ML, MatriXX_1 |  | 2023-02-15 | 2.003 | 1.0025 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 6
20230216 | 0.45 | 0.34 | 0.15 | -0.04 | 0.77 | 0.16 | 0 | 2.011 | 8.7 | 9.5 | 8.6 | 8.9 | 119.5 | 119.9 | 0.19 | 0.42 | 0.03 | -0.25 | 0.67 | 0.06 | -0.27 | 2 | 11.9 | 11.7 | 11.5 | 11.4 | 119.3 | 119.6 | ML, MatriXX_3 |  | 2023-02-16 | 2.003 | 1.004 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 7
20230217 | 0.56 | 0.45 | 0.31 | 0.24 | 0.95 | 0.26 | 0.27 | 2.005 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 119.9 | 0.33 | 0.12 | 0.11 | 0 | 0.55 | 0.12 | 0.01 | 2.001 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | EH SR OJ, MatriXX_2 | MatriXX_2 | 2023-02-17 | 2.003 | 1.001 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 8
20230217 | 0.56 | 0.45 | 0.31 | 0.24 | 0.95 | 0.26 | 0.27 | 2.005 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 119.9 | 0.33 | 0.12 | 0.11 | 0 | 0.55 | 0.12 | 0.01 | 2.001 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | ML, MatriXX_2 |  | 2023-02-17 | 2.003 | 1.001 | 2 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 9
20230221 | 0.52 | 0.36 | 0.25 | 0.16 | 0.88 | 0.23 | 0.18 | 2.001 | 9.1 | 9.2 | 9 | 8.5 | 119.6 | 120 | 0.37 | 0.13 | 0.11 | 0.01 | 0.62 | 0.11 | 0.03 | 1.994 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | SM, GF, AR, OJ, MatriXX_2 | MatriXX_2 | 2023-02-21 | 2.003 | 0.999 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 10
20230221 | 0.52 | 0.36 | 0.25 | 0.16 | 0.88 | 0.23 | 0.18 | 2.001 | 9.1 | 9.2 | 9 | 8.5 | 119.6 | 120 | 0.37 | 0.13 | 0.11 | 0.01 | 0.62 | 0.11 | 0.03 | 1.994 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | ML, MatriXX_2 |  | 2023-02-21 | 2.003 | 0.999 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 11
[29 empty] | Intermission _ Matrixx from DailyQA |  | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 12
20230222 | 0.44 | 0.46 | 0.16 | 0.23 | 0.89 | 0.2 | 0.26 | 1.998 | 9.1 | 9.2 | 8.9 | 8.6 | 119.6 | 119.9 | 0.3 | 0.16 | 0.07 | 0 | 0.59 | 0.11 | 0.03 | 1.99 | 11.5 | 11.6 | 11.6 | 11.5 | 119.3 | 119.7 | AR | MatriXX_2 | 2023-02-22 | 2.003 | 0.9975 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230223 | 0.46 | 0.41 | 0.18 | 0.22 | 0.89 | 0.2 | 0.24 | 2 | 9.1 | 9.2 | 8.7 | 8.8 | 119.6 | 119.9 | 0.29 | 0.14 | 0.05 | 0.02 | 0.55 | 0.11 | 0.01 | 1.993 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-02-23 | 2.003 | 0.9985 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230224 | 0.47 | 0.38 | 0.26 | 0.18 | 0.83 | 0.23 | 0.22 | 1.996 | 9.3 | 9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.25 | 0.18 | 0.06 | 0.03 | 0.48 | 0.09 | 0.01 | 1.985 | 11.4 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | EH | MatriXX_2 | 2023-02-24 | 2.003 | 0.99651 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230227 | 0.48 | 0.41 | 0.3 | 0.18 | 0.86 | 0.25 | 0.21 | 1.989 | 8.2 | 9.7 | 8.7 | 8.8 | 119.5 | 120 | 0.37 | 0.17 | 0.13 | -0.05 | 0.54 | 0.16 | -0.04 | 1.981 | 11.9 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | AC | MatriXX_2 | 2023-02-27 | 2.003 | 0.99301 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230228 | 0.41 | 0.43 | 0.18 | 0.22 | 0.91 | 0.2 | 0.25 | 1.999 | 9.2 | 9 | 8.8 | 8.5 | 119.6 | 120 | 0.36 | 0.21 | 0.14 | 0.03 | 0.61 | 0.13 | -0.02 | 1.994 | 11.4 | 11.4 | 11.4 | 11.1 | 119.4 | 119.8 | AR, AJ | MatriXX_2 | 2023-02-28 | 2.003 | 0.998 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230301 | 0.4 | 0.46 | 0.16 | 0.23 | 0.87 | 0.19 | 0.24 | 2.01 | 9.1 | 9.2 | 8.8 | 8.6 | 119.6 | 120 | 0.29 | 0.15 | 0.05 | 0.01 | 0.57 | 0.09 | 0.03 | 2.003 | 11.4 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | AR, JR | MatriXX_2 | 2023-03-01 | 2.003 | 1.0035 | 2 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230302 | 0.43 | 0.46 | 0.15 | 0.25 | 0.78 | 0.17 | 0.22 | 2.003 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.28 | 0.19 | 0.08 | -0.02 | 0.46 | 0.09 | -0.01 | 1.994 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | SM, SF | MatriXX_2 | 2023-03-02 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230303 | 0.41 | 0.48 | 0.23 | 0.24 | 0.9 | 0.24 | 0.25 | 2.006 | 9 | 9.3 | 8.8 | 8.6 | 119.6 | 119.9 | 0.3 | 0.12 | 0.07 | -0.02 | 0.53 | 0.11 | 0 | 1.996 | 11.5 | 11.6 | 11.5 | 11.4 | 119.3 | 119.7 | OJ | MatriXX_2 | 2023-03-03 | 2.003 | 1.0015 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230306 | 0.48 | 0.47 | 0.23 | 0.16 | 0.89 | 0.2 | 0.2 | 2.003 | 9.1 | 9.2 | 8.6 | 8.9 | 119.6 | 119.9 | 0.33 | 0.18 | 0.09 | -0.08 | 0.52 | 0.1 | -0.04 | 1.994 | 11.5 | 11.6 | 11.5 | 11.3 | 119.3 | 119.6 | AD, GB, run 2, after ion source replacement | MatriXX_2 | 2023-03-06 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230307 | 0.44 | 0.41 | 0.18 | 0.18 | 0.85 | 0.21 | 0.23 | 2.006 | 9.2 | 9.1 | 8.8 | 8.6 | 119.6 | 120 | 0.3 | 0.13 | 0.09 | -0.05 | 0.46 | 0.09 | -0.02 | 1.997 | 11.4 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | OJ | MatriXX_2 | 2023-03-07 | 2.003 | 1.0015 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230308 | 0.44 | 0.42 | 0.22 | 0.2 | 0.85 | 0.17 | 0.22 | 2.003 | 9.3 | 8.9 | 9 | 8.3 | 119.6 | 120 | 0.35 | 0.12 | 0.12 | -0.02 | 0.53 | 0.1 | -0.01 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.3 | 119.7 | JR, After beam alignment | MatriXX_2 | 2023-03-08 | 2.003 | 1 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230309 | 0.45 | 0.46 | 0.21 | 0.23 | 0.81 | 0.2 | 0.22 | 2.002 | 9.2 | 9 | 8.8 | 8.6 | 119.6 | 120 | 0.27 | 0.15 | 0.08 | -0.01 | 0.48 | 0.09 | 0 | 1.996 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | JR | MatriXX_2 | 2023-03-09 | 2.003 | 0.9995 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230310 | 0.33 | 0.63 | 0.18 | 0.29 | 0.87 | 0.16 | 0.29 | 2.006 | 9 | 9.2 | 8.6 | 8.8 | 119.6 | 120 | 0.27 | 0.16 | 0.07 | 0 | 0.47 | 0.08 | -0.01 | 1.997 | 11.3 | 11.5 | 11.4 | 11.2 | 119.4 | 119.7 | SM, AR | MatriXX_2 | 2023-03-10 | 2.003 | 1.0015 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230313 | 0.39 | 0.46 | 0.13 | 0.25 | 0.76 | 0.16 | 0.23 | 2 | 9.3 | 9 | 8.8 | 8.8 | 119.6 | 119.9 | 0.2 | 0.11 | 0.03 | -0.01 | 0.5 | 0.09 | 0.01 | 1.991 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | SM, SF | MatriXX_2 | 2023-03-13 | 2.003 | 0.9985 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230314 | 0.55 | 0.43 | 0.27 | 0.2 | 0.9 | 0.23 | 0.2 | 1.999 | 8.9 | 9.3 | 8.8 | 8.6 | 119.6 | 120 | 0.34 | 0.13 | 0.12 | -0.02 | 0.59 | 0.14 | 0.01 | 1.991 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | AE, AR | MatriXX_2 | 2023-03-14 | 2.003 | 0.998 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230315 | 0.42 | 0.4 | 0.18 | 0.2 | 0.94 | 0.2 | 0.24 | 1.995 | 9.3 | 9 | 8.7 | 8.8 | 119.6 | 120 | 0.27 | 0.16 | 0.06 | 0.01 | 0.52 | 0.09 | 0.03 | 1.989 | 11.5 | 11.4 | 11.5 | 11.2 | 119.3 | 119.7 | SM, AE | MatriXX_2 | 2023-03-15 | 2.003 | 0.99601 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230316 | 0.45 | 0.56 | 0.27 | 0.26 | 0.94 | 0.24 | 0.28 | 1.992 | 9.1 | 9.1 | 8.6 | 8.8 | 119.6 | 120 | 0.41 | 0.17 | 0.18 | 0.02 | 0.58 | 0.16 | -0.01 | 1.983 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | JR | MatriXX_2 | 2023-03-16 | 2.003 | 0.99451 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230320 | 0.43 | 0.42 | 0.2 | 0.17 | 0.74 | 0.18 | 0.2 | 1.999 | 9.2 | 9.1 | 8.6 | 8.9 | 119.6 | 119.9 | 0.3 | 0.12 | 0.08 | -0.02 | 0.53 | 0.09 | 0 | 1.989 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.7 | AC | MatriXX_2 | 2023-03-20 | 2.003 | 0.998 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230321 | 0.47 | 0.43 | 0.15 | 0.25 | 0.97 | 0.16 | 0.23 | 1.997 | 9.2 | 9.1 | 9 | 8.4 | 119.6 | 120 | 0.31 | 0.15 | -0.02 | 0.01 | 0.53 | 0.04 | 0.02 | 1.988 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | ML, ML corrected the analysis | MatriXX_2 | 2023-03-21 | 2.003 | 0.997 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 2
20230322 | 0.26 | 0.63 | 0.14 | 0.28 | 1.01 | 0.15 | 0.3 | 2 | 9.3 | 8.9 | 8.5 | 9 | 119.6 | 119.9 | 0.26 | 0.16 | 0.03 | 0.01 | 0.51 | 0.06 | 0.02 | 1.989 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | GK,SM | MatriXX_2 | 2023-03-22 | 2.003 | 0.9985 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 3
20230323 | 0.51 | 0.45 | 0.27 | 0.23 | 0.89 | 0.24 | 0.26 | 2.003 | 9 | 9.3 | 8.8 | 8.6 | 119.6 | 120 | 0.4 | 0.13 | 0.15 | -0.03 | 0.56 | 0.14 | 0 | 1.994 | 11.7 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | EH | MatriXX_2 | 2023-03-23 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 4
20230324 | 0.48 | 0.51 | 0.16 | 0.28 | 0.92 | 0.18 | 0.26 | 2.005 | 9.2 | 9 | 8.8 | 8.7 | 119.6 | 120 | 0.27 | 0.17 | 0.04 | 0.01 | 0.54 | 0.07 | 0.02 | 1.997 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | AE | MatriXX_2 | 2023-03-24 | 2.003 | 1.001 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 5
20230327 | 0.38 | 0.47 | 0.14 | 0.23 | 0.74 | 0.14 | 0.21 | 1.991 | 9.2 | 9 | 8.9 | 8.5 | 119.6 | 120 | 0.27 | 0.15 | 0.04 | 0 | 0.5 | 0.08 | 0.02 | 1.983 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | SM, AR | MatriXX_2 | 2023-03-27 | 2.003 | 0.99401 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 6
20230328 | 0.5 | 0.46 | 0.23 | 0.23 | 0.95 | 0.22 | 0.24 | 2.003 | 9.3 | 8.9 | 8.7 | 8.9 | 119.6 | 119.9 | 0.28 | 0.15 | 0.07 | -0.02 | 0.53 | 0.1 | 0.01 | 1.991 | 11.5 | 11.4 | 11.6 | 11.4 | 119.4 | 119.7 | GK,ON,SF | MatriXX_2 | 2023-03-28 | 2.003 | 1 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 7
20230329 | 0.43 | 0.44 | 0.18 | 0.22 | 0.9 | 0.19 | 0.25 | 2.001 | 9.3 | 9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.24 | 0.19 | 0.02 | 0.03 | 0.53 | 0.07 | 0.05 | 1.989 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.7 | JR | MatriXX_2 | 2023-03-29 | 2.003 | 0.999 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 8
20230330 | 0.42 | 0.5 | 0.19 | 0.28 | 0.93 | 0.22 | 0.3 | 2.001 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 120 | 0.3 | 0.13 | 0.07 | -0.02 | 0.63 | 0.12 | 0.05 | 1.992 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | SM, OJ | MatriXX_2 | 2023-03-30 | 2.003 | 0.999 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 9
20230331 | 0.45 | 0.42 | 0.16 | 0.21 | 0.76 | 0.15 | 0.24 | 2.002 | 9.3 | 9 | 8.7 | 8.9 | 119.6 | 119.9 | 0.27 | 0.14 | 0.06 | -0.05 | 0.54 | 0.07 | -0.01 | 1.991 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-03-31 | 2.003 | 0.9995 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 10
20230403 | 0.39 | 0.59 | 0.24 | 0.29 | 0.97 | 0.2 | 0.32 | 2.004 | 9.3 | 8.9 | 8.6 | 8.9 | 119.6 | 119.9 | 0.34 | 0.13 | 0.09 | 0.01 | 0.61 | 0.11 | 0.02 | 1.994 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-04-03 | 2.003 | 1.0005 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 11
20230404 | 0.55 | 0.43 | 0.23 | 0.2 | 0.96 | 0.2 | 0.24 | 2 | 9.2 | 9.1 | 8.7 | 8.8 | 119.6 | 120 | 0.25 | 0.12 | 0.04 | 0 | 0.48 | 0.07 | 0.02 | 1.992 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | AE, GF | MatriXX_2 | 2023-04-04 | 2.003 | 0.9985 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 12
20230405 | 0.47 | 0.46 | 0.25 | 0.23 | 0.85 | 0.22 | 0.23 | 2.004 | 9.4 | 8.8 | 8.6 | 8.7 | 119.6 | 120 | 0.36 | 0.12 | 0.14 | -0.03 | 0.51 | 0.13 | -0.02 | 1.995 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.8 | ON, SM, JR | MatriXX_2 | 2023-04-05 | 2.003 | 1.0005 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 13
20230406 | 0.45 | 0.51 | 0.2 | 0.25 | 0.93 | 0.22 | 0.27 | 2.005 | 8.7 | 9.4 | 8.8 | 8.7 | 119.6 | 120 | 0.29 | 0.12 | 0.08 | 0 | 0.55 | 0.12 | 0.03 | 1.996 | 11.9 | 11.7 | 11.6 | 11.7 | 119.4 | 119.8 | OJ | MatriXX_2 | 2023-04-06 | 2.003 | 1.001 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 14
20230407 | 0.45 | 0.44 | 0.21 | 0.21 | 0.85 | 0.21 | 0.27 | 2.008 | 9.1 | 9.2 | 8.7 | 8.9 | 119.6 | 119.9 | 0.33 | 0.13 | 0.11 | 0 | 0.6 | 0.13 | 0.03 | 2 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | JR,SM | MatriXX_2 | 2023-04-07 | 2.003 | 1.0025 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 15
20230411 | 0.41 | 0.43 | 0.18 | 0.21 | 0.85 | 0.19 | 0.26 | 2.004 | 9.1 | 9.1 | 8.9 | 8.7 | 119.6 | 119.9 | 0.34 | 0.17 | 0.13 | -0.01 | 0.59 | 0.13 | -0.01 | 1.992 | 11.5 | 11.6 | 11.6 | 11.6 | 119.4 | 119.7 | OJ, GK | MatriXX_2 | 2023-04-11 | 2.003 | 1.0005 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 16
20230412 | 0.46 | 0.46 | 0.19 | 0.22 | 0.91 | 0.21 | 0.23 | 1.996 | 9.3 | 9 | 8.7 | 8.8 | 119.6 | 120 | 0.28 | 0.15 | 0.06 | -0.02 | 0.58 | 0.1 | 0.02 | 1.986 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | GK,EH | MatriXX_2 | 2023-04-12 | 2.003 | 0.99651 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 17
20230413 | 0.39 | 0.5 | 0.16 | 0.26 | 0.91 | 0.19 | 0.21 | 1.999 | 9.1 | 9.1 | 8.9 | 8.4 | 119.6 | 120 | 0.31 | 0.14 | 0.11 | 0 | 0.55 | 0.12 | -0.01 | 1.989 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | AR | MatriXX_2 | 2023-04-13 | 2.003 | 0.998 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 18
20230414 | 0.48 | 0.46 | 0.21 | 0.22 | 0.85 | 0.22 | 0.23 | 1.994 | 9.1 | 9.1 | 8.9 | 8.6 | 119.6 | 120 | 0.37 | 0.16 | 0.14 | -0.04 | 0.51 | 0.14 | -0.01 | 1.985 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | AR, JR | MatriXX_2 | 2023-04-14 | 2.003 | 0.99551 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 19
20230417 | 0.29 | 0.65 | 0.2 | 0.28 | 0.93 | 0.17 | 0.29 | 2.003 | 9.3 | 9 | 8.5 | 8.9 | 119.6 | 119.9 | 0.29 | 0.18 | 0.06 | -0.04 | 0.52 | 0.09 | 0 | 1.992 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | GK,SM | MatriXX_2 | 2023-04-17 | 2.003 | 1 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 20
20230418 | 0.49 | 0.53 | 0.18 | 0.23 | 0.92 | 0.17 | 0.19 | 2.004 | 9.1 | 9.1 | 8.9 | 8.6 | 119.6 | 120 | 0.22 | 0.15 | 0.04 | 0.02 | 0.56 | 0.06 | 0.03 | 1.994 | 11.5 | 11.5 | 11.5 | 11.5 | 119.4 | 119.8 | AJ | MatriXX_2 | 2023-04-18 | 2.003 | 1.0005 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 21
20230419 | 0.44 | 0.5 | 0.23 | 0.28 | 0.85 | 0.2 | 0.25 | 2.003 | 9.4 | 8.8 | 8.8 | 8.7 | 119.6 | 119.9 | 0.37 | 0.14 | 0.16 | 0.03 | 0.56 | 0.12 | 0 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | JR | MatriXX_2 | 2023-04-19 | 2.003 | 1 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 22
20230420 | 0.45 | 0.56 | 0.2 | 0.3 | 0.83 | 0.18 | 0.29 | 2 | 8.9 | 9.3 | 8.9 | 8.5 | 119.6 | 120 | 0.33 | 0.2 | 0.07 | 0.01 | 0.61 | 0.13 | 0.1 | 1.992 | 11.6 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | JR,AR | MatriXX_2 | 2023-04-20 | 2.003 | 0.9985 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 23
20230421 | 0.53 | 0.46 | 0.23 | 0.21 | 0.91 | 0.21 | 0.24 | 2.006 | 9.4 | 8.8 | 8.7 | 8.7 | 119.6 | 120 | 0.26 | 0.11 | 0.06 | 0 | 0.57 | 0.09 | 0.01 | 1.993 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.8 | GK,AR | MatriXX_2 | 2023-04-21 | 2.003 | 1.0015 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 24
20230424 | 0.47 | 0.49 | 0.18 | 0.23 | 0.84 | 0.19 | 0.22 | 1.996 | 9.1 | 9.2 | 8.9 | 8.5 | 119.6 | 120 | 0.28 | 0.17 | 0.06 | -0.02 | 0.56 | 0.1 | 0.02 | 1.989 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | AR, AJ | MatriXX_2 | 2023-04-24 | 2.003 | 0.99651 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 25
20230425 | 0.48 | 0.43 | 0.24 | 0.18 | 0.88 | 0.2 | 0.19 | 2.001 | 9.2 | 9 | 9 | 8.3 | 119.6 | 120 | 0.32 | 0.17 | 0.11 | 0.02 | 0.5 | 0.12 | 0.01 | 1.989 | 11.5 | 11.5 | 11.6 | 11.7 | 119.4 | 119.7 | OJ, AR | MatriXX_2 | 2023-04-25 | 2.003 | 0.999 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 26
20230426 | 0.43 | 0.4 | 0.18 | 0.17 | 0.91 | 0.22 | 0.26 | 2.012 | 9.3 | 9 | 8.6 | 8.6 | 119.6 | 120 | 0.28 | 0.18 | 0.02 | -0.01 | 0.54 | 0.1 | 0 | 2.002 | 11.3 | 11.4 | 11.3 | 11.2 | 119.4 | 119.8 | ON, SM | MatriXX_2 | 2023-04-26 | 2.003 | 1.0045 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 27
20230427 | 0.49 | 0.44 | 0.24 | 0.24 | 0.84 | 0.22 | 0.24 | 2.001 | 9.1 | 9.2 | 8.6 | 8.8 | 119.6 | 120 | 0.35 | 0.19 | 0.13 | 0.04 | 0.61 | 0.15 | 0.03 | 1.992 | 11.5 | 11.6 | 11.5 | 11.3 | 119.4 | 119.7 | SM, JR | MatriXX_2 | 2023-04-27 | 2.003 | 0.999 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 28
20230428 | 0.51 | 0.52 | 0.22 | 0.25 | 0.93 | 0.21 | 0.27 | 2.003 | 9.3 | 8.9 | 8.8 | 8.6 | 119.6 | 120 | 0.29 | 0.12 | 0.06 | -0.01 | 0.58 | 0.1 | 0.02 | 1.994 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | OJ, SM | MatriXX_2 | 2023-04-28 | 2.003 | 1 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 29
20230502 | 0.44 | 0.49 | 0.19 | 0.25 | 0.86 | 0.18 | 0.23 | 1.999 | 9.1 | 9.1 | 8.9 | 8.4 | 119.6 | 120 | 0.26 | 0.2 | 0.05 | 0.04 | 0.6 | 0.1 | 0.05 | 1.99 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | EN,AJ | MatriXX_2 | 2023-05-02 | 2.003 | 0.998 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 30
20230503 | 0.49 | 0.64 | 0.31 | 0.28 | 1.01 | 0.21 | 0.31 | 1.998 | 9.2 | 9.1 | 8.2 | 9.1 | 119.6 | 119.9 | 0.28 | 0.14 | 0.09 | 0.01 | 0.53 | 0.11 | 0.02 | 1.987 | 11.5 | 11.5 | 11.7 | 11.3 | 119.4 | 119.8 | JR | MatriXX_2 | 2023-05-03 | 2.003 | 0.9975 | 2 | 0.9935 | 1.01 | 0.99 | 1.02 | 0.98 | 31
20230504 | 0.39 | 0.45 | 0.13 | 0.23 | 0.87 | 0.2 | 0.27 | 2.003 | 9 | 9.2 | 8.7 | 8.6 | 119.6 | 120 | 0.27 | 0.14 | 0.06 | 0.01 | 0.59 | 0.1 | 0.02 | 1.997 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.7 | GK,AR,EF | MatriXX_2 | 2023-05-04 | 2.003 | 1 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 32
20230504 | 0.52 | 0.41 | 0.21 | 0.23 | 0.89 | 0.16 | 0.27 | 1.997 | 9.6 | 8.4 | 8.6 | 8.8 | 119.6 | 119.9 | 0.37 | 0.13 | 0.11 | 0.01 | 0.55 | 0.11 | 0.01 | 1.998 | 11.8 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | JR | MatriXX_2 | 2023-05-04 | 2.003 | 0.997 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 33
20230504 | 0.52 | 0.41 | 0.21 | 0.23 | 0.89 | 0.16 | 0.27 | 1.997 | 9.6 | 8.4 | 8.6 | 8.8 | 119.6 | 119.9 | 0.37 | 0.13 | 0.11 | 0.01 | 0.55 | 0.11 | 0.01 | 1.998 | 11.8 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | JR | MatriXX_2 | 2023-05-04 | 2.003 | 0.997 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 34
20230505 | 0.46 | 0.42 | 0.24 | 0.23 | 0.85 | 0.2 | 0.23 | 2.003 | 8.8 | 9.4 | 8.8 | 8.5 | 119.6 | 120 | 0.37 | 0.13 | 0.11 | 0.01 | 0.55 | 0.11 | 0.01 | 1.998 | 11.8 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | OJ | MatriXX_2 | 2023-05-05 | 2.003 | 1 | 2 | 0.999 | 1.01 | 0.99 | 1.02 | 0.98 | 35
20230508 | 0.49 | 0.4 | 0.26 | 0.22 | 0.88 | 0.21 | 0.22 | 2 | 9 | 9.3 | 8.9 | 8.4 | 119.6 | 120 | 0.26 | 0.13 | 0.09 | 0 | 0.54 | 0.1 | 0.02 | 1.994 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | SM, AJ | MatriXX_2 | 2023-05-08 | 2.003 | 0.9985 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 36
20230509 | 0.53 | 0.47 | 0.25 | 0.26 | 0.95 | 0.23 | 0.26 | 1.998 | 9.1 | 9.2 | 8.7 | 8.6 | 119.6 | 120 | 0.29 | 0.15 | 0.1 | -0.04 | 0.55 | 0.1 | -0.03 | 1.988 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | EN | MatriXX_2 | 2023-05-09 | 2.003 | 0.9975 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 37
20230510 | 0.19 | 0.61 | 0.09 | 0.27 | 0.87 | 0.1 | 0.29 | 2.004 | 9.5 | 8.7 | 8.5 | 8.8 | 119.6 | 120 | 0.2 | 0.15 | 0.02 | 0.02 | 0.48 | 0.06 | 0.07 | 1.992 | 11.6 | 11.5 | 11.5 | 11.3 | 119.4 | 119.8 | AR | MatriXX_2 | 2023-05-10 | 2.003 | 1.0005 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 38
20230511 | 0.4 | 0.45 | 0.15 | 0.21 | 0.83 | 0.17 | 0.25 | 2.007 | 9.4 | 8.8 | 8.6 | 8.7 | 119.6 | 120 | 0.21 | 0.13 | 0.05 | -0.01 | 0.52 | 0.09 | 0.02 | 1.993 | 11.5 | 11.5 | 11.4 | 11.2 | 119.4 | 119.7 | JR | MatriXX_2 | 2023-05-11 | 2.003 | 1.002 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 39
20230512 | 0.47 | 0.4 | 0.21 | 0.24 | 0.85 | 0.2 | 0.23 | 1.999 | 9.2 | 9 | 8.7 | 8.6 | 119.6 | 120 | 0.32 | 0.17 | 0.12 | 0.02 | 0.58 | 0.11 | 0.03 | 1.99 | 11.4 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EN,GK,ON | MatriXX_2 | 2023-05-12 | 2.003 | 0.998 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 40
20230515 | 0.41 | 0.51 | 0.19 | 0.28 | 0.92 | 0.21 | 0.27 | 1.992 | 9.1 | 9.1 | 8.7 | 8.8 | 119.6 | 119.9 | 0.34 | 0.21 | 0.1 | 0.06 | 0.57 | 0.11 | 0.04 | 1.985 | 11.5 | 11.5 | 11.4 | 11.2 | 119.4 | 119.8 | AR | MatriXX_2 | 2023-05-15 | 2.003 | 0.99451 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 41
20230516 | 0.4 | 0.42 | 0.13 | 0.23 | 0.85 | 0.19 | 0.25 | 2.001 | 9.2 | 9 | 8.8 | 8.6 | 119.6 | 120 | 0.22 | 0.14 | 0.03 | 0.02 | 0.55 | 0.09 | 0.05 | 1.991 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.8 | GF, AE, EF | MatriXX_2 | 2023-05-16 | 2.003 | 0.999 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 42
20230517 | 0.42 | 0.47 | 0.2 | 0.21 | 0.88 | 0.19 | 0.24 | 2.004 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 120 | 0.26 | 0.15 | 0.09 | 0 | 0.52 | 0.1 | -0.01 | 1.992 | 11.4 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | SM, GK | MatriXX_2 | 2023-05-17 | 2.003 | 1.0005 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 43
20230519 | 0.32 | 0.58 | 0.18 | 0.28 | 0.94 | 0.15 | 0.3 | 2.012 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.29 | 0.17 | 0.07 | 0.04 | 0.51 | 0.08 | 0.03 | 2.002 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | EF, AJ, OJ | MatriXX_2 | 2023-05-19 | 2.003 | 1.0045 | 2 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 44
20230522 | 0.39 | 0.64 | 0.26 | 0.32 | 1.08 | 0.2 | 0.31 | 2.002 | 9.4 | 8.9 | 8.6 | 9 | 119.6 | 119.9 | 0.31 | 0.18 | 0.1 | 0.03 | 0.55 | 0.11 | 0.01 | 1.994 | 11.6 | 11.5 | 11.6 | 11.3 | 119.4 | 119.7 | AJ, OJ | MatriXX_2 | 2023-05-22 | 2.003 | 0.9995 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 45
20230523 | 0.42 | 0.43 | 0.17 | 0.19 | 0.88 | 0.2 | 0.23 | 2 | 9.4 | 8.7 | 8.7 | 8.7 | 119.6 | 119.9 | 0.28 | 0.14 | 0.1 | 0.01 | 0.53 | 0.11 | -0.01 | 1.991 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | AE, EF | MatriXX_2 | 2023-05-23 | 2.003 | 0.9985 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 46
20230524 | 0.52 | 0.41 | 0.21 | 0.23 | 0.89 | 0.16 | 0.27 | 1.997 | 9.6 | 8.4 | 8.6 | 8.8 | 119.6 | 119.9 | 0.23 | 0.16 | 0.04 | 0.04 | 0.54 | 0.05 | 0.07 | 1.986 | 11.6 | 11.6 | 11.5 | 11.2 | 119.4 | 119.7 | JR | MatriXX_2 | 2023-05-24 | 2.003 | 0.997 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 47
20230525 | 0.47 | 0.31 | 0.14 | -0.04 | 0.77 | 0.13 | -0.04 | 1.998 | 9 | 9.3 | 8.7 | 8.8 | 119.6 | 119.9 | 0.19 | 0.43 | 0.04 | -0.23 | 0.65 | 0.05 | -0.25 | 1.985 | 11.5 | 11.5 | 11.4 | 11.4 | 119.3 | 119.6 | GB, YM | MatriXX_3 | 2023-05-25 | 2.003 | 0.9975 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 48
20230525 | 0.5 | 0.5 | 0.21 | 0.28 | 0.78 | 0.18 | 0.26 | 2.004 | 9.1 | 9.1 | 8.7 | 8.8 | 119.6 | 119.9 | 1.21 | 0.47 | -1 | 0.47 | 1.68 | -0.88 | 0.14 | 4.854 | 13.6 | 13.5 | 15.8 | 13.5 | 105.4 | 100.8 | GK,JR, Checked and ok with Matrixx3 /GB - Weird valus on R31, did the test 2 times. | MatriXX_2 | 2023-05-25 | 2.003 | 1.0005 | 2 | 2.427 | 1.01 | 0.99 | 1.02 | 0.98 | 49
20230526 | 0.45 | 0.48 | 0.21 | 0.26 | 0.92 | 0.21 | 0.29 | 2.002 | 9.3 | 8.9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.31 | 0.15 | 0.11 | 0 | 0.61 | 0.13 | 0.02 | 1.993 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.7 | GK,SF | MatriXX_2 | 2023-05-26 | 2.003 | 0.9995 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 50
20230529 | 0.4 | 0.46 | 0.17 | 0.24 | 0.88 | 0.21 | 0.27 | 1.995 | 9.2 | 9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.29 | 0.16 | 0.09 | 0.03 | 0.57 | 0.11 | 0.03 | 1.986 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | GK,SF | MatriXX_2 | 2023-05-29 | 2.003 | 0.99601 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 51
20230530 | 0.5 | 0.51 | 0.17 | 0.27 | 1.01 | 0.2 | 0.27 | 1.997 | 9.3 | 8.9 | 8.7 | 8.7 | 119.6 | 120 | 0.28 | 0.09 | 0.05 | -0.02 | 0.52 | 0.08 | -0.01 | 1.988 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | SM, JR | MatriXX_2 | 2023-05-30 | 2.003 | 0.997 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 52
20230531 | 0.42 | 0.36 | 0.17 | 0.21 | 0.87 | 0.19 | 0.24 | 2.003 | 9.4 | 8.8 | 8.7 | 8.9 | 119.6 | 119.9 | 0.27 | 0.16 | 0.05 | 0.01 | 0.52 | 0.09 | 0.02 | 1.992 | 11.5 | 11.4 | 11.4 | 11.3 | 119.4 | 119.7 | GF, SF | MatriXX_2 | 2023-05-31 | 2.003 | 1 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 53
20230601 | 0.47 | 0.52 | 0.19 | 0.29 | 0.93 | 0.21 | 0.29 | 2.003 | 9 | 9.2 | 8.8 | 8.8 | 119.6 | 119.9 | 0.29 | 0.12 | 0.08 | -0.02 | 0.56 | 0.1 | 0.02 | 1.991 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-06-01 | 2.003 | 1 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 54
20230602 | 0.47 | 0.51 | 0.2 | 0.25 | 0.8 | 0.15 | 0.22 | 1.984 | 9.2 | 9 | 9 | 8.4 | 119.6 | 120 | 0.27 | 0.13 | 0.07 | 0 | 0.53 | 0.09 | 0.01 | 1.976 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.7 | EF | MatriXX_2 | 2023-06-02 | 2.003 | 0.99051 | 2 | 0.988 | 1.01 | 0.99 | 1.02 | 0.98 | 55
20230605 | 0.46 | 0.4 | 0.23 | 0.19 | 0.94 | 0.22 | 0.21 | 2.005 | 9.3 | 9 | 8.9 | 8.5 | 119.6 | 119.9 | 0.26 | 0.1 | 0.07 | -0.01 | 0.59 | 0.12 | 0.03 | 1.994 | 11.5 | 11.4 | 11.6 | 11.4 | 119.4 | 119.7 | AR | MatriXX_2 | 2023-06-05 | 2.003 | 1.001 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 56
20230607 | 0.21 | 0.66 | 0.09 | 0.29 | 0.9 | 0.12 | 0.34 | 2 | 9.3 | 8.8 | 8.3 | 9 | 119.6 | 119.9 | 0.3 | 0.19 | 0.07 | 0.05 | 0.55 | 0.07 | 0.06 | 1.99 | 11.4 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | AJ, EF | MatriXX_2 | 2023-06-07 | 2.003 | 0.9985 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 57
20230608 | 0.35 | 0.51 | 0.11 | 0.25 | 0.8 | 0.18 | 0.29 | 2 | 9.2 | 9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.25 | 0.2 | 0.07 | 0.04 | 0.56 | 0.08 | 0.03 | 1.993 | 11.4 | 11.4 | 11.4 | 11.2 | 119.3 | 119.7 | EF | MatriXX_2 | 2023-06-08 | 2.003 | 0.9985 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 58
20230609 | 0.48 | 0.49 | 0.17 | 0.24 | 0.92 | 0.17 | 0.23 | 1.991 | 9.3 | 8.9 | 8.8 | 8.6 | 119.6 | 119.9 | 0.26 | 0.19 | 0.04 | 0.05 | 0.56 | 0.1 | 0.06 | 1.983 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-06-09 | 2.003 | 0.99401 | 2 | 0.9915 | 1.01 | 0.99 | 1.02 | 0.98 | 59
20230612 | 0.44 | 0.5 | 0.23 | 0.24 | 0.86 | 0.22 | 0.27 | 1.994 | 9.2 | 9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.37 | 0.11 | 0.14 | -0.01 | 0.6 | 0.17 | 0.02 | 1.986 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | EF | MatriXX_2 | 2023-06-12 | 2.003 | 0.99551 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 60
20230613 | 0.54 | 0.43 | 0.18 | 0.22 | 0.92 | 0.21 | 0.24 | 2.001 | 9.1 | 9.2 | 8.8 | 8.6 | 119.6 | 119.9 | 0.24 | 0.14 | 0.05 | 0.02 | 0.56 | 0.09 | 0.02 | 1.992 | 11.4 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | EF | MatriXX_2 | 2023-06-13 | 2.003 | 0.999 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 61
20230614 | 0.41 | 0.72 | 0.25 | 0.34 | 1.07 | 0.2 | 0.35 | 2.002 | 9.3 | 9 | 8.6 | 8.9 | 119.6 | 119.9 | 0.24 | 0.13 | 0.06 | -0.01 | 0.54 | 0.1 | 0.03 | 1.994 | 11.5 | 11.5 | 11.4 | 11.2 | 119.4 | 119.7 | GK,ON | MatriXX_2 | 2023-06-14 | 2.003 | 0.9995 | 2 | 0.997 | 1.01 | 0.99 | 1.02 | 0.98 | 62
20230615 | 0.49 | 0.5 | 0.2 | 0.26 | 0.92 | 0.18 | 0.3 | 2.002 | 9.3 | 8.8 | 8.9 | 8.7 | 119.6 | 119.9 | 0.23 | 0.14 | 0.04 | 0.02 | 0.59 | 0.1 | 0.08 | 1.992 | 11.5 | 11.4 | 11.6 | 11.5 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-06-15 | 2.003 | 0.9995 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 63
20230616 | 0.47 | 0.49 | 0.21 | 0.23 | 0.87 | 0.22 | 0.28 | 1.994 | 9.2 | 9.1 | 8.9 | 8.6 | 119.6 | 119.9 | 0.36 | 0.11 | 0.14 | -0.03 | 0.57 | 0.14 | 0.01 | 1.986 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | EN | MatriXX_2 | 2023-06-16 | 2.003 | 0.99551 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 64
20230619 | 0.5 | 0.68 | 0.29 | 0.3 | 1.09 | 0.24 | 0.32 | 2.033 | 9.5 | 8.8 | 8.2 | 9.2 | 119.5 | 119.8 | 0.3 | 0.16 | 0.11 | 0 | 0.59 | 0.15 | 0.05 | 2.017 | 11.6 | 11.7 | 11.6 | 11.3 | 119.3 | 119.7 | GK, After LLRF calibration /ML | MatriXX_2 | 2023-06-19 | 2.003 | 1.015 | 2 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 65
20230620 | 0.42 | 0.51 | 0.22 | 0.29 | 0.96 | 0.23 | 0.31 | 1.998 | 9.1 | 9.2 | 8.9 | 8.7 | 119.6 | 119.9 | 0.32 | 0.15 | 0.13 | 0.03 | 0.58 | 0.13 | 0.03 | 1.986 | 11.5 | 11.6 | 11.6 | 11.7 | 119.4 | 119.7 | EN | MatriXX_2 | 2023-06-20 | 2.003 | 0.9975 | 2 | 0.993 | 1.01 | 0.99 | 1.02 | 0.98 | 66
20230621 | 0.43 | 0.45 | 0.19 | 0.22 | 0.95 | 0.22 | 0.28 | 2.002 | 9.3 | 8.9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.35 | 0.1 | 0.16 | 0.01 | 0.65 | 0.16 | 0.07 | 1.99 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | ON | MatriXX_2 | 2023-06-21 | 2.003 | 0.9995 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 67
20230622 | 0.46 | 0.49 | 0.16 | 0.22 | 0.92 | 0.19 | 0.2 | 2.001 | 9 | 9.2 | 8.8 | 8.6 | 119.6 | 119.9 | 0.33 | 0.16 | 0.09 | 0.04 | 0.53 | 0.11 | 0.03 | 1.992 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-06-22 | 2.003 | 0.999 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 68
20230626 | 0.29 | 0.61 | 0.13 | 0.31 | 0.79 | 0.15 | 0.3 | 2.008 | 9.3 | 8.8 | 8.7 | 8.8 | 119.6 | 119.9 | 0.28 | 0.12 | 0.11 | 0.02 | 0.53 | 0.12 | 0 | 2 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-06-26 | 2.003 | 1.0025 | 2 | 1 | 1.01 | 0.99 | 1.02 | 0.98 | 69
20230627 | 0.47 | 0.56 | 0.28 | 0.32 | 0.95 | 0.25 | 0.32 | 1.999 | 9.4 | 8.7 | 8.9 | 8.6 | 119.6 | 119.9 | 0.38 | 0.17 | 0.17 | 0.07 | 0.57 | 0.15 | 0.05 | 1.992 | 11.5 | 11.4 | 11.5 | 11.2 | 119.3 | 119.7 | GK | MatriXX_2 | 2023-06-27 | 2.003 | 0.998 | 2 | 0.996 | 1.01 | 0.99 | 1.02 | 0.98 | 70
20230628 | 0.37 | 0.52 | 0.19 | 0.33 | 0.96 | 0.25 | 0.31 | 2.002 | 9.2 | 8.9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.36 | 0.2 | 0.16 | 0.1 | 0.6 | 0.15 | 0.03 | 1.993 | 11.4 | 11.3 | 11.3 | 11.2 | 119.3 | 119.7 | EN | MatriXX_2 | 2023-06-28 | 2.003 | 0.9995 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 71
20230629 | 0.45 | 0.58 | 0.26 | 0.32 | 0.98 | 0.26 | 0.3 | 1.997 | 9 | 9.1 | 8.8 | 8.6 | 119.6 | 119.9 | 0.42 | 0.16 | 0.19 | 0.06 | 0.67 | 0.18 | 0.03 | 1.99 | 11.3 | 11.5 | 11.3 | 11.1 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-06-29 | 2.003 | 0.997 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 72
20230630 | 0.4 | 0.72 | 0.22 | 0.32 | 1.06 | 0.2 | 0.34 | 2.004 | 9.3 | 8.9 | 8.5 | 8.9 | 119.6 | 119.9 | 0.38 | 0.29 | 0.12 | 0.12 | 0.64 | 0.11 | 0.07 | 1.995 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | AR | MatriXX_2 | 2023-06-30 | 2.003 | 1.0005 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 73
20230703 | 0.39 | 0.5 | 0.2 | 0.25 | 0.86 | 0.19 | 0.26 | 1.999 | 9 | 9.2 | 8.9 | 8.4 | 119.6 | 119.9 | 0.31 | 0.16 | 0.15 | 0.02 | 0.55 | 0.13 | 0.02 | 1.989 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | OJ | MatriXX_2 | 2023-07-03 | 2.003 | 0.998 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 74
20230704 | 0.52 | 0.46 | 0.2 | 0.23 | 0.98 | 0.21 | 0.27 | 2 | 9.2 | 9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.33 | 0.15 | 0.13 | 0.03 | 0.55 | 0.12 | 0.04 | 1.991 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | AR | MatriXX_2 | 2023-07-04 | 2.003 | 0.9985 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 75
20230705 | 0.49 | 0.52 | 0.29 | 0.28 | 0.94 | 0.25 | 0.32 | 1.993 | 9 | 9.3 | 8.8 | 8.7 | 119.6 | 119.9 | 0.3 | 0.13 | 0.11 | 0.03 | 0.54 | 0.12 | 0.03 | 1.985 | 11.6 | 11.6 | 11.5 | 11.5 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-07-05 | 2.003 | 0.99501 | 2 | 0.9925 | 1.01 | 0.99 | 1.02 | 0.98 | 76
20230706 | 0.45 | 0.48 | 0.27 | 0.26 | 0.96 | 0.25 | 0.29 | 1.997 | 9.1 | 9.1 | 8.8 | 8.7 | 119.6 | 119.9 | 0.36 | 0.16 | 0.17 | 0.04 | 0.59 | 0.14 | 0.04 | 1.99 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-07-06 | 2.003 | 0.997 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 77
20230707 | 0.41 | 0.53 | 0.27 | 0.27 | 1.07 | 0.26 | 0.32 | 1.987 | 9.3 | 8.9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.37 | 0.15 | 0.15 | 0.04 | 0.61 | 0.14 | 0.05 | 1.981 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | AR | MatriXX_2 | 2023-07-07 | 2.003 | 0.99201 | 2 | 0.9905 | 1.01 | 0.99 | 1.02 | 0.98 | 78
20230710 | 0.38 | 0.67 | 0.27 | 0.33 | 1.05 | 0.26 | 0.31 | 2 | 9.6 | 8.6 | 8.6 | 8.9 | 119.6 | 119.9 | 0.4 | 0.11 | 0.16 | -0.01 | 0.58 | 0.13 | -0.01 | 1.99 | 11.5 | 11.3 | 11.4 | 11.2 | 119.3 | 119.7 | AR | MatriXX_2 | 2023-07-10 | 2.003 | 0.9985 | 2 | 0.995 | 1.01 | 0.99 | 1.02 | 0.98 | 79
20230711 | 0.41 | 0.53 | 0.17 | 0.28 | 0.97 | 0.18 | 0.28 | 2.005 | 9.3 | 8.9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.28 | 0.16 | 0.09 | 0.02 | 0.63 | 0.1 | 0.05 | 1.995 | 11.5 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | EN | MatriXX_2 | 2023-07-11 | 2.003 | 1.001 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 80
20230712 | 0.25 | 0.67 | 0.12 | 0.32 | 1.02 | 0.14 | 0.35 | 1.998 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.27 | 0.17 | 0.11 | 0.03 | 0.57 | 0.1 | 0.05 | 1.988 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-07-12 | 2.003 | 0.9975 | 2 | 0.994 | 1.01 | 0.99 | 1.02 | 0.98 | 81
20230713 | 0.31 | 0.69 | 0.15 | 0.31 | 0.95 | 0.14 | 0.33 | 2 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.23 | 0.16 | 0.08 | 0.04 | 0.45 | 0.08 | 0.03 | 1.989 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | OJ | MatriXX_2 | 2023-07-13 | 2.003 | 0.9985 | 2 | 0.9945 | 1.01 | 0.99 | 1.02 | 0.98 | 82
20230714 | 0.43 | 0.6 | 0.26 | 0.32 | 0.99 | 0.24 | 0.32 | 2.011 | 9.2 | 9 | 8.3 | 8.9 | 119.6 | 119.9 | 0.4 | 0.11 | 0.2 | -0.01 | 0.53 | 0.17 | -0.02 | 2.004 | 11.3 | 11.3 | 11.4 | 11.1 | 119.4 | 119.8 | OJ | MatriXX_2 | 2023-07-14 | 2.003 | 1.004 | 2 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 83
20230717 | 0.35 | 0.49 | 0.21 | 0.29 | 0.9 | 0.21 | 0.29 | 2.002 | 9.2 | 9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.32 | 0.13 | 0.13 | -0.04 | 0.57 | 0.14 | 0.01 | 1.991 | 11.4 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | EN | MatriXX_2 | 2023-07-17 | 2.003 | 0.9995 | 2 | 0.9955 | 1.01 | 0.99 | 1.02 | 0.98 | 84
20230718 | 0.56 | 0.8 | 0.33 | 0.48 | 1.23 | 0.31 | 0.38 | 2.013 | 9.4 | 9.1 | 9 | 8.6 | 119.6 | 120 | 0.48 | 0.43 | 0.26 | 0.3 | 0.9 | 0.24 | 0.22 | 2.005 | 11.6 | 11.6 | 11.5 | 11.5 | 119.4 | 119.8 | AR | MatriXX_2 | 2023-07-18 | 2.003 | 1.005 | 2 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 85
20230719 | 0.43 | 0.64 | 0.28 | 0.31 | 1.02 | 0.22 | 0.32 | 2.007 | 9.3 | 8.9 | 8.4 | 8.8 | 119.6 | 120 | 0.34 | 0.17 | 0.13 | 0.01 | 0.57 | 0.12 | 0.04 | 1.996 | 11.4 | 11.3 | 11.4 | 11.2 | 119.4 | 119.7 | AR | MatriXX_2 | 2023-07-19 | 2.003 | 1.002 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 86
20230720 | 0.33 | 0.71 | 0.18 | 0.33 | 1.03 | 0.18 | 0.35 | 2.002 | 9.2 | 9.1 | 8.4 | 9 | 119.6 | 119.9 | 0.27 | 0.11 | 0.12 | 0.01 | 0.58 | 0.13 | 0.04 | 1.993 | 11.5 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | ON | MatriXX_2 | 2023-07-20 | 2.003 | 0.9995 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 87
20230721 | 0.43 | 0.62 | 0.16 | 0.34 | 0.99 | 0.17 | 0.3 | 2.005 | 9.2 | 9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.26 | 0.21 | 0.11 | 0.09 | 0.61 | 0.11 | 0.07 | 1.995 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | EN | MatriXX_2 | 2023-07-21 | 2.003 | 1.001 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 88
20230724 | 0.39 | 0.62 | 0.25 | 0.31 | 1.07 | 0.26 | 0.35 | 2.002 | 9 | 9.3 | 8.6 | 8.8 | 119.6 | 119.9 | 0.33 | 0.15 | 0.15 | -0.01 | 0.62 | 0.15 | 0.02 | 1.995 | 11.5 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | ON | MatriXX_2 | 2023-07-24 | 2.003 | 0.9995 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 89
20230725 | 0.41 | 0.49 | 0.22 | 0.24 | 0.8 | 0.18 | 0.26 | 2.004 | 9.2 | 9 | 8.6 | 8.8 | 119.6 | 120 | 0.24 | 0.13 | 0.07 | 0.02 | 0.43 | 0.07 | 0.01 | 1.997 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | AR | MatriXX_2 | 2023-07-25 | 2.003 | 1.0005 | 2 | 0.9985 | 1.01 | 0.99 | 1.02 | 0.98 | 90
20230726 | 0.32 | 0.73 | 0.19 | 0.34 | 1.09 | 0.2 | 0.37 | 2.006 | 8.9 | 9.3 | 8.6 | 8.9 | 119.6 | 119.9 | 0.28 | 0.19 | 0.08 | 0.03 | 0.49 | 0.1 | 0.03 | 1.996 | 11.4 | 11.5 | 11.4 | 11.2 | 119.4 | 119.8 | EN | MatriXX_2 | 2023-07-26 | 2.003 | 1.0015 | 2 | 0.998 | 1.01 | 0.99 | 1.02 | 0.98 | 91
20230727 | 0.41 | 0.54 | 0.21 | 0.3 | 0.96 | 0.21 | 0.31 | 2.003 | 9 | 9.2 | 8.8 | 8.6 | 119.6 | 119.9 | 0.28 | 0.23 | 0.13 | 0.09 | 0.63 | 0.13 | 0.06 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | EN | MatriXX_2 | 2023-07-27 | 2.003 | 1 | 2 | 0.9965 | 1.01 | 0.99 | 1.02 | 0.98 | 92
20230728 | 0.44 | 0.61 | 0.28 | 0.3 | 1.04 | 0.22 | 0.36 | 2.005 | 8.9 | 9.3 | 8.5 | 8.9 | 119.6 | 119.9 | 0.29 | 0.21 | 0.1 | 0.07 | 0.61 | 0.1 | 0.06 | 1.995 | 11.6 | 11.6 | 11.5 | 11.2 | 119.4 | 119.8 | EN | MatriXX_2 | 2023-07-28 | 2.003 | 1.001 | 2 | 0.9975 | 1.01 | 0.99 | 1.02 | 0.98 | 93
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 94
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 95
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 96
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 97
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 98
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 99
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 100
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 101
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 102
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 103
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 104
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 105
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 106
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 107
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 108
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 109
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 110
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 111
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 112
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 113
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 114
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 115
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 116
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 117
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 118
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 119
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 120
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 121
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 122
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 123
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 124
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 125
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 126
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 127
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 128
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 129
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 130
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 131
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 132
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 133
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 134
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 135
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 136
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 137
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 138
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 139
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 140
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 141
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 142
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 143
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 144
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 145
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 146
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 147
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 148
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 149
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 150
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 151
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 152
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 153
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 154
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 155
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 156
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 157
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 158
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 159
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 160
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 161
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 162
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 163
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 164
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 165
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 166
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 167
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 168
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 169
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 170
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 171
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 172
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 173
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 174
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 175
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 176
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 177
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 178
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 179
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 180
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 181
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 182
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 183
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 184
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 185
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 186
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 187
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 188
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 189
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 190
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 191
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 192
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 193
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 194
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 195
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 196
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 197
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 198
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 199
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 200
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 201
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 202
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 203
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 204
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 205
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 206
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 207
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 208
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 209
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 210
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 211
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 212
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 213
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 214
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 215
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 216
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 217
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 218
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 219
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 220
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 221
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 222
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 223
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 224
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 225
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 226
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 227
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 228
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 229
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 230
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 231
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 232
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 233
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 234
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 235
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 236
[31 empty] | #VALUE! | 2.003 | 0 | 2 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 237

## GTR2
Daily cubes parameters | [13 empty] | Reference | from measurements for Matrixx 1 | [6 empty] | 2019 |  |  |  |  | 2019
MatriXX PT | [14 empty] | R15_SOBP10_2Gy |  | R31_SOBP10_2Gy |  |  |  |  | R15_SOBP10_2Gy |  | R31_SOBP10 |  |  | R15_SOBP10_2Gy |  | R31_SOBP10
[11 empty] | S | F |  | D (from meas) | 1.9774 |  | 1.9991 |  |  | P |  | 8.5927 |  | 11.7214 |  | FWHM | 119.9065 |  | 119.6661
Set-up: MatriXX position adjusted using lasers or X-rays (chambers at iso); MyQA Project: GTR2_MAtrixMorgonQA | [10 empty] | 0.5 | 1 |  | 1% | 0.019774 | 1.9972 | 0.019991 | 2.0191 |  | 1 |  | 9.5927 |  | 12.7214 |  | 1 | 120.9065 |  | 120.6661
| F | S | Dose | FWHM (mm) | Penumbra (mm) | [5 empty] | -0.5 | -1 |  | -1% | -0.019774 | 1.9576 | -0.019991 | 1.9791 |  | -1 |  | 7.5927 |  | 10.7214 |  | -1 | 118.9065 |  | 118.6661
warrning | 1% | 0.5% | 1% | 1 | 1 | [5 empty] | 1 | 2 |  | 2% | 0.039548 | 2.0169 | 0.039982 | 2.0391 |  | 2 |  | 10.5927 |  | 13.7214 |  | 2 | 121.9065 |  | 121.6661
action level | 2% | 1% | 2% | 2 | 2 | [5 empty] | -1 | -2 |  | -2% | -0.039548 | 1.9379 | -0.039982 | 1.9591 |  | -2 |  | 6.5927 |  | 9.7214 |  | -2 | 117.9065 |  | 117.6661
Date | R15_SOBP10_2Gy, 10 cm WET, 9.2 SW | [13 empty] | R31_SOBP10_2Gy , 26 cm WET, 24.8 SW | [13 empty] | Notes | From DailyQA
| Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) |  |  |  | ref R15_SOBP10_2Gy | Diff R15_SOBP10_2Gy | ref R31_SOBP10_2Gy | Diff R31_SOBP10_2Gy | 1% | -1% | 2% | -2% | Ref.
20190220 | 0.48 | 0.7 | -0.26 | 0.36 | 1.32 | -0.27 | 0.38 | 1.98 | 9 | 9.1 | 7.1 | 9.4 | 119.8 | 119.9 | 0.25 | 0.35 | -0.09 | 0.06 | 0.71 | -0.05 | 0.08 | 2.005 | 11.6 | 11.6 | 12.2 | 11.6 | 119.4 | 119.7 | LS |  | 2019-02-20 | 1.9774 | 1.0013 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190313 | 0.64 | 0.65 | -0.36 | 0.2 | 1.2 | -0.35 | 0.13 | 1.988 | 8.2 | 9.6 | 9.5 | 6.9 | 119.8 | 120.1 | 0.32 | 0.46 | -0.16 | 0.21 | 0.83 | -0.12 | 0.17 | 2.016 | 12.1 | 11.8 | 11.9 | 12.2 | 119.5 | 119.8 | LS |  | 2019-03-13 | 1.9774 | 1.0054 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190409 | 0.63 | 0.62 | -0.36 | 0.17 | 1.17 | -0.3 | 0.16 | 1.994 | 9.1 | 9 | 9.4 | 7.4 | 119.8 | 120 | 0.32 | 0.41 | -0.17 | 0.11 | 0.9 | -0.13 | 0.15 | 2.014 | 11.6 | 11.8 | 11.7 | 12 | 119.6 | 119.9 | LS |  | 2019-04-09 | 1.9774 | 1.0084 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190424 | 0.59 | 0.57 | -0.38 | 0.13 | 1.19 | -0.32 | 0.15 | 1.979 | 8.6 | 9.1 | 9.4 | 6.9 | 119.9 | 120.1 | 0.15 | 0.24 | 0.01 | 0.02 | 0.64 | -0.04 | 0.03 | 2 | 11.6 | 11.5 | 11.7 | 12 | 119.6 | 120 | LS |  | 2019-04-24 | 1.9774 | 1.0008 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190508 | 0.59 | 0.62 | -0.29 | 0.19 | 1.25 | -0.24 | 0.21 | 1.977 | 8.2 | 9.4 | 9.6 | 6.9 | 119.8 | 120.1 | 0.28 | 0.44 | -0.11 | 0.13 | 0.81 | -0.08 | 0.15 | 2.037 | 11.6 | 11.7 | 12 | 12 | 118.8 | 118.9 | LS |  | 2019-05-08 | 1.9774 | 0.9998 | 1.9991 | 1.019 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190515 | 0.43 | 0.6 | -0.25 | 0.22 | 1.1 | -0.21 | 0.22 | 1.982 | 8.8 | 8.9 | 9.5 | 6.9 | 119.9 | 120.1 | 0.12 | 0.44 | -0.03 | 0.15 | 0.75 | -0.06 | 0.14 | 2.005 | 11.5 | 11.5 | 11.8 | 12.1 | 119.7 | 120 | LS |  | 2019-05-15 | 1.9774 | 1.0023 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190528 | 0.47 | 0.69 | -0.28 | 0.2 | 1.23 | -0.26 | 0.23 | 1.986 | 8.4 | 9.2 | 9.7 | 7 | 119.9 | 119.9 | 0.16 | 0.43 | -0.04 | 0.14 | 0.76 | -0.06 | 0.14 | 2.005 | 11.7 | 11.5 | 11.8 | 12 | 119.7 | 119.9 | LS |  | 2019-05-28 | 1.9774 | 1.0043 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190618 | 0.46 | 0.7 | -0.29 | 0.31 | 1.21 | -0.28 | 0.26 | 1.987 | 8.8 | 8.9 | 9.4 | 6.9 | 119.9 | 120.1 | 0.1 | 0.35 | -0.03 | 0.07 | 0.73 | -0.03 | 0.09 | 2.013 | 11.5 | 11.4 | 11.6 | 11.9 | 119.7 | 120 | LS |  | 2019-06-18 | 1.9774 | 1.0049 | 1.9991 | 1.007 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190625 | 0.41 | 0.53 | -0.25 | 0.21 | 1.15 | -0.27 | 0.21 | 1.986 | 8.3 | 9.3 | 9.7 | 6.9 | 119.9 | 120.1 | 0.17 | 0.41 | -0.05 | 0.1 | 0.77 | -0.05 | 0.11 | 2.009 | 11.9 | 11.6 | 11.7 | 12 | 119.7 | 120 | LS |  | 2019-06-25 | 1.9774 | 1.0043 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190702 | 0.38 | 0.66 | -0.21 | 0.27 | 1.07 | -0.16 | 0.26 | 1.972 | 8.9 | 8.7 | 9.5 | 6.9 | 119.9 | 120.1 | 0.19 | 0.37 | -0.09 | 0.1 | 0.82 | -0.06 | 0.13 | 1.999 | 11.4 | 11.4 | 11.7 | 11.9 | 119.6 | 119.9 | LS |  | 2019-07-02 | 1.9774 | 0.99727 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190717 | 0.33 | 0.58 | -0.21 | 0.24 | 1.11 | -0.19 | 0.21 | 1.982 | 8.8 | 8.9 | 9.5 | 6.9 | 119.9 | 120.1 | 0.18 | 0.43 | -0.09 | 0.11 | 0.78 | -0.06 | 0.1 | 2.005 | 11.6 | 11.5 | 11.7 | 11.9 | 119.6 | 120 | LS |  | 2019-07-17 | 1.9774 | 1.0023 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190724 | 0.46 | 0.68 | -0.23 | 0.22 | 1.22 | -0.27 | 0.23 | 1.983 | 8.5 | 9.3 | 9.4 | 7.2 | 119.8 | 120.1 | 0.24 | 0.45 | -0.1 | 0.13 | 0.73 | -0.06 | 0.12 | 2.006 | 11.8 | 11.6 | 11.7 | 11.9 | 119.5 | 119.9 | LS |  | 2019-07-24 | 1.9774 | 1.0028 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190730 | 0.43 | 0.54 | -0.24 | 0.3 | 1.15 | -0.27 | 0.26 | 1.981 | 8.6 | 9.1 | 9.8 | 7 | 119.9 | 120 | 0.24 | 0.51 | -0.12 | 0.14 | 0.86 | -0.08 | 0.11 | 2.01 | 11.7 | 11.5 | 11.9 | 11.9 | 119.6 | 119.8 | LS |  | 2019-07-30 | 1.9774 | 1.0018 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190809 | 0.55 | 0.67 | -0.32 | 0.25 | 1.21 | -0.26 | 0.23 | 1.991 | 8.6 | 9.2 | 9.7 | 6.9 | 119.8 | 120 | 0.21 | 0.4 | -0.1 | 0.07 | 0.73 | -0.06 | 0.07 | 2.015 | 11.5 | 11.5 | 11.8 | 11.9 | 119.5 | 119.8 | LS |  | 2019-08-09 | 1.9774 | 1.0069 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190819 | 0.43 | 0.65 | -0.3 | 0.26 | 1.2 | -0.28 | 0.25 | 1.99 | 8.7 | 9.1 | 9.2 | 7.7 | 119.9 | 120.1 | [14 empty] | LS |  | 2019-08-19 | 1.9774 | 1.0064 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190905 | 0.48 | 0.78 | -0.27 | 0.3 | 1.26 | -0.24 | 0.3 | 1.989 | 8.6 | 9.1 | 9.9 | 6.9 | 119.8 | 119.9 | 0.2 | 0.47 | -0.07 | 0.15 | 0.87 | -0.03 | 0.14 | 2.01 | 11.6 | 11.5 | 11.9 | 12 | 119.6 | 119.8 | LS |  | 2019-09-05 | 1.9774 | 1.0059 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190912 | 0.59 | 0.68 | -0.33 | 0.21 | 1.29 | -0.28 | 0.19 | 1.991 | 9.3 | 8.8 | 9.9 | 7 | 119.8 | 119.7 | [14 empty] | LS |  | 2019-09-12 | 1.9774 | 1.0069 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20190920 | 0.59 | 0.58 | -0.3 | 0.15 | 1.2 | -0.24 | 0.2 | 1.986 | 9.4 | 8 | 9.2 | 7.4 | 119.9 | 120.1 | 0.21 | 0.44 | -0.1 | 0.16 | 0.86 | -0.08 | 0.15 | 2.01 | 11.6 | 11.7 | 11.8 | 12 | 119.6 | 119.9 | LS |  | 2019-09-20 | 1.9774 | 1.0043 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191002 | 0.47 | 0.55 | -0.27 | 0.2 | 1.1 | -0.26 | 0.18 | 1.976 | 8.2 | 9.4 | 9.5 | 7.1 | 119.8 | 120 | 0.18 | 0.41 | -0.05 | 0.1 | 0.75 | -0.02 | 0.11 | 1.998 | 11.8 | 11.6 | 11.7 | 11.9 | 119.6 | 119.9 | LS |  | 2019-10-02 | 1.9774 | 0.99929 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191008 | 0.44 | 0.66 | -0.2 | 0.24 | 1.2 | -0.27 | 0.21 | 1.986 | 8.1 | 9.5 | 9.6 | 6.9 | 119.8 | 119.9 | 0.18 | 0.38 | -0.05 | 0.09 | 0.74 | -0.05 | 0.07 | 2.01 | 11.8 | 11.6 | 11.7 | 11.9 | 119.5 | 119.9 | LS |  | 2019-10-08 | 1.9774 | 1.0043 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191014 | 0.41 | 0.7 | -0.22 | 0.26 | 1.17 | -0.26 | 0.25 | 1.987 | 8.4 | 9.2 | 9.8 | 6.9 | 119.9 | 120 | 0.16 | 0.44 | 0.03 | 0.14 | 0.84 | -0.03 | 0.12 | 2.004 | 11.5 | 11.4 | 11.8 | 12 | 119.7 | 120 | LS |  | 2019-10-14 | 1.9774 | 1.0049 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191017 | 0.43 | 0.7 | -0.23 | 0.27 | 1.22 | -0.27 | 0.26 | 1.982 | 8.6 | 9.2 | 9.6 | 7 | 119.8 | 119.9 | [14 empty] | LS |  | 2019-10-17 | 1.9774 | 1.0023 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191023 | 0.51 | 0.57 | -0.3 | 0.22 | 1.2 | -0.27 | 0.2 | 1.987 | 8.8 | 9 | 9.7 | 6.9 | 119.8 | 119.9 | 0.23 | 0.46 | -0.05 | 0.12 | 0.83 | -0.05 | 0.1 | 2.008 | 11.4 | 11.4 | 11.9 | 12.1 | 119.5 | 119.7 | LS |  | 2019-10-23 | 1.9774 | 1.0049 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191024 | 0.5 | 0.67 | -0.29 | 0.24 | 1.23 | -0.22 | 0.22 | 1.964 | 8.9 | 8.9 | 9.6 | 7 | 119.8 | 120 | 0.27 | 0.51 | -0.08 | 0.14 | 0.87 | -0.06 | 0.11 | 1.99 | 11.4 | 11.4 | 11.8 | 12.1 | 119.5 | 119.8 | LS |  | 2019-10-24 | 1.9774 | 0.99322 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191031 | 0.51 | 0.66 | -0.32 | 0.29 | 1.27 | -0.29 | 0.24 | 1.981 | 8.9 | 8.8 | 9.1 | 8 | 119.9 | 120 | 0.21 | 0.41 | -0.05 | 0.15 | 0.75 | -0.06 | 0.1 | 2.01 | 11.4 | 11.4 | 11.8 | 11.9 | 119.5 | 119.6 | LS |  | 2019-10-31 | 1.9774 | 1.0018 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191108 | 0.31 | 0.72 | -0.21 | 0.28 | 1.17 | -0.19 | 0.26 | 1.977 | 9 | 8.7 | 9.6 | 7.2 | 119.9 | 119.9 | 0.19 | 0.43 | -0.06 | 0.16 | 0.82 | -0.05 | 0.13 | 2.001 | 11.5 | 11.5 | 11.8 | 12 | 119.5 | 119.7 | LS |  | 2019-11-08 | 1.9774 | 0.9998 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191114 | 0.53 | 0.51 | -0.28 | 0.21 | 1.2 | -0.21 | 0.25 | 1.959 | 10.1 | 7.1 | 9.5 | 7.2 | 119.4 | 119.9 | 0.22 | 0.49 | -0.05 | 0.16 | 0.85 | -0.09 | 0.13 | 1.99 | 11.7 | 12 | 11.8 | 12 | 119.3 | 119.7 | LS |  | 2019-11-14 | 1.9774 | 0.99069 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191126 | 0.62 | 0.66 | -0.38 | 0.22 | 1.23 | -0.32 | 0.19 | 1.992 | 8.8 | 9.1 | 9.5 | 7.2 | 119.8 | 119.9 | 0.24 | 0.42 | -0.1 | 0.13 | 0.76 | -0.07 | 0.11 | 2.016 | 11.6 | 11.5 | 11.7 | 11.9 | 119.5 | 119.7 | LS |  | 2019-11-26 | 1.9774 | 1.0074 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191205 | 0.56 | 0.58 | -0.35 | 0.2 | 1.18 | -0.32 | 0.2 | 1.983 | 8.9 | 9 | 9.6 | 7 | 119.8 | 119.9 | 0.27 | 0.42 | -0.15 | 0.22 | 0.79 | -0.12 | 0.17 | 2.007 | 11.4 | 11.4 | 11.9 | 12 | 119.6 | 119.7 | LS |  | 2019-12-05 | 1.9774 | 1.0028 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191210 | 0.6 | 0.64 | -0.36 | 0.22 | 1.26 | -0.29 | 0.23 | 1.966 | 9.1 | 8.8 | 9.2 | 7.8 | 119.8 | 120 | 0.26 | 0.49 | -0.11 | 0.15 | 0.8 | -0.1 | 0.15 | 1.995 | 11.4 | 11.4 | 11.7 | 11.9 | 119.5 | 119.8 | LS |  | 2019-12-10 | 1.9774 | 0.99423 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191220 | 0.5 | 0.67 | -0.34 | 0.2 | 1.31 | -0.29 | 0.23 | 1.983 | 8.8 | 9 | 9.8 | 6.9 | 119.8 | 119.7 | 0.28 | 0.49 | -0.11 | 0.16 | 0.84 | -0.1 | 0.15 | 2.004 | 11.4 | 11.4 | 11.8 | 11.8 | 119.5 | 119.6 | LS |  | 2019-12-20 | 1.9774 | 1.0028 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20191229 | 0.61 | 0.66 | -0.38 | 0.22 | 1.43 | -0.36 | 0.22 | 1.978 | 8.2 | 9.5 | 9.8 | 7 | 119.8 | 119.8 | 0.33 | 0.43 | -0.15 | 0.12 | 0.86 | -0.11 | 0.13 | 1.997 | 11.8 | 11.5 | 11.9 | 11.9 | 119.5 | 119.6 | LS |  | 2019-12-29 | 1.9774 | 1.0003 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200104 | 0.45 | 0.62 | -0.2 | 0.18 | 1.26 | -0.21 | 0.22 | 1.983 | 9.3 | 8.3 | 9.9 | 6.9 | 119.8 | 119.8 | 0.25 | 0.5 | -0.13 | 0.17 | 0.8 | -0.12 | 0.16 | 2.009 | 11.5 | 11.6 | 11.9 | 12 | 119.5 | 119.5 | LS |  | 2020-01-04 | 1.9774 | 1.0028 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200115 | 0.53 | 0.54 | -0.31 | 0.2 | 1.14 | -0.28 | 0.2 | 1.978 | 8.5 | 9 | 9.9 | 6.9 | 119.9 | 119.8 | 0.32 | 0.48 | -0.19 | 0.13 | 0.81 | -0.1 | 0.15 | 2 | 11.4 | 11.4 | 11.9 | 11.8 | 119.6 | 119.6 | LS |  | 2020-01-15 | 1.9774 | 1.0003 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200122 | 0.34 | 0.63 | -0.22 | 0.18 | 1.07 | -0.27 | 0.21 | 1.975 | 8.7 | 9 | 9.2 | 7.6 | 119.9 | 120 | 0.31 | 0.41 | -0.1 | 0.1 | 0.75 | -0.12 | 0.1 | 1.998 | 11.4 | 11.4 | 11.7 | 11.9 | 119.5 | 119.8 | LS, after AI upgrade and iso change |  | 2020-01-22 | 1.9774 | 0.99879 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200129 | 0.37 | 0.62 | -0.26 | 0.24 | 1.06 | -0.28 | 0.25 | 1.982 | 8.8 | 9 | 8.9 | 8 | 119.8 | 120 | 0.22 | 0.45 | -0.05 | 0.16 | 0.79 | -0.11 | 0.15 | 2.007 | 11.4 | 11.4 | 11.7 | 11.7 | 119.6 | 119.8 | LS |  | 2020-01-29 | 1.9774 | 1.0023 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200203 | 0.39 | 0.6 | -0.24 | 0.27 | 1.12 | -0.27 | 0.27 | 1.991 | 9 | 8.7 | 8.9 | 8 | 119.8 | 120 | 0.19 | 0.35 | -0.06 | 0.1 | 0.79 | -0.11 | 0.12 | 2.015 | 11.3 | 11.3 | 11.6 | 11.7 | 119.5 | 119.8 | LS |  | 2020-02-03 | 1.9774 | 1.0069 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200213 | 0.33 | 0.68 | -0.19 | 0.21 | 0.98 | -0.22 | 0.21 | 1.986 | 9.1 | 8.7 | 8.6 | 8.4 | 119.8 | 120 | 0.23 | 0.46 | -0.03 | 0.15 | 0.77 | -0.08 | 0.11 | 2.007 | 11.5 | 11.3 | 11.5 | 11.2 | 119.5 | 119.7 | LS |  | 2020-02-13 | 1.9774 | 1.0043 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200214 | 0.39 | 0.57 | -0.23 | 0.18 | 1 | -0.26 | 0.18 | 1.984 | 8.8 | 9 | 8.9 | 8 | 119.8 | 120 | 0.23 | 0.39 | -0.08 | 0.1 | 0.76 | -0.14 | 0.09 | 2.005 | 11.3 | 11.4 | 11.6 | 11.7 | 119.5 | 119.8 | LS |  | 2020-02-14 | 1.9774 | 1.0033 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200220 | 0.47 | 0.59 | -0.32 | 0.25 | 0.98 | -0.3 | 0.24 | 1.984 | 8.5 | 9.2 | 9.1 | 7.7 | 119.8 | 120 | 0.41 | 0.4 | -0.12 | 0.16 | 0.83 | -0.1 | 0.13 | 2.005 | 11.2 | 11.3 | 11.6 | 11.5 | 119.5 | 119.7 | LS |  | 2020-02-20 | 1.9774 | 1.0033 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200227 | 0.42 | 0.69 | -0.27 | 0.23 | 1.05 | -0.29 | 0.23 | 1.989 | 8.7 | 9.1 | 9.2 | 7.8 | 119.8 | 120 | 0.25 | 0.39 | -0.1 | 0.11 | 0.81 | -0.13 | 0.1 | 2.01 | 11.4 | 11.4 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-02-27 | 1.9774 | 1.0059 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200311 | 0.41 | 0.62 | -0.27 | 0.2 | 1.06 | -0.3 | 0.23 | 1.982 | 8.6 | 9.1 | 8.9 | 7.9 | 119.8 | 120 | 0.26 | 0.43 | -0.08 | 0.1 | 0.77 | -0.11 | 0.12 | 2.005 | 11.3 | 11.4 | 11.6 | 11.7 | 119.6 | 119.8 | LS |  | 2020-03-11 | 1.9774 | 1.0023 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200318 | 0.43 | 0.64 | -0.26 | 0.21 | 1.08 | -0.29 | 0.22 | 1.989 | 8.7 | 9.1 | 9.2 | 7.8 | 119.8 | 120 | 0.28 | 0.46 | -0.08 | 0.1 | 0.79 | -0.1 | 0.12 | 2.014 | 11.4 | 11.4 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-03-18 | 1.9774 | 1.0059 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200326 | 0.3 | 0.65 | -0.22 | 0.23 | 1.03 | -0.25 | 0.27 | 1.984 | 8.8 | 8.9 | 9.1 | 7.7 | 119.8 | 120 | 0.26 | 0.45 | -0.07 | 0.15 | 0.85 | -0.11 | 0.17 | 2.009 | 11.4 | 11.3 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-03-26 | 1.9774 | 1.0033 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200401 | 0.39 | 0.55 | -0.26 | 0.21 | 1.02 | -0.32 | 0.21 | 1.985 | 8.5 | 9.3 | 9.2 | 7.7 | 119.8 | 120 | 0.29 | 0.37 | -0.09 | 0.07 | 0.74 | -0.09 | 0.09 | 2.011 | 11.3 | 11.5 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-04-01 | 1.9774 | 1.0038 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200408 | 0.3 | 0.65 | -0.2 | 0.22 | 0.94 | -0.24 | 0.26 | 1.983 | 8.4 | 9.3 | 9.1 | 8 | 119.8 | 119.9 | 0.31 | 0.46 | -0.09 | 0.14 | 0.83 | -0.11 | 0.16 | 2.002 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-04-08 | 1.9774 | 1.0028 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200417 | 0.38 | 0.67 | -0.27 | 0.19 | 1.04 | -0.31 | 0.2 | 1.981 | 8.5 | 9.2 | 9.2 | 7.5 | 119.8 | 120 | 0.27 | 0.47 | -0.1 | 0.1 | 0.86 | -0.13 | 0.11 | 2.004 | 11.4 | 11.5 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-04-17 | 1.9774 | 1.0018 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200429 | 0.3 | 0.68 | -0.2 | 0.27 | 1.12 | -0.25 | 0.29 | 1.971 | 8.6 | 9.1 | 9 | 7.9 | 119.8 | 120 | 0.28 | 0.35 | -0.1 | 0.11 | 0.81 | -0.09 | 0.17 | 1.997 | 11.4 | 11.4 | 11.7 | 11.8 | 119.5 | 119.7 | LS |  | 2020-04-29 | 1.9774 | 0.99676 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200506 | 0.43 | 0.61 | -0.28 | 0.21 | 1.03 | -0.31 | 0.23 | 1.984 | 8.6 | 9.2 | 9 | 8.1 | 119.8 | 120 | 0.27 | 0.48 | -0.08 | 0.14 | 0.82 | -0.13 | 0.13 | 2.001 | 11.4 | 11.5 | 11.7 | 11.8 | 119.5 | 119.8 | LS |  | 2020-05-06 | 1.9774 | 1.0033 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200513 | 0.35 | 0.67 | -0.22 | 0.23 | 0.94 | -0.23 | 0.21 | 1.982 | 8.6 | 9.1 | 8.9 | 8.1 | 119.8 | 120 | 0.2 | 0.44 | -0.03 | 0.13 | 0.76 | -0.09 | 0.13 | 2.008 | 11.6 | 11.6 | 11.8 | 12 | 119.5 | 119.7 | GB |  | 2020-05-13 | 1.9774 | 1.0023 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200519 | 0.4 | 0.57 | -0.27 | 0.23 | 0.98 | -0.31 | 0.23 | 1.977 | 8.5 | 9.2 | 8.6 | 8.3 | 119.8 | 120.1 | 0.27 | 0.44 | -0.1 | 0.12 | 0.78 | -0.12 | 0.11 | 2 | 11.5 | 11.5 | 11.6 | 11.7 | 119.5 | 119.8 | AA |  | 2020-05-19 | 1.9774 | 0.9998 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200525 | 0.35 | 0.64 | -0.25 | 0.2 | 1.04 | -0.29 | 0.2 | 1.987 | 8.1 | 9.5 | 9 | 8.1 | 119.7 | 120 | 0.24 | 0.39 | -0.09 | 0.08 | 0.77 | -0.11 | 0.11 | 2.006 | 11.9 | 11.6 | 11.6 | 11.6 | 119.5 | 119.8 | LS |  | 2020-05-25 | 1.9774 | 1.0049 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200604 | 0.36 | 0.62 | -0.23 | 0.16 | 1.06 | -0.29 | 0.2 | 1.985 | 8 | 9.6 | 9.1 | 8 | 119.7 | 119.9 | 0.23 | 0.48 | 0 | 0.02 | 0.85 | -0.09 | 0.03 | 2.006 | 11.7 | 11.6 | 11.7 | 11.9 | 119.4 | 119.7 | AA |  | 2020-06-04 | 1.9774 | 1.0038 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200611 | 0.34 | 0.58 | -0.24 | 0.21 | 0.99 | -0.27 | 0.2 | 1.986 | 8.9 | 9 | 8.8 | 8.2 | 119.8 | 120 | 0.26 | 0.39 | -0.07 | 0.07 | 0.77 | -0.1 | 0.08 | 2.011 | 11.8 | 11.6 | 11.5 | 11.2 | 119.4 | 119.7 | LS |  | 2020-06-11 | 1.9774 | 1.0043 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200616 | 0.15 | 0.76 | -0.1 | 0.3 | 1.06 | -0.15 | 0.33 | 1.98 | 8.7 | 8.9 | 8.3 | 8.3 | 119.9 | 120.1 | 0.17 | 0.36 | -0.05 | 0.03 | 0.73 | -0.07 | 0.08 | 2.007 | 11.3 | 11.4 | 11.5 | 11.3 | 119.6 | 119.9 | AA |  | 2020-06-16 | 1.9774 | 1.0013 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200618 | 0.37 | 0.56 | -0.2 | 0.2 | 0.99 | -0.23 | 0.22 | 1.983 | 8.7 | 9.1 | 8.9 | 7.7 | 119.8 | 120.1 | 0.25 | 0.4 | -0.06 | 0.09 | 0.73 | -0.08 | 0.09 | 2.008 | 11.4 | 11.4 | 11.7 | 11.8 | 119.4 | 119.8 | AA |  | 2020-06-18 | 1.9774 | 1.0028 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200625 | 0.44 | 0.67 | -0.26 | 0.23 | 1.07 | -0.29 | 0.23 | 1.983 | 8.6 | 9.2 | 8.9 | 8 | 119.8 | 120 | 0.23 | 0.48 | -0.05 | 0.16 | 0.77 | -0.1 | 0.14 | 2.009 | 11.4 | 11.5 | 11.7 | 11.7 | 119.4 | 119.7 | LS |  | 2020-06-25 | 1.9774 | 1.0028 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200701 | 0.38 | 0.58 | -0.21 | 0.24 | 0.92 | -0.28 | 0.2 | 1.974 | 8.3 | 9.2 | 9.1 | 7.3 | 119.8 | 120.1 | 0.24 | 0.33 | -0.05 | 0.05 | 0.77 | -0.08 | 0.05 | 2.002 | 11.5 | 11.5 | 11.7 | 11.8 | 119.5 | 119.8 | AA |  | 2020-07-01 | 1.9774 | 0.99828 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200706 | 0.32 | 0.61 | -0.23 | 0.19 | 0.96 | -0.26 | 0.24 | 1.982 | 8.3 | 9.3 | 9.1 | 7.5 | 119.8 | 120.1 | 0.3 | 0.41 | -0.11 | 0.1 | 0.8 | -0.08 | 0.14 | 2.006 | 11.4 | 11.5 | 11.7 | 11.7 | 119.5 | 119.8 | LS |  | 2020-07-06 | 1.9774 | 1.0023 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200707 | 0.25 | 0.3 | 0.13 | 0.07 | 0.42 | 0.07 | 0.08 | 2.003 | 8.8 | 9.3 | 9 | 8.5 | 119.6 | 119.9 | [14 empty] | LS |  | 2020-07-07 | 1.9774 | 1.0129 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200709 | 0.43 | 0.61 | -0.27 | 0.2 | 0.95 | -0.28 | 0.19 | 1.982 | 8.1 | 9.5 | 8.9 | 7.8 | 119.7 | 120.1 | 0.24 | 0.4 | -0.09 | 0.08 | 0.77 | -0.11 | 0.09 | 2.008 | 11.6 | 11.5 | 11.5 | 11.3 | 119.4 | 119.8 | LS |  | 2020-07-09 | 1.9774 | 1.0023 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200728 | 0.46 | 0.61 | -0.26 | 0.24 | 1.01 | -0.29 | 0.25 | 1.98 | 8.2 | 9.4 | 8.7 | 7.9 | 119.8 | 120.1 | 0.2 | 0.36 | 0 | 0.07 | 0.76 | -0.05 | 0.07 | 2.007 | 11.6 | 11.5 | 11.5 | 11.4 | 119.4 | 119.9 | CVD |  | 2020-07-28 | 1.9774 | 1.0013 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200806 | 0.4 | 0.56 | -0.21 | 0.25 | 1.11 | -0.24 | 0.29 | 1.978 | 9.1 | 8.6 | 8.5 | 8 | 119.8 | 120.1 | 0.15 | 0.32 | -0.03 | 0.07 | 0.76 | -0.09 | 0.08 | 2.002 | 11.4 | 11.2 | 11.4 | 11.3 | 119.5 | 119.9 | LS |  | 2020-08-06 | 1.9774 | 1.0003 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200813 | 0.36 | 0.59 | -0.24 | 0.2 | 0.95 | -0.28 | 0.19 | 1.977 | 8.2 | 9.4 | 9 | 7.6 | 119.7 | 120.1 | 0.23 | 0.48 | -0.06 | 0.14 | 0.79 | -0.11 | 0.1 | 2 | 11.7 | 11.5 | 11.6 | 11.8 | 119.5 | 119.9 | LS |  | 2020-08-13 | 1.9774 | 0.9998 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200820 | 0.33 | 0.47 | -0.22 | 0.17 | 0.96 | -0.24 | 0.21 | 1.981 | 9 | 8.8 | 8.8 | 7.9 | 119.8 | 120.1 | 0.28 | 0.42 | -0.09 | 0.1 | 0.81 | -0.1 | 0.11 | 2.01 | 11.4 | 11.3 | 11.6 | 11.7 | 119.4 | 119.8 | LS |  | 2020-08-20 | 1.9774 | 1.0018 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200827 | 0.29 | 0.65 | -0.19 | 0.23 | 1 | -0.22 | 0.26 | 1.988 | 8.6 | 9.1 | 8.7 | 8 | 119.8 | 120.1 | 0.21 | 0.44 | -0.06 | 0.12 | 0.81 | -0.1 | 0.13 | 2.012 | 11.2 | 11.4 | 11.6 | 11.6 | 119.5 | 119.8 | AD |  | 2020-08-27 | 1.9774 | 1.0054 | 1.9991 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200904 | 0.33 | 0.66 | -0.23 | 0.24 | 0.95 | -0.27 | 0.23 | 1.984 | 8.4 | 9.3 | 8.7 | 8 | 119.7 | 120.1 | 0.27 | 0.45 | -0.09 | 0.12 | 0.77 | -0.08 | 0.11 | 2.011 | 11.5 | 11.5 | 11.6 | 11.6 | 119.4 | 119.7 | AD |  | 2020-09-04 | 1.9774 | 1.0033 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200911 | 0.44 | 0.55 | -0.28 | 0.22 | 0.95 | -0.33 | 0.2 | 1.988 | 8.2 | 9.4 | 9 | 7.7 | 119.7 | 120.1 | 0.21 | 0.43 | -0.07 | 0.15 | 0.83 | -0.13 | 0.14 | 2.014 | 11.4 | 11.5 | 11.6 | 11.7 | 119.5 | 119.8 | AD |  | 2020-09-11 | 1.9774 | 1.0054 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200917 | 0.36 | 0.7 | -0.22 | 0.28 | 1.11 | -0.22 | 0.32 | 1.996 | 8.8 | 9 | 8.2 | 8.5 | 119.8 | 120.1 | 0.21 | 0.42 | -0.06 | 0.09 | 0.75 | -0.1 | 0.08 | 2.024 | 11.2 | 11.3 | 11.5 | 11.4 | 119.5 | 119.8 | AD |  | 2020-09-17 | 1.9774 | 1.0094 | 1.9991 | 1.0125 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20200924 | 0.43 | 0.65 | -0.29 | 0.23 | 0.96 | -0.3 | 0.22 | 1.987 | 8.6 | 9.2 | 8.8 | 8.1 | 119.7 | 120 | 0.29 | 0.45 | -0.11 | 0.11 | 0.8 | -0.13 | 0.09 | 2.009 | 11.3 | 11.4 | 11.6 | 11.6 | 119.4 | 119.8 | AD |  | 2020-09-24 | 1.9774 | 1.0049 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201001 | 0.37 | 0.64 | -0.24 | 0.23 | 0.94 | -0.27 | 0.22 | 1.987 | 8.5 | 9.3 | 8.9 | 7.9 | 119.7 | 120 | 0.34 | 0.45 | -0.1 | 0.13 | 0.81 | -0.1 | 0.12 | 2.017 | 11.3 | 11.5 | 11.6 | 11.6 | 119.3 | 119.7 | AD |  | 2020-10-01 | 1.9774 | 1.0049 | 1.9991 | 1.009 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201008 | 0.49 | 0.74 | -0.35 | 0.31 | 1.18 | -0.31 | 0.34 | 1.987 | 9.1 | 8.8 | 8.5 | 8.4 | 119.7 | 120.1 | 0.31 | 0.48 | -0.16 | 0.13 | 0.82 | -0.16 | 0.14 | 2.016 | 11.4 | 11.3 | 11.5 | 11.4 | 119.4 | 119.8 | AD |  | 2020-10-08 | 1.9774 | 1.0049 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201011 | 0.41 | 0.61 | -0.22 | 0.25 | 1.06 | -0.28 | 0.2 | 1.986 | 8.9 | 8.9 | 9 | 7.7 | 119.8 | 120.1 | 0.23 | 0.42 | -0.05 | 0.12 | 0.83 | -0.11 | 0.08 | 2.014 | 11.6 | 11.5 | 11.7 | 11.8 | 119.4 | 119.7 | AD |  | 2020-10-11 | 1.9774 | 1.0043 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201022 | 0.35 | 0.53 | -0.21 | 0.21 | 0.96 | -0.26 | 0.22 | 1.989 | 8.3 | 9.4 | 9 | 8 | 119.7 | 119.9 | 0.27 | 0.47 | -0.11 | 0.13 | 0.82 | -0.12 | 0.14 | 2.01 | 11.6 | 11.5 | 11.7 | 11.8 | 119.4 | 119.6 | ML |  | 2020-10-22 | 1.9774 | 1.0059 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201029 | 0.4 | 0.62 | -0.27 | 0.25 | 1.01 | -0.29 | 0.25 | 1.981 | 8.4 | 9.4 | 9 | 8.1 | 119.7 | 119.9 | 0.25 | 0.43 | -0.06 | 0.14 | 0.85 | -0.09 | 0.13 | 1.998 | 11.6 | 11.6 | 11.7 | 11.8 | 119.5 | 119.7 | ML |  | 2020-10-29 | 1.9774 | 1.0018 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201106 | 0.43 | 0.66 | -0.28 | 0.24 | 1.02 | -0.29 | 0.21 | 1.977 | 8.9 | 8.9 | 8.9 | 8.3 | 119.8 | 119.9 | 0.2 | 0.47 | -0.04 | 0.14 | 0.74 | -0.1 | 0.1 | 2.002 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.6 | ML |  | 2020-11-06 | 1.9774 | 0.9998 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201112 | 0.51 | 0.65 | -0.3 | 0.34 | 1.18 | -0.27 | 0.34 | 1.984 | 9 | 8.8 | 8.3 | 8.7 | 119.8 | 119.9 | 0.28 | 0.39 | -0.11 | 0.09 | 0.78 | -0.12 | 0.1 | 2.006 | 11.5 | 11.4 | 11.6 | 11.5 | 119.5 | 119.8 | ML |  | 2020-11-12 | 1.9774 | 1.0033 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201118 | 0.53 | 0.55 | -0.32 | 0.2 | 1.08 | -0.36 | 0.17 | 1.988 | 8.5 | 9.2 | 8.8 | 8.1 | 119.8 | 120 | 0.3 | 0.29 | -0.05 | 0.04 | 0.75 | -0.11 | 0 | 2.009 | 11.7 | 11.6 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2020-11-18 | 1.9774 | 1.0054 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201126 | 0.69 | 0.77 | -0.31 | 0.51 | 1.18 | -0.31 | 0.46 | 1.967 | 8.7 | 8.8 | 8.1 | 8.3 | 119.9 | 120.2 | 0.38 | 0.36 | -0.16 | 0.17 | 0.67 | -0.13 | 0.2 | 1.997 | 11.4 | 11.4 | 11.4 | 11.5 | 119.5 | 119.9 | ML, Matrix_1 |  | 2020-11-26 | 1.9774 | 0.99474 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201204 | 0.76 | 0.75 | -0.33 | 0.49 | 1.17 | -0.29 | 0.42 | 1.968 | 8.9 | 8.7 | 8.3 | 8.5 | 119.8 | 120 | 0.3 | 0.35 | -0.08 | 0.16 | 0.6 | -0.11 | 0.16 | 1.993 | 11.4 | 11.4 | 11.4 | 11.5 | 119.6 | 119.8 | ML, Matrix_1 |  | 2020-12-04 | 1.9774 | 0.99525 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201210 | 0.79 | 0.83 | -0.4 | 0.49 | 1.27 | -0.32 | 0.45 | 1.978 | 8.3 | 9.2 | 9 | 7.8 | 119.8 | 119.9 | 0.33 | 0.47 | -0.11 | 0.27 | 0.73 | -0.12 | 0.22 | 2.007 | 11.4 | 11.4 | 11.5 | 11.5 | 119.5 | 119.7 | ML, Matrix_1 |  | 2020-12-10 | 1.9774 | 1.0003 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201217 | 0.79 | 0.72 | -0.4 | 0.34 | 1.13 | -0.33 | 0.32 | 1.978 | 8.1 | 9.4 | 8.8 | 8 | 119.8 | 120 | 0.26 | 0.37 | -0.07 | 0.17 | 0.57 | -0.11 | 0.15 | 2 | 11.7 | 11.5 | 11.5 | 11.7 | 119.5 | 119.6 | ML, Matrix_1 |  | 2020-12-17 | 1.9774 | 1.0003 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20201223 | 0.81 | 0.89 | -0.4 | 0.5 | 1.22 | -0.35 | 0.43 | 1.971 | 8.9 | 8.7 | 8.6 | 8.3 | 119.8 | 119.9 | 0.29 | 0.38 | -0.12 | -0.02 | 0.58 | -0.12 | -0.02 | 1.991 | 11.3 | 11.3 | 11.4 | 11.5 | 119.6 | 119.7 | ML, Matrix_1 |  | 2020-12-23 | 1.9774 | 0.99676 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210103 | 0.79 | 0.78 | -0.39 | 0.42 | 1.51 | -0.38 | 0.44 | 1.976 | 8.4 | 9.1 | 9 | 7.6 | 119.9 | 120 | 0.35 | 0.43 | -0.16 | -0.03 | 0.71 | -0.15 | 0.01 | 1.999 | 11.2 | 11.3 | 11.5 | 11.7 | 119.6 | 119.7 | ML, Matrix_1 |  | 2021-01-03 | 1.9774 | 0.99929 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210114 | 0.71 | 0.81 | -0.37 | 0.53 | 1.08 | -0.28 | 0.44 | 1.966 | 8.5 | 8.9 | 7.8 | 8.3 | 119.9 | 120.3 | 0.34 | 0.37 | -0.18 | 0.22 | 0.65 | -0.15 | 0.23 | 1.998 | 11.3 | 11.4 | 11.3 | 11.4 | 119.5 | 120 | ML, Matrix_1 |  | 2021-01-14 | 1.9774 | 0.99423 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210121 | 0.85 | 0.71 | -0.41 | 0.47 | 1.11 | -0.34 | 0.39 | 1.973 | 8.1 | 9.3 | 8.1 | 8.1 | 119.8 | 120.2 | 0.29 | 0.31 | -0.1 | 0.06 | 0.53 | -0.11 | 0.07 | 1.995 | 11.7 | 11.5 | 11.4 | 11.3 | 119.6 | 120 | ML, Matrix_1 |  | 2021-01-21 | 1.9774 | 0.99777 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210128 | 0.78 | 0.7 | -0.38 | 0.39 | 1.11 | -0.33 | 0.36 | 1.974 | 7.6 | 9.6 | 7.7 | 8.5 | 119.7 | 120.2 | 0.31 | 0.3 | -0.12 | 0.11 | 0.65 | -0.11 | 0.14 | 1.998 | 11.5 | 11.5 | 11.3 | 11.5 | 119.5 | 120 | ML, Matrix_1 |  | 2021-01-28 | 1.9774 | 0.99828 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210205 | 0.69 | 0.66 | -0.37 | 0.43 | 1.15 | -0.34 | 0.41 | 1.971 | 8.4 | 9.1 | 8.3 | 7.9 | 119.9 | 120.2 | 0.38 | 0.39 | -0.16 | 0.19 | 0.74 | -0.14 | 0.21 | 1.997 | 11.5 | 11.5 | 11.4 | 11.7 | 119.5 | 119.9 | ML, Matrix_1 |  | 2021-02-05 | 1.9774 | 0.99676 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210211 | 0.57 | 0.71 | -0.4 | 0.23 | 1.32 | -0.41 | 0.23 | 1.97 | 8.8 | 8.7 | 8.4 | 7.8 | 119.9 | 120.3 | 0.28 | 0.43 | 0.01 | 0.14 | 0.6 | -0.04 | 0.1 | 1.99 | 11.2 | 11.5 | 11.4 | 11.3 | 119.6 | 120.1 | ML, Matrix_1, rot 270 deg |  | 2021-02-11 | 1.9774 | 0.99626 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210215 | 0.81 | 0.71 | -0.41 | 0.47 | 1.16 | -0.38 | 0.42 | 1.977 | 8.1 | 9.4 | 7.9 | 8.5 | 119.8 | 120.2 | 0.27 | 0.32 | -0.11 | -0.03 | 0.53 | -0.12 | -0.02 | 1.991 | 11.7 | 11.5 | 11.6 | 11.4 | 119.5 | 119.9 | ML, Matrix_1 |  | 2021-02-15 | 1.9774 | 0.9998 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210226 | 0.69 | 0.68 | -0.33 | 0.47 | 1.18 | -0.32 | 0.46 | 1.974 | 8.4 | 9.2 | 8.5 | 7.8 | 119.8 | 120.2 | 0.37 | 0.28 | -0.16 | 0.11 | 0.64 | -0.14 | 0.14 | 1.996 | 11.4 | 11.4 | 11.3 | 11.6 | 119.5 | 120 | ML, Matrix_1 |  | 2021-02-26 | 1.9774 | 0.99828 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210304 | 0.86 | 0.7 | -0.45 | 0.39 | 1.2 | -0.36 | 0.35 | 1.978 | 8.4 | 9.2 | 8.7 | 7.6 | 119.9 | 120.2 | 0.27 | 0.3 | -0.1 | 0.06 | 0.56 | -0.1 | 0.09 | 2 | 11.6 | 11.4 | 11.4 | 11.9 | 119.6 | 120 | ML, Matrix_1 |  | 2021-03-04 | 1.9774 | 1.0003 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210306 | 0.73 | 0.83 | -0.36 | 0.54 | 1.32 | -0.33 | 0.49 | 1.964 | 8.7 | 9 | 7.9 | 8.5 | 119.8 | 120.2 | 0.31 | 0.4 | -0.13 | -0.05 | 0.61 | -0.13 | -0.04 | 1.978 | 11.4 | 11.4 | 11.7 | 11.5 | 119.6 | 120.1 | ML, Matrix_1 |  | 2021-03-06 | 1.9774 | 0.99322 | 1.9991 | 0.98945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210311 | 0.8 | 0.68 | -0.41 | 0.39 | 1.17 | -0.31 | 0.38 | 1.984 | 9.5 | 7.7 | 8.5 | 7.9 | 119.8 | 120.2 | 0.31 | 0.46 | -0.12 | 0.18 | 0.75 | -0.16 | 0.17 | 2.007 | 11.5 | 11.7 | 11.5 | 11.7 | 119.5 | 119.9 | ML, Matrix_1 |  | 2021-03-11 | 1.9774 | 1.0033 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210323 | 0.8 | 0.69 | -0.37 | 0.43 | 1.08 | -0.32 | 0.41 | 1.974 | 9.2 | 8.2 | 8.1 | 8.2 | 119.8 | 120.2 | 0.33 | 0.39 | -0.14 | -0.05 | 0.55 | -0.14 | -0.01 | 1.992 | 11.4 | 11.5 | 11.5 | 11.4 | 119.6 | 120 | ML, Matrix_1 |  | 2021-03-23 | 1.9774 | 0.99828 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210323 | 0.44 | 0.73 | -0.23 | 0.36 | 1.09 | -0.24 | 0.36 | 1.986 | 8.5 | 9.1 | 8.9 | 7.8 | 119.8 | 120.1 | 0.29 | 0.53 | -0.12 | 0.22 | 0.94 | -0.09 | 0.24 | 2.009 | 11.7 | 11.5 | 11.6 | 11.5 | 119.5 | 119.9 | ML, Matrix_2 |  | 2021-03-23 | 1.9774 | 1.0043 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210401 | 0.75 | 0.6 | -0.38 | 0.32 | 1.18 | -0.32 | 0.33 | 1.972 | 8.4 | 9.2 | 8.6 | 7.9 | 119.8 | 120.2 | 0.33 | 0.38 | -0.11 | 0.16 | 0.67 | -0.1 | 0.2 | 1.995 | 11.6 | 11.5 | 11.5 | 11.8 | 119.5 | 120 | ML, MatriXX_1 |  | 2021-04-01 | 1.9774 | 0.99727 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210409 | 0.51 | 0.83 | -0.24 | 0.38 | 1.16 | -0.24 | 0.36 | 1.978 | 8.6 | 9.3 | 9.1 | 7.9 | 119.7 | 120 | 0.28 | 0.57 | -0.1 | 0.24 | 0.88 | -0.06 | 0.23 | 1.999 | 11.7 | 11.6 | 11.8 | 11.8 | 119.4 | 119.8 | ML, MatriXX_2 |  | 2021-04-09 | 1.9774 | 1.0003 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210415 | 0.7 | 0.87 | -0.34 | 0.57 | 1.19 | -0.31 | 0.46 | 1.981 | 8.8 | 8.8 | 8.1 | 8.4 | 119.9 | 120.2 | 0.35 | 0.38 | -0.14 | 0.19 | 0.71 | -0.14 | 0.19 | 2.006 | 11.6 | 11.5 | 11.7 | 11.4 | 119.6 | 120 | ML, MatriXX_1 |  | 2021-04-15 | 1.9774 | 1.0018 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210421 | 0.71 | 0.68 | -0.42 | 0.35 | 1.47 | -0.49 | 0.31 | 1.974 | 8.9 | 9 | 8.2 | 8.3 | 119.8 | 120.3 | 0.37 | 0.57 | -0.13 | 0.21 | 1.04 | -0.21 | 0.15 | 1.998 | 11.6 | 11.5 | 11.8 | 11.4 | 119.6 | 120.1 | ML, MatriXX_2 |  | 2021-04-21 | 1.9774 | 0.99828 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.76 | 0.61 | -0.42 | 0.32 | 1.13 | -0.35 | 0.3 | 1.993 | 8.5 | 9.1 | 8.6 | 7.8 | 119.8 | 120.2 | 0.29 | 0.35 | -0.09 | 0.04 | 0.54 | -0.12 | 0.04 | 2.014 | 11.4 | 11.4 | 11.4 | 11.8 | 119.5 | 119.9 | ML, MatriXX_1 |  | 2021-04-30 | 1.9774 | 1.0079 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.93 | 0.59 | -0.5 | 0.39 | 1.19 | -0.42 | 0.35 | 1.983 | 8.3 | 9.2 | 8 | 8.3 | 119.8 | 120.2 | 0.34 | 0.36 | -0.11 | 0.02 | 0.56 | -0.13 | 0.02 | 2.003 | 11.7 | 11.5 | 11.4 | 11.4 | 119.6 | 120 | ML, MatriXX_1, run 1 |  | 2021-04-30 | 1.9774 | 1.0028 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.82 | 0.59 | -0.43 | 0.37 | 1.11 | -0.37 | 0.33 | 1.98 | 8.4 | 9.2 | 8.1 | 8.2 | 119.8 | 120.2 | 0.31 | 0.37 | -0.11 | -0.02 | 0.54 | -0.12 | -0.02 | 2.003 | 11.7 | 11.5 | 11.5 | 11.4 | 119.6 | 120 | ML, MatriXX_1, run 2 |  | 2021-04-30 | 1.9774 | 1.0013 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210430 | 0.8 | 0.58 | -0.44 | 0.37 | 1.16 | -0.4 | 0.34 | 1.982 | 8.4 | 9.2 | 8 | 8.2 | 119.8 | 120.2 | 0.31 | 0.38 | -0.11 | 0 | 0.54 | -0.11 | 0.02 | 2.004 | 11.7 | 11.5 | 11.4 | 11.4 | 119.6 | 120 | ML, MatriXX_1, run 3 |  | 2021-04-30 | 1.9774 | 1.0023 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210505 | 0.76 | 0.7 | -0.41 | 0.45 | 1.18 | -0.33 | 0.4 | 1.972 | 8.9 | 8.7 | 8.3 | 8.1 | 119.8 | 120.2 | 0.48 | 0.36 | -0.23 | 0.15 | 0.75 | -0.2 | 0.18 | 1.998 | 11.6 | 11.6 | 11.6 | 11.8 | 119.5 | 119.7 | ML, MatriXX_1 |  | 2021-05-05 | 1.9774 | 0.99727 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210509 | 0.78 | 0.61 | -0.42 | 0.39 | 1.14 | -0.34 | 0.36 | 1.963 | 8.9 | 8.7 | 7.8 | 8.4 | 119.9 | 120.2 | 0.27 | 0.44 | -0.13 | -0.09 | 0.57 | -0.14 | -0.05 | 1.975 | 11.5 | 11.5 | 11.7 | 11.5 | 119.6 | 120 | ML, Matrixx 1 |  | 2021-05-09 | 1.9774 | 0.99272 | 1.9991 | 0.98794 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210521 | 0.73 | 0.83 | -0.37 | 0.5 | 1.15 | -0.33 | 0.45 | 1.975 | 8.8 | 8.8 | 7.7 | 8.5 | 119.9 | 120.2 | 0.26 | 0.29 | -0.12 | 0.12 | 0.62 | -0.11 | 0.17 | 1.992 | 11.3 | 11.3 | 11.6 | 11.4 | 119.7 | 120 | ML, Matrixx 1 |  | 2021-05-21 | 1.9774 | 0.99879 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210604 | 0.41 | 0.78 | -0.21 | 0.32 | 1.11 | -0.22 | 0.34 | 1.995 | 8.5 | 9.3 | 9 | 7.9 | 119.7 | 120.1 | 0.33 | 0.58 | -0.08 | 0.25 | 0.9 | -0.06 | 0.25 | 2.019 | 11.6 | 11.6 | 11.8 | 11.8 | 119.4 | 119.7 | ML, Matrixx 2 |  | 2021-06-04 | 1.9774 | 1.0089 | 1.9991 | 1.01 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210607 | 0.63 | 0.68 | -0.3 | 0.47 | 1.12 | -0.32 | 0.42 | 1.975 | 8.2 | 9.3 | 8.1 | 8.4 | 119.8 | 120.2 | 0.27 | 0.44 | -0.07 | 0.22 | 0.61 | -0.11 | 0.19 | 2 | 11.8 | 11.6 | 11.6 | 11.3 | 119.6 | 120 | ML, Matrixx_1 |  | 2021-06-07 | 1.9774 | 0.99879 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210609 | 0.66 | 0.77 | -0.36 | 0.46 | 1.09 | -0.3 | 0.4 | 1.967 | 8.6 | 9 | 8 | 8.3 | 119.9 | 120.2 | 0.35 | 0.38 | -0.13 | 0.21 | 0.67 | -0.12 | 0.2 | 1.993 | 11.6 | 11.5 | 11.7 | 11.3 | 119.5 | 120 | CVD, Mx1 |  | 2021-06-09 | 1.9774 | 0.99474 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210610 | 0.74 | 0.72 | -0.38 | 0.44 | 1.19 | -0.32 | 0.4 | 1.976 | 7.8 | 9.5 | 8.4 | 8.3 | 119.8 | 120.1 | 0.31 | 0.38 | -0.11 | 0.14 | 0.67 | -0.13 | 0.14 | 2 | 11.9 | 11.6 | 11.5 | 11.4 | 119.5 | 119.9 | CVD, Mx1 |  | 2021-06-10 | 1.9774 | 0.99929 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210611 | 0.57 | 0.66 | -0.31 | 0.35 | 0.97 | -0.24 | 0.33 | 1.974 | 8.6 | 8.9 | 8.2 | 8.3 | 119.9 | 120.1 | 0.29 | 0.4 | -0.11 | 0.17 | 0.69 | -0.12 | 0.16 | 2 | 11.5 | 11.5 | 11.6 | 11.3 | 119.6 | 119.9 | CVD, Mx1 |  | 2021-06-11 | 1.9774 | 0.99828 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210616 | 0.81 | 0.74 | -0.4 | 0.46 | 1.15 | -0.34 | 0.41 | 1.962 | 8.5 | 9.1 | 8.3 | 8.1 | 119.8 | 120.2 | 0.22 | 0.38 | -0.09 | -0.01 | 0.52 | -0.09 | 0.01 | 1.977 | 11.5 | 11.5 | 11.3 | 11.7 | 119.6 | 120 | CVD, Mx1 |  | 2021-06-16 | 1.9774 | 0.99221 | 1.9991 | 0.98895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210618 | 0.7 | 0.93 | -0.35 | 0.57 | 1.2 | -0.32 | 0.48 | 1.977 | 8.2 | 9.3 | 8.5 | 8 | 119.8 | 120.2 | 0.42 | 0.38 | -0.2 | 0.21 | 0.73 | -0.19 | 0.2 | 2.001 | 11.6 | 11.5 | 11.4 | 11.7 | 119.5 | 119.9 | GB |  | 2021-06-18 | 1.9774 | 0.9998 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210621 | 0.73 | 0.75 | -0.42 | 0.48 | 1.17 | -0.36 | 0.43 | 1.981 | 8.4 | 9.1 | 8.1 | 8.3 | 119.8 | 120.2 | 0.37 | 0.43 | -0.12 | 0.18 | 0.7 | -0.15 | 0.17 | 2.005 | 11.4 | 11.4 | 11.5 | 11.3 | 119.6 | 120 | GB |  | 2021-06-21 | 1.9774 | 1.0018 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210623 | 0.64 | 0.68 | -0.33 | 0.42 | 1.12 | -0.29 | 0.41 | 1.976 | 8.2 | 9.2 | 8.2 | 8.1 | 119.9 | 120.2 | [14 empty] | GB |  | 2021-06-23 | 1.9774 | 0.99929 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210624 | 0.8 | 0.85 | -0.41 | 0.55 | 1.31 | -0.28 | 0.53 | 1.981 | 7.3 | 9.7 | 8.4 | 8.1 | 119.8 | 120.3 | 0.31 | 0.31 | -0.12 | 0.11 | 0.6 | -0.11 | 0.15 | 1.995 | 11.9 | 11.6 | 11.5 | 11.5 | 119.6 | 120.1 | GB |  | 2021-06-24 | 1.9774 | 1.0018 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210628 | 0.67 | 0.76 | -0.34 | 0.52 | 1.15 | -0.28 | 0.48 | 1.972 | 8.5 | 9.1 | 8 | 8.4 | 119.9 | 120.2 | 0.39 | 0.44 | -0.14 | 0.24 | 0.73 | -0.16 | 0.21 | 1.996 | 11.5 | 11.5 | 11.7 | 11.3 | 119.6 | 120 | GB |  | 2021-06-28 | 1.9774 | 0.99727 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210630 | 0.64 | 0.74 | -0.36 | 0.47 | 1.08 | -0.31 | 0.44 | 1.973 | 8.3 | 9.3 | 8.1 | 8.4 | 119.8 | 120.2 | 0.33 | 0.4 | -0.1 | 0.24 | 0.71 | -0.12 | 0.24 | 1.993 | 11.7 | 11.5 | 11.4 | 11.6 | 119.6 | 120 | GB |  | 2021-06-30 | 1.9774 | 0.99777 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210702 | 0.68 | 0.64 | -0.34 | 0.42 | 1.09 | -0.31 | 0.36 | 1.986 | 8.2 | 9.3 | 7.9 | 8.3 | 119.8 | 120.2 | 0.24 | 0.38 | -0.05 | 0.21 | 0.57 | -0.1 | 0.22 | 2.013 | 11.8 | 11.6 | 11.3 | 11.2 | 119.4 | 119.9 | GB |  | 2021-07-02 | 1.9774 | 1.0043 | 1.9991 | 1.007 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210705 | 0.88 | 0.89 | -0.49 | 0.58 | 1.24 | -0.33 | 0.5 | 1.975 | 8.5 | 9 | 8.4 | 8 | 119.8 | 120.2 | 0.46 | 0.44 | -0.2 | 0.24 | 0.74 | -0.19 | 0.24 | 2.002 | 11.4 | 11.4 | 11.3 | 11.3 | 119.5 | 120 | ML |  | 2021-07-05 | 1.9774 | 0.99879 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210707 | 0.67 | 0.78 | -0.35 | 0.57 | 1.22 | -0.33 | 0.51 | 1.976 | 9 | 8.6 | 8.2 | 8.2 | 119.8 | 120.2 | 0.3 | 0.33 | -0.11 | 0.18 | 0.61 | -0.12 | 0.2 | 1.996 | 11.3 | 11.3 | 11.3 | 11.5 | 119.6 | 120 | ML |  | 2021-07-07 | 1.9774 | 0.99929 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210708 | 0.66 | 0.8 | -0.38 | 0.44 | 1.16 | -0.31 | 0.39 | 1.985 | 8.1 | 9.4 | 8.2 | 8.3 | 119.8 | 120.1 | 0.37 | 0.33 | -0.12 | 0.17 | 0.63 | -0.12 | 0.16 | 2.012 | 11.7 | 11.5 | 11.4 | 11.2 | 119.5 | 119.9 | ML |  | 2021-07-08 | 1.9774 | 1.0038 | 1.9991 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210712 | 0.78 | 0.69 | -0.42 | 0.42 | 1.18 | -0.36 | 0.4 | 1.971 | 8.5 | 9.1 | 8.1 | 8.2 | 119.8 | 120.2 | 0.33 | 0.33 | -0.14 | 0.17 | 0.66 | -0.16 | 0.18 | 1.995 | 11.5 | 11.5 | 11.4 | 11.2 | 119.5 | 120 | ML |  | 2021-07-12 | 1.9774 | 0.99676 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210714 | 0.72 | 0.87 | -0.35 | 0.51 | 1.2 | -0.31 | 0.47 | 1.98 | 8.2 | 9.3 | 8.2 | 8.1 | 119.8 | 120.2 | 0.36 | 0.35 | -0.17 | 0.22 | 0.73 | -0.18 | 0.25 | 2.004 | 11.7 | 11.5 | 11.4 | 11.2 | 119.5 | 120 | ML |  | 2021-07-14 | 1.9774 | 1.0013 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210715 | 0.8 | 0.84 | -0.48 | 0.56 | 1.28 | -0.39 | 0.51 | 1.973 | 8.1 | 9.4 | 7.9 | 8.4 | 119.8 | 120.2 | 0.34 | 0.39 | -0.18 | -0.04 | 0.65 | -0.18 | -0.01 | 1.99 | 11.8 | 11.5 | 11.4 | 11.4 | 119.6 | 120 | CVD |  | 2021-07-15 | 1.9774 | 0.99777 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210716 | 0.7 | 0.68 | -0.37 | 0.42 | 1.15 | -0.32 | 0.4 | 1.974 | 7.8 | 9.5 | 8.2 | 8.1 | 119.8 | 120.2 | 0.29 | 0.39 | -0.1 | 0.18 | 0.68 | -0.14 | 0.19 | 1.998 | 11.8 | 11.6 | 11.5 | 11.3 | 119.6 | 120 | ML |  | 2021-07-16 | 1.9774 | 0.99828 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210720 | 0.68 | 0.89 | -0.35 | 0.61 | 1.33 | -0.33 | 0.52 | 1.973 | 7.9 | 9.4 | 8.3 | 8.2 | 119.8 | 120.2 | 0.36 | 0.39 | -0.14 | 0.23 | 0.73 | -0.15 | 0.25 | 1.997 | 11.8 | 11.6 | 11.6 | 11.3 | 119.5 | 120 | ML |  | 2021-07-20 | 1.9774 | 0.99777 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210721 | 0.63 | 0.71 | -0.34 | 0.46 | 1.16 | -0.33 | 0.4 | 1.977 | 8.9 | 8.7 | 8.4 | 8.2 | 119.8 | 120.2 | 0.45 | 0.42 | -0.22 | 0.23 | 0.78 | -0.22 | 0.26 | 1.997 | 11.4 | 11.5 | 11.5 | 11.7 | 119.5 | 119.9 | ML |  | 2021-07-21 | 1.9774 | 0.9998 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210722 | 0.62 | 0.74 | -0.35 | 0.49 | 1.17 | -0.3 | 0.47 | 1.971 | 8.2 | 9.3 | 7.9 | 8.5 | 119.8 | 120.2 | 0.33 | 0.21 | -0.17 | 0.05 | 0.58 | -0.15 | 0.12 | 1.986 | 11.8 | 11.5 | 11.7 | 11.4 | 119.6 | 120.1 | ML |  | 2021-07-22 | 1.9774 | 0.99676 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210723 | 0.61 | 0.77 | -0.31 | 0.44 | 1.16 | -0.32 | 0.4 | 1.977 | 8.3 | 9.2 | 8.2 | 8.3 | 119.9 | 120.2 | 0.41 | 0.43 | -0.18 | 0.26 | 0.8 | -0.2 | 0.26 | 1.989 | 11.9 | 11.5 | 11.6 | 11.9 | 119.6 | 120.1 | ML |  | 2021-07-23 | 1.9774 | 0.9998 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210726 | 0.71 | 0.74 | -0.35 | 0.51 | 1.17 | -0.33 | 0.47 | 1.978 | 8.7 | 9 | 8.1 | 8.4 | 119.9 | 120.2 | 0.38 | 0.37 | -0.12 | 0.2 | 0.68 | -0.18 | 0.18 | 2.001 | 11.4 | 11.5 | 11.6 | 11.3 | 119.6 | 120 | ML |  | 2021-07-26 | 1.9774 | 1.0003 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210728 | 0.69 | 0.93 | -0.39 | 0.68 | 1.26 | -0.33 | 0.57 | 1.975 | 8.8 | 8.8 | 6.9 | 9.1 | 119.9 | 120.1 | 0.39 | 0.46 | -0.18 | 0.3 | 0.79 | -0.21 | 0.29 | 1.999 | 11.3 | 11.4 | 11.9 | 11.4 | 119.5 | 119.9 | ML |  | 2021-07-28 | 1.9774 | 0.99879 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210729 | 0.69 | 0.85 | -0.4 | 0.58 | 1.36 | -0.38 | 0.53 | 1.974 | 8.2 | 9.3 | 8.3 | 8.2 | 119.8 | 120.2 | 0.39 | 0.27 | -0.21 | 0.08 | 0.73 | -0.2 | 0.11 | 1.994 | 11.7 | 11.5 | 11.4 | 11.3 | 119.5 | 120.1 | ML |  | 2021-07-29 | 1.9774 | 0.99828 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210730 | 0.74 | 0.79 | -0.42 | 0.53 | 1.27 | -0.35 | 0.49 | 1.972 | 8.5 | 9.1 | 8 | 8.4 | 119.8 | 120.2 | 0.36 | 0.36 | -0.14 | 0.19 | 0.64 | -0.15 | 0.22 | 1.998 | 11.4 | 11.4 | 11.7 | 11.3 | 119.5 | 120 | ML |  | 2021-07-30 | 1.9774 | 0.99727 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210802 | 0.76 | 0.75 | -0.43 | 0.52 | 1.38 | -0.41 | 0.49 | 1.978 | 8.5 | 9.1 | 8.3 | 8.3 | 119.8 | 120.2 | 0.37 | 0.42 | -0.15 | 0.23 | 0.77 | -0.17 | 0.24 | 2.007 | 11.4 | 11.4 | 11.4 | 11.2 | 119.5 | 120 | ML |  | 2021-08-02 | 1.9774 | 1.0003 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210804 | 0.67 | 0.68 | -0.36 | 0.43 | 1.18 | -0.34 | 0.41 | 1.986 | 8.7 | 9 | 8.1 | 8.4 | 119.8 | 120.2 | 0.32 | 0.4 | -0.1 | 0.21 | 0.68 | -0.14 | 0.2 | 2.001 | 11.4 | 11.4 | 11.5 | 11.3 | 119.5 | 120 | ML |  | 2021-08-04 | 1.9774 | 1.0043 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210811 | 0.64 | 0.74 | -0.31 | 0.52 | 1.33 | -0.36 | 0.5 | 1.988 | 8.5 | 9.1 | 8.2 | 8.3 | 119.8 | 120.2 | 0.38 | 0.42 | -0.14 | 0.22 | 0.76 | -0.16 | 0.23 | 2.01 | 11.4 | 11.4 | 11.4 | 11.2 | 119.5 | 120 | CVD |  | 2021-08-11 | 1.9774 | 1.0054 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210812 | 0.59 | 0.87 | -0.29 | 0.54 | 1.21 | -0.3 | 0.48 | 1.985 | 8.3 | 9.2 | 8.3 | 8.3 | 119.8 | 120.2 | 0.34 | 0.46 | -0.1 | 0.27 | 0.71 | -0.13 | 0.25 | 2.007 | 11.5 | 11.4 | 11.4 | 11.2 | 119.5 | 120 | ML |  | 2021-08-12 | 1.9774 | 1.0038 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210813 | 0.77 | 0.81 | -0.44 | 0.55 | 1.33 | -0.37 | 0.48 | 1.982 | 7.9 | 9.4 | 8 | 8.5 | 119.8 | 120.2 | 0.32 | 0.24 | -0.16 | 0.1 | 0.6 | -0.19 | 0.14 | 2.001 | 11.5 | 11.4 | 11.5 | 11.2 | 119.6 | 120.1 | ML |  | 2021-08-13 | 1.9774 | 1.0023 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210816 | 0.65 | 0.81 | -0.35 | 0.6 | 1.48 | -0.35 | 0.56 | 1.975 | 7.6 | 9.6 | 8.2 | 8.3 | 119.8 | 120.2 | 0.28 | 0.46 | -0.13 | 0.21 | 0.77 | -0.13 | 0.24 | 1.999 | 11.8 | 11.6 | 11.4 | 11.2 | 119.5 | 120 | ML |  | 2021-08-16 | 1.9774 | 0.99879 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210818 | 0.76 | 0.87 | -0.43 | 0.55 | 1.32 | -0.35 | 0.47 | 1.983 | 7.8 | 9.4 | 8.4 | 8.1 | 119.8 | 120.2 | 0.36 | 0.44 | -0.12 | 0.23 | 0.77 | -0.16 | 0.24 | 2.005 | 11.7 | 11.5 | 11.4 | 11.3 | 119.5 | 120 | CVD |  | 2021-08-18 | 1.9774 | 1.0028 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210820 | 0.76 | 0.85 | -0.4 | 0.52 | 1.34 | -0.36 | 0.46 | 1.971 | 8 | 9.3 | 8.4 | 8 | 119.8 | 120.2 | 0.32 | 0.5 | -0.14 | -0.12 | 0.67 | -0.14 | -0.06 | 1.983 | 11.8 | 11.5 | 11.2 | 11.5 | 119.6 | 120.1 | ML |  | 2021-08-20 | 1.9774 | 0.99676 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210820 | 0.76 | 0.85 | -0.4 | 0.52 | 1.34 | -0.36 | 0.46 | 1.971 | 8 | 9.3 | 8.4 | 8 | 119.8 | 120.2 | 0.3 | 0.47 | -0.15 | -0.14 | 0.71 | -0.14 | -0.07 | 1.982 | 11.8 | 11.5 | 11.2 | 11.6 | 119.6 | 120.1 | ML/drop not confirmed by measurements done with IC |  | 2021-08-20 | 1.9774 | 0.99676 | 1.9991 | 0.99145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210823 | 0.72 | 0.78 | -0.41 | 0.5 | 1.34 | -0.38 | 0.48 | 1.991 | 7.7 | 9.6 | 8.1 | 8.3 | 119.7 | 120.2 | 0.37 | 0.37 | -0.15 | 0.15 | 0.72 | -0.17 | 0.16 | 2.019 | 11.7 | 11.5 | 11.4 | 11.5 | 119.5 | 120 | ML |  | 2021-08-23 | 1.9774 | 1.0069 | 1.9991 | 1.01 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210826 | 0.8 | 0.75 | -0.48 | 0.52 | 1.26 | -0.39 | 0.47 | 1.99 | 8.2 | 9.3 | 8.1 | 8.3 | 119.8 | 120.2 | 0.36 | 0.44 | -0.17 | 0.23 | 0.75 | -0.17 | 0.23 | 2.005 | 11.6 | 11.4 | 11.4 | 11.5 | 119.6 | 120 | GB |  | 2021-08-26 | 1.9774 | 1.0064 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210827 | 0.75 | 0.78 | -0.41 | 0.51 | 1.22 | -0.37 | 0.46 | 1.979 | 8.9 | 8.8 | 8.2 | 8.4 | 119.8 | 120.2 | 0.35 | 0.47 | -0.13 | 0.25 | 0.77 | -0.17 | 0.24 | 2.009 | 11.3 | 11.3 | 11.4 | 11.2 | 119.5 | 120 | GB |  | 2021-08-27 | 1.9774 | 1.0008 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210830 | 0.77 | 0.83 | -0.45 | 0.59 | 1.28 | -0.35 | 0.51 | 1.986 | 7.6 | 9.6 | 8 | 8.4 | 119.7 | 120.2 | 0.37 | 0.57 | -0.16 | 0.37 | 0.83 | -0.17 | 0.35 | 2.003 | 11.8 | 11.5 | 11.6 | 11.2 | 119.5 | 120 | GB |  | 2021-08-30 | 1.9774 | 1.0043 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210901 | 0.63 | 0.72 | -0.32 | 0.4 | 1.25 | -0.36 | 0.38 | 1.982 | 8.8 | 8.7 | 8.1 | 8.2 | 119.9 | 120.2 | 0.4 | 0.48 | -0.14 | 0.31 | 0.82 | -0.18 | 0.3 | 2.01 | 11.4 | 11.3 | 11.4 | 11.5 | 119.6 | 120.1 | GB |  | 2021-09-01 | 1.9774 | 1.0023 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210902 | 0.75 | 0.8 | -0.37 | 0.47 | 1.15 | -0.3 | 0.41 | 1.979 | 8.8 | 8.8 | 8.1 | 8.2 | 119.8 | 120.2 | 0.29 | 0.42 | -0.13 | -0.1 | 0.69 | -0.15 | -0.07 | 1.992 | 11.2 | 11.2 | 11.2 | 11.7 | 119.7 | 120.1 | GB |  | 2021-09-02 | 1.9774 | 1.0008 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210903 | 0.66 | 0.71 | -0.34 | 0.47 | 1.13 | -0.35 | 0.4 | 1.984 | 8.5 | 9.1 | 8.3 | 8 | 119.8 | 120.2 | 0.39 | 0.4 | -0.15 | 0.22 | 0.76 | -0.17 | 0.22 | 2.012 | 11.5 | 11.4 | 11.3 | 11.2 | 119.5 | 120 | GB |  | 2021-09-03 | 1.9774 | 1.0033 | 1.9991 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210906 | 0.56 | 0.72 | -0.34 | 0.38 | 1.18 | -0.3 | 0.36 | 1.977 | 8.6 | 9.2 | 8.6 | 8.1 | 119.8 | 120.1 | 0.4 | 0.44 | -0.19 | 0.14 | 0.9 | -0.12 | 0.18 | 1.999 | 11.6 | 11.5 | 11.5 | 11.3 | 119.5 | 119.9 | ML, Matrixx 2 |  | 2021-09-06 | 1.9774 | 0.9998 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210907 | 0.89 | 0.65 | -0.51 | 0.44 | 1.28 | -0.43 | 0.41 | 1.98 | 8.5 | 9.1 | 8.2 | 8.3 | 119.8 | 120.2 | 0.44 | 0.34 | -0.18 | 0.16 | 0.75 | -0.21 | 0.17 | 2.001 | 11.5 | 11.5 | 11.5 | 11.4 | 119.5 | 119.9 | CVD |  | 2021-09-07 | 1.9774 | 1.0013 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210909 | 0.66 | 0.79 | -0.38 | 0.54 | 1.26 | -0.32 | 0.5 | 1.981 | 8.3 | 9.3 | 7.9 | 8.6 | 119.8 | 120.2 | 0.35 | 0.41 | -0.13 | 0.26 | 0.73 | -0.16 | 0.26 | 2.001 | 11.5 | 11.4 | 11.3 | 11.3 | 119.6 | 120 | GB |  | 2021-09-09 | 1.9774 | 1.0018 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210910 | 0.65 | 0.79 | -0.36 | 0.56 | 1.28 | -0.34 | 0.53 | 1.974 | 8.2 | 9.3 | 8.3 | 8.2 | 119.8 | 120.2 | 0.39 | 0.5 | -0.16 | 0.35 | 0.77 | -0.16 | 0.31 | 1.998 | 11.6 | 11.5 | 11.4 | 11.5 | 119.5 | 120 | GB |  | 2021-09-10 | 1.9774 | 0.99828 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210914 | 0.75 | 0.75 | -0.47 | 0.44 | 1.2 | -0.38 | 0.42 | 1.985 | 8.3 | 9.3 | 7.6 | 8.7 | 119.8 | 120.2 | 0.4 | 0.41 | -0.13 | 0.23 | 0.78 | -0.18 | 0.2 | 2.01 | 11.6 | 11.5 | 11.6 | 11.3 | 119.5 | 120 | ML |  | 2021-09-14 | 1.9774 | 1.0038 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210915 | 0.6 | 0.72 | -0.26 | 0.49 | 1.17 | -0.26 | 0.42 | 1.971 | 7.9 | 9.4 | 8.3 | 8.1 | 119.8 | 120.2 | 0.41 | 0.43 | -0.18 | 0.26 | 0.81 | -0.19 | 0.27 | 1.992 | 11.5 | 11.5 | 11.6 | 11.8 | 119.6 | 120 | GB, Matrixx 1 |  | 2021-09-15 | 1.9774 | 0.99676 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210916 | 0.44 | 0.53 | -0.26 | 0.31 | 0.96 | -0.28 | 0.26 | 1.986 | 8.8 | 8.8 | 8.1 | 8.3 | 119.9 | 120.2 | 0.28 | 0.29 | -0.1 | 0.17 | 0.47 | -0.16 | 0.1 | 2.011 | 11.3 | 11.4 | 11.6 | 11.2 | 119.6 | 120.1 | CVD |  | 2021-09-16 | 1.9774 | 1.0043 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210917 | 0.48 | 0.47 | -0.28 | 0.26 | 0.9 | -0.32 | 0.26 | 1.974 | 8.5 | 9.1 | 8.1 | 8.4 | 119.8 | 120.2 | 0.32 | 0.34 | -0.13 | 0.19 | 0.71 | -0.17 | 0.18 | 1.995 | 11.8 | 11.6 | 11.6 | 11.8 | 119.6 | 119.9 | GB |  | 2021-09-17 | 1.9774 | 0.99828 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210920 | 0.48 | 0.54 | -0.27 | 0.3 | 0.97 | -0.3 | 0.27 | 1.983 | 8.9 | 8.7 | 8.1 | 8.3 | 119.9 | 120.2 | 0.32 | 0.25 | -0.13 | 0.12 | 0.52 | -0.18 | 0.09 | 2.009 | 11.4 | 11.4 | 11.4 | 11.6 | 119.6 | 120 | ML |  | 2021-09-20 | 1.9774 | 1.0028 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210922 | 0.57 | 0.73 | -0.27 | 0.45 | 1.11 | -0.23 | 0.44 | 1.98 | 8.3 | 9.3 | 8.3 | 8.4 | 119.8 | 120 | 0.3 | 0.35 | -0.1 | 0.14 | 0.66 | -0.11 | 0.19 | 2.001 | 11.7 | 11.5 | 11.6 | 11.4 | 119.6 | 119.8 | GB |  | 2021-09-22 | 1.9774 | 1.0013 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210923 | 0.71 | 0.83 | -0.38 | 0.44 | 1.28 | -0.34 | 0.4 | 1.972 | 8 | 9.3 | 9 | 7.5 | 119.9 | 120.1 | 0.33 | 0.48 | -0.12 | 0.28 | 0.77 | -0.15 | 0.28 | 1.993 | 11.9 | 11.6 | 11.5 | 11.7 | 119.6 | 120 | ML |  | 2021-09-23 | 1.9774 | 0.99727 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210924 | 0.54 | 0.79 | -0.26 | 0.47 | 1.02 | -0.22 | 0.39 | 1.972 | 8.2 | 9.3 | 8.2 | 8.2 | 119.8 | 120.1 | 0.35 | 0.57 | -0.13 | 0.32 | 0.79 | -0.12 | 0.28 | 1.992 | 11.6 | 11.5 | 11.4 | 11.4 | 119.6 | 119.9 | ML |  | 2021-09-24 | 1.9774 | 0.99727 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20210927 | 0.66 | 0.78 | -0.33 | 0.48 | 1.23 | -0.3 | 0.44 | 1.977 | 8.6 | 9 | 8.2 | 8.5 | 119.8 | 120.1 | 0.31 | 0.42 | -0.09 | 0.24 | 0.7 | -0.13 | 0.22 | 2 | 11.6 | 11.5 | 11.7 | 11.4 | 119.6 | 119.9 | GB |  | 2021-09-27 | 1.9774 | 0.9998 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211006 | 0.66 | 0.82 | -0.38 | 0.55 | 1.23 | -0.33 | 0.48 | 1.976 | 8.9 | 8.7 | 7.8 | 8.8 | 119.9 | 120.1 | 0.36 | 0.49 | -0.13 | 0.25 | 0.77 | -0.15 | 0.24 | 1.997 | 11.5 | 11.6 | 11.6 | 11.5 | 119.6 | 120.1 | GB |  | 2021-10-06 | 1.9774 | 0.99929 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211007 | 0.71 | 0.78 | -0.44 | 0.5 | 1.17 | -0.34 | 0.45 | 1.978 | 7.9 | 9.5 | 7.8 | 8.6 | 119.8 | 120.2 | 0.38 | 0.56 | -0.14 | 0.37 | 0.8 | -0.15 | 0.33 | 2.002 | 11.9 | 11.6 | 11.9 | 11.4 | 119.5 | 120 | GB |  | 2021-10-07 | 1.9774 | 1.0003 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211008 | 0.81 | 0.84 | -0.53 | 0.52 | 1.22 | -0.36 | 0.47 | 1.982 | 8.3 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.38 | 0.46 | -0.14 | 0.31 | 0.76 | -0.15 | 0.27 | 2.002 | 11.8 | 11.5 | 11.8 | 11.3 | 119.6 | 120 | GB |  | 2021-10-08 | 1.9774 | 1.0023 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211011 | 0.68 | 0.84 | -0.44 | 0.55 | 1.21 | -0.35 | 0.49 | 1.985 | 8.4 | 9.2 | 8.1 | 8.5 | 119.8 | 120.1 | 0.37 | 0.45 | -0.13 | 0.26 | 0.78 | -0.16 | 0.23 | 2.001 | 11.7 | 11.5 | 11.8 | 11.3 | 119.6 | 120 | CVD |  | 2021-10-11 | 1.9774 | 1.0038 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211013 | 0.6 | 0.84 | -0.34 | 0.55 | 1.17 | -0.3 | 0.48 | 1.979 | 7.7 | 9.5 | 8.1 | 8.5 | 119.8 | 120.2 | 0.38 | 0.55 | -0.15 | 0.34 | 0.85 | -0.16 | 0.3 | 1.998 | 11.9 | 11.6 | 11.5 | 11.6 | 119.6 | 120 | CVD |  | 2021-10-13 | 1.9774 | 1.0008 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211014 | 0.68 | 0.77 | -0.38 | 0.47 | 1.15 | -0.33 | 0.41 | 1.979 | 7.7 | 9.6 | 8.5 | 8.1 | 119.8 | 120.2 | 0.41 | 0.52 | -0.14 | 0.32 | 0.82 | -0.16 | 0.28 | 2.001 | 12 | 11.6 | 11.5 | 11.4 | 119.5 | 120 | CVD |  | 2021-10-14 | 1.9774 | 1.0008 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211018 | 0.81 | 0.57 | -0.47 | 0.41 | 1.21 | -0.36 | 0.39 | 1.989 | 7.8 | 9.6 | 7.7 | 8.7 | 119.7 | 120.1 | 0.35 | 0.44 | -0.16 | 0.29 | 0.7 | -0.16 | 0.27 | 2.007 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | ML, Matrixx_1 |  | 2021-10-18 | 1.9774 | 1.0059 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211020 | 0.67 | 0.79 | -0.37 | 0.51 | 1.11 | -0.29 | 0.45 | 1.992 | 7.9 | 9.4 | 8.1 | 8.4 | 119.8 | 120.1 | 0.4 | 0.34 | -0.17 | 0.19 | 0.75 | -0.15 | 0.21 | 2.017 | 11.6 | 11.5 | 11.4 | 11.2 | 119.5 | 119.9 | ML |  | 2021-10-20 | 1.9774 | 1.0074 | 1.9991 | 1.009 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211021 | 0.68 | 0.97 | -0.45 | 0.63 | 1.26 | -0.36 | 0.53 | 1.981 | 8.5 | 9 | 6.9 | 9.8 | 119.8 | 119.6 | 0.4 | 0.46 | -0.16 | 0.3 | 0.78 | -0.18 | 0.26 | 1.992 | 11.1 | 11.2 | 11.7 | 11.4 | 119.6 | 119.8 | GB |  | 2021-10-21 | 1.9774 | 1.0018 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211025 | 0.67 | 0.81 | -0.36 | 0.48 | 1.19 | -0.35 | 0.42 | 1.991 | 8 | 9.4 | 8.1 | 8.5 | 119.8 | 120.1 | 0.37 | 0.45 | -0.12 | 0.25 | 0.74 | -0.16 | 0.21 | 2.009 | 11.6 | 11.5 | 11.4 | 11.3 | 119.6 | 119.9 | GB |  | 2021-10-25 | 1.9774 | 1.0069 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211027 | 0.76 | 0.81 | -0.43 | 0.55 | 1.26 | -0.36 | 0.49 | 1.987 | 7.9 | 9.5 | 8.5 | 8.4 | 119.7 | 119.9 | 0.37 | 0.49 | -0.13 | 0.31 | 0.72 | -0.15 | 0.28 | 2.007 | 11.6 | 11.5 | 11.3 | 11.2 | 119.5 | 119.7 | GB |  | 2021-10-27 | 1.9774 | 1.0049 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211028 | 0.75 | 0.74 | -0.43 | 0.56 | 1.21 | -0.39 | 0.45 | 1.993 | 8.7 | 8.9 | 8.1 | 8.7 | 119.8 | 120 | 0.39 | 0.54 | -0.14 | 0.32 | 0.89 | -0.18 | 0.27 | 2.016 | 11.3 | 11.4 | 11.4 | 11.3 | 119.5 | 119.7 | GB |  | 2021-10-28 | 1.9774 | 1.0079 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211029 | 0.69 | 0.6 | -0.37 | 0.43 | 1.17 | -0.31 | 0.4 | 1.984 | 7.2 | 9.7 | 7.6 | 8.8 | 119.6 | 120 | 0.26 | 0.37 | -0.08 | 0.17 | 0.61 | -0.12 | 0.17 | 2.007 | 11.8 | 11.6 | 11.5 | 11.4 | 119.4 | 119.8 | KA |  | 2021-10-29 | 1.9774 | 1.0033 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211101 | 0.67 | 0.83 | -0.38 | 0.51 | 1.09 | -0.33 | 0.41 | 1.987 | 7.6 | 9.5 | 8.2 | 8.5 | 119.7 | 120 | 0.33 | 0.45 | -0.1 | 0.25 | 0.75 | -0.14 | 0.21 | 2.008 | 11.8 | 11.6 | 11.3 | 11.2 | 119.5 | 119.8 | CVD |  | 2021-11-01 | 1.9774 | 1.0049 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211103 | 0.79 | 0.76 | -0.43 | 0.43 | 1.14 | -0.38 | 0.38 | 1.981 | 8.1 | 9.3 | 8 | 8.5 | 119.8 | 120.1 | 0.29 | 0.38 | -0.12 | 0.18 | 0.67 | -0.15 | 0.21 | 2.005 | 11.6 | 11.4 | 11.2 | 11.2 | 119.5 | 119.8 | GB |  | 2021-11-03 | 1.9774 | 1.0018 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211104 | 0.71 | 0.69 | -0.44 | 0.43 | 1.21 | -0.38 | 0.38 | 1.99 | 8.5 | 9.1 | 7.9 | 8.6 | 119.8 | 120.1 | 0.32 | 0.35 | -0.13 | 0.05 | 0.54 | -0.12 | 0.05 | 2.001 | 11.4 | 11.4 | 11.3 | 11.4 | 119.6 | 119.8 | ML |  | 2021-11-04 | 1.9774 | 1.0064 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211108 | 0.78 | 1.03 | -0.43 | 0.55 | 1.3 | -0.38 | 0.45 | 1.973 | 8.1 | 9.4 | 8.3 | 8.2 | 119.8 | 120.1 | 0.33 | 0.31 | -0.17 | 0.07 | 0.63 | -0.18 | 0.06 | 1.988 | 11.7 | 11.5 | 11.2 | 11.2 | 119.5 | 119.8 | GB/ wrong calib. Factor for Matrixx 1 |  | 2021-11-08 | 1.9774 | 0.99777 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211109 | 0.74 | 0.69 | -0.42 | 0.47 | 1.23 | -0.36 | 0.41 | 1.986 | 8.1 | 9.4 | 8 | 8.5 | 119.8 | 120.1 | 0.4 | 0.45 | -0.15 | 0.25 | 0.8 | -0.18 | 0.22 | 2.005 | 11.8 | 11.5 | 11.4 | 11.3 | 119.5 | 119.9 | ML |  | 2021-11-09 | 1.9774 | 1.0043 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211112 | 0.76 | 0.65 | -0.39 | 0.45 | 1.1 | -0.34 | 0.39 | 1.983 | 7.9 | 9.5 | 8.2 | 8.4 | 119.8 | 120.1 | 0.27 | 0.37 | -0.09 | 0.18 | 0.65 | -0.11 | 0.2 | 2.005 | 11.8 | 11.6 | 11.3 | 11.3 | 119.5 | 119.8 | ML |  | 2021-11-12 | 1.9774 | 1.0028 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211117 | 0.63 | 0.71 | -0.34 | 0.32 | 1.12 | -0.36 | 0.24 | 1.98 | 8.2 | 9.2 | 8.6 | 7.9 | 119.9 | 120.1 | 0.4 | 0.48 | -0.12 | 0.23 | 0.71 | -0.17 | 0.17 | 2.003 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | GB |  | 2021-11-17 | 1.9774 | 1.0013 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211119 | 0.75 | 0.72 | -0.45 | 0.48 | 1.15 | -0.4 | 0.4 | 1.991 | 8.3 | 9.2 | 8.1 | 8.4 | 119.8 | 120.1 | 0.35 | 0.37 | -0.13 | 0.15 | 0.66 | -0.16 | 0.16 | 2.013 | 11.6 | 11.5 | 11.3 | 11.3 | 119.5 | 119.8 | ML |  | 2021-11-19 | 1.9774 | 1.0069 | 1.9991 | 1.007 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211122 | 0.72 | 0.79 | -0.38 | 0.49 | 1.18 | -0.34 | 0.44 | 1.995 | 7.6 | 9.6 | 8.3 | 8.5 | 119.7 | 120 | 0.39 | 0.46 | -0.16 | 0.26 | 0.77 | -0.16 | 0.24 | 2.009 | 11.8 | 11.5 | 11.5 | 11.5 | 119.5 | 119.8 | ML |  | 2021-11-22 | 1.9774 | 1.0089 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211124 | 0.64 | 0.82 | -0.41 | 0.49 | 1.1 | -0.33 | 0.4 | 1.994 | 8 | 9.4 | 8 | 8.6 | 119.8 | 120 | 0.37 | 0.46 | -0.12 | 0.25 | 0.67 | -0.14 | 0.2 | 2.014 | 11.8 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2021-11-24 | 1.9774 | 1.0084 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211125 | 0.74 | 0.84 | -0.39 | 0.55 | 1.27 | -0.36 | 0.49 | 1.987 | 7.8 | 9.4 | 8.2 | 8.4 | 119.8 | 120.1 | 0.34 | 0.28 | -0.14 | 0.05 | 0.53 | -0.13 | 0.04 | 1.997 | 11.6 | 11.4 | 11.3 | 11.6 | 119.6 | 119.9 | CVD |  | 2021-11-25 | 1.9774 | 1.0049 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211126 | 0.7 | 0.7 | -0.38 | 0.44 | 1.09 | -0.33 | 0.39 | 1.991 | 8.2 | 9.3 | 8.5 | 8.2 | 119.8 | 120.1 | 0.38 | 0.39 | -0.15 | 0.19 | 0.72 | -0.15 | 0.21 | 2.008 | 11.7 | 11.5 | 11.4 | 11.3 | 119.5 | 119.8 | GB |  | 2021-11-26 | 1.9774 | 1.0069 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211129 | 0.66 | 0.85 | -0.39 | 0.47 | 1.12 | -0.33 | 0.37 | 1.993 | 8.2 | 9.3 | 8.4 | 8.4 | 119.8 | 120 | 0.3 | 0.54 | -0.09 | 0.31 | 0.74 | -0.15 | 0.25 | 2.006 | 11.7 | 11.5 | 11.5 | 11.5 | 119.6 | 119.9 | GB |  | 2021-11-29 | 1.9774 | 1.0079 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211202 | 0.74 | 0.81 | -0.45 | 0.47 | 1.21 | -0.4 | 0.41 | 1.998 | 8.3 | 9.3 | 8.3 | 8.5 | 119.8 | 120 | 0.43 | 0.46 | -0.18 | 0.24 | 0.79 | -0.2 | 0.21 | 2.013 | 11.5 | 11.5 | 11.4 | 11.3 | 119.5 | 119.8 | GB |  | 2021-12-02 | 1.9774 | 1.0104 | 1.9991 | 1.007 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211207 | 0.76 | 0.79 | -0.42 | 0.45 | 1.11 | -0.35 | 0.39 | 1.995 | 8.2 | 9.3 | 8.4 | 8.3 | 119.8 | 120 | 0.38 | 0.38 | -0.16 | 0.18 | 0.71 | -0.13 | 0.2 | 2.01 | 11.6 | 11.5 | 11.4 | 11.3 | 119.6 | 119.8 | GB |  | 2021-12-07 | 1.9774 | 1.0089 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211209 | 0.7 | 0.68 | -0.38 | 0.37 | 1.01 | -0.31 | 0.31 | 1.996 | 8 | 9.4 | 8.5 | 8.2 | 119.8 | 120 | 0.35 | 0.44 | -0.1 | 0.25 | 0.66 | -0.15 | 0.19 | 2.012 | 11.8 | 11.5 | 11.4 | 11.2 | 119.5 | 119.8 | ML |  | 2021-12-09 | 1.9774 | 1.0094 | 1.9991 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211210 | 0.73 | 0.72 | -0.45 | 0.42 | 1.15 | -0.35 | 0.37 | 1.995 | 8.1 | 9.3 | 8.2 | 8.6 | 119.8 | 120 | 0.35 | 0.45 | -0.13 | 0.26 | 0.73 | -0.15 | 0.23 | 2.013 | 11.7 | 11.5 | 11.4 | 11.4 | 119.6 | 119.8 | ML |  | 2021-12-10 | 1.9774 | 1.0089 | 1.9991 | 1.007 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211213 | 0.85 | 0.68 | -0.45 | 0.39 | 1.25 | -0.38 | 0.37 | 1.991 | 8.4 | 9.1 | 8.3 | 8.2 | 119.8 | 120.1 | 0.42 | 0.51 | -0.17 | 0.3 | 0.8 | -0.16 | 0.27 | 2.012 | 11.4 | 11.4 | 11.4 | 11.5 | 119.6 | 119.8 | ML |  | 2021-12-13 | 1.9774 | 1.0069 | 1.9991 | 1.0065 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211216 | 0.75 | 0.77 | -0.39 | 0.46 | 1.22 | -0.33 | 0.4 | 1.998 | 8.9 | 8.7 | 8.5 | 8.2 | 119.8 | 120 | 0.36 | 0.52 | -0.11 | 0.29 | 0.75 | -0.16 | 0.21 | 2.016 | 11.4 | 11.3 | 11.4 | 11.5 | 119.6 | 119.8 | GB |  | 2021-12-16 | 1.9774 | 1.0104 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211220 | 0.66 | 0.76 | -0.33 | 0.41 | 1.13 | -0.3 | 0.38 | 1.996 | 8.7 | 8.9 | 8.3 | 8.4 | 119.9 | 120 | 0.41 | 0.5 | -0.15 | 0.28 | 0.83 | -0.16 | 0.24 | 2.015 | 11.4 | 11.4 | 11.3 | 11.2 | 119.5 | 119.8 | CVD |  | 2021-12-20 | 1.9774 | 1.0094 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211222 | 0.71 | 0.74 | -0.33 | 0.46 | 1.13 | -0.3 | 0.41 | 1.993 | 8.9 | 8.7 | 8.1 | 8.5 | 119.8 | 120 | 0.41 | 0.45 | -0.17 | 0.29 | 0.8 | -0.15 | 0.28 | 2.009 | 11.4 | 11.3 | 11.4 | 11.4 | 119.6 | 119.8 | LL |  | 2021-12-22 | 1.9774 | 1.0079 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20211223 | 0.74 | 0.66 | -0.42 | 0.41 | 1.03 | -0.34 | 0.34 | 1.993 | 8.6 | 9 | 8.3 | 8.2 | 119.8 | 120.1 | 0.4 | 0.4 | -0.19 | 0.23 | 0.73 | -0.18 | 0.22 | 2.006 | 11.3 | 11.3 | 11.4 | 11.5 | 119.6 | 119.9 | LL |  | 2021-12-23 | 1.9774 | 1.0079 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220104 | 0.68 | 0.71 | -0.46 | 0.34 | 1.16 | -0.36 | 0.34 | 1.997 | 8.6 | 9 | 8.3 | 8.6 | 119.9 | 120 | 0.37 | 0.58 | -0.12 | 0.24 | 0.94 | -0.11 | 0.24 | 1.996 | 11.7 | 12 | 11.8 | 12 | 119.7 | 119.9 | GB |  | 2022-01-04 | 1.9774 | 1.0099 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220104 | 0.68 | 0.71 | -0.46 | 0.34 | 1.16 | -0.36 | 0.34 | 1.997 | 8.6 | 9 | 8.3 | 8.6 | 119.9 | 120 | 0.37 | 0.58 | -0.12 | 0.24 | 0.94 | -0.11 | 0.24 | 1.996 | 11.7 | 12 | 11.8 | 12 | 119.7 | 119.9 | ML |  | 2022-01-04 | 1.9774 | 1.0099 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220105 | 0.54 | 0.72 | -0.24 | 0.29 | 1.17 | -0.22 | 0.3 | 1.979 | 9.2 | 8.2 | 9.1 | 7.5 | 120 | 120.1 | 0.37 | 0.58 | -0.12 | 0.24 | 0.94 | -0.11 | 0.24 | 1.996 | 11.7 | 12 | 11.8 | 12 | 119.7 | 119.7 | LL |  | 2022-01-05 | 1.9774 | 1.0008 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220106 | 0.72 | 0.86 | -0.36 | 0.57 | 1.22 | -0.34 | 0.47 | 1.983 | 8.6 | 9 | 8.4 | 8.4 | 119.9 | 120.1 | 0.34 | 0.52 | -0.13 | 0.33 | 0.74 | -0.15 | 0.27 | 1.998 | 11.5 | 11.4 | 11.5 | 11.3 | 119.6 | 119.8 | GB |  | 2022-01-06 | 1.9774 | 1.0028 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220107 | 0.37 | 0.79 | -0.19 | 0.32 | 1.12 | -0.21 | 0.33 | 1.983 | 8.5 | 9.2 | 8.7 | 8.3 | 119.8 | 120 | 0.33 | 0.53 | -0.13 | 0.18 | 0.89 | -0.08 | 0.22 | 2.001 | 11.7 | 11.6 | 11.6 | 11.3 | 119.5 | 119.8 | GB |  | 2022-01-07 | 1.9774 | 1.0028 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220110 | 0.73 | 0.61 | -0.37 | 0.35 | 1.09 | -0.34 | 0.35 | 1.984 | 8.5 | 9.1 | 8.3 | 8.4 | 119.9 | 120.1 | 0.36 | 0.4 | -0.13 | 0.15 | 0.76 | -0.11 | 0.17 | 2.005 | 11.8 | 11.5 | 11.5 | 11.3 | 119.6 | 119.8 | ML |  | 2022-01-10 | 1.9774 | 1.0033 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220112 | 0.74 | 0.63 | -0.4 | 0.28 | 1.15 | -0.37 | 0.27 | 1.996 | 8.3 | 9.3 | 8.6 | 8.2 | 119.8 | 120 | 0.34 | 0.45 | -0.11 | 0.22 | 0.76 | -0.15 | 0.2 | 2.016 | 11.7 | 11.5 | 11.4 | 11.2 | 119.5 | 119.8 | GB |  | 2022-01-12 | 1.9774 | 1.0094 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220113 | 0.68 | 0.73 | -0.39 | 0.45 | 1.16 | -0.36 | 0.41 | 1.988 | 8.5 | 9.1 | 8.4 | 8.4 | 119.8 | 120 | 0.39 | 0.48 | -0.13 | 0.25 | 0.76 | -0.14 | 0.22 | 2.006 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | GB |  | 2022-01-13 | 1.9774 | 1.0054 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220114 | 0.81 | 0.72 | -0.37 | 0.44 | 1.12 | -0.34 | 0.35 | 1.995 | 8.7 | 8.9 | 8.3 | 8.4 | 119.9 | 120 | 0.36 | 0.49 | -0.11 | 0.27 | 0.72 | -0.15 | 0.23 | 2.007 | 11.6 | 11.5 | 11.4 | 11.2 | 119.6 | 119.8 | GB |  | 2022-01-14 | 1.9774 | 1.0089 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220116 | 0.45 | 0.79 | -0.25 | 0.4 | 1.22 | -0.24 | 0.39 | 1.976 | 8.7 | 9.1 | 8.8 | 7.9 | 119.8 | 120.2 | 0.35 | 0.52 | -0.14 | 0.26 | 0.75 | -0.15 | 0.2 | 2.011 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.9 | ML |  | 2022-01-16 | 1.9774 | 0.99929 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220117 | 0.63 | 0.85 | -0.38 | 0.49 | 1.29 | -0.38 | 0.4 | 1.989 | 8.4 | 9.2 | 8.4 | 8.5 | 119.8 | 120 | 0.35 | 0.52 | -0.14 | 0.26 | 0.75 | -0.15 | 0.2 | 2.011 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2022-01-17 | 1.9774 | 1.0059 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220119 | 0.55 | 0.67 | -0.36 | 0.34 | 1.09 | -0.36 | 0.31 | 1.998 | 8.5 | 9.1 | 8.4 | 8.4 | 119.9 | 120 | 0.3 | 0.43 | -0.13 | 0.19 | 0.74 | -0.16 | 0.15 | 2.022 | 11.6 | 11.4 | 11.5 | 11.3 | 119.6 | 119.8 | LL |  | 2022-01-19 | 1.9774 | 1.0104 | 1.9991 | 1.0115 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220120 | 0.68 | 0.79 | -0.35 | 0.48 | 1.26 | -0.35 | 0.41 | 1.986 | 8.3 | 9.2 | 8.3 | 8.4 | 119.8 | 120 | 0.37 | 0.52 | -0.17 | 0.31 | 0.79 | -0.15 | 0.3 | 2.004 | 11.7 | 11.4 | 11.4 | 11.5 | 119.6 | 119.8 | ML |  | 2022-01-20 | 1.9774 | 1.0043 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220121 | 0.87 | 0.7 | -0.56 | 0.37 | 1.22 | -0.4 | 0.36 | 1.994 | 8.4 | 9.2 | 8.3 | 8.5 | 119.8 | 120 | 0.37 | 0.41 | -0.13 | 0.22 | 0.75 | -0.15 | 0.2 | 2.009 | 11.8 | 11.5 | 11.6 | 11.3 | 119.5 | 119.8 | ML |  | 2022-01-21 | 1.9774 | 1.0084 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220124 | 0.76 | 0.69 | -0.4 | 0.43 | 1.14 | -0.34 | 0.39 | 1.993 | 8.8 | 8.9 | 8.3 | 8.4 | 119.8 | 120 | 0.36 | 0.41 | -0.13 | 0.24 | 0.73 | -0.15 | 0.23 | 2.01 | 11.5 | 11.5 | 11.4 | 11.3 | 119.5 | 119.9 | ML |  | 2022-01-24 | 1.9774 | 1.0079 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220126 | 0.84 | 0.65 | -0.46 | 0.4 | 1.19 | -0.33 | 0.38 | 1.986 | 9.1 | 8.6 | 8 | 8.7 | 119.8 | 120 | 0.41 | 0.42 | -0.16 | 0.22 | 0.77 | -0.15 | 0.2 | 2.005 | 11.5 | 11.5 | 11.7 | 11.4 | 119.5 | 119.8 | ML |  | 2022-01-26 | 1.9774 | 1.0043 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220127 | 0.69 | 0.72 | -0.35 | 0.38 | 1.13 | -0.33 | 0.34 | 1.996 | 8.7 | 8.9 | 8.3 | 8.4 | 119.9 | 120 | 0.4 | 0.47 | -0.13 | 0.21 | 0.71 | -0.15 | 0.15 | 2.017 | 11.5 | 11.4 | 11.4 | 11.3 | 119.5 | 119.8 | GB |  | 2022-01-27 | 1.9774 | 1.0094 | 1.9991 | 1.009 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220202 | 0.6 | 0.74 | -0.35 | 0.34 | 1.09 | -0.33 | 0.32 | 1.995 | 8.9 | 8.8 | 8.6 | 8.2 | 119.9 | 120 | 0.32 | 0.46 | -0.12 | 0.21 | 0.71 | -0.16 | 0.18 | 2.014 | 11.6 | 11.7 | 11.6 | 11.7 | 119.6 | 119.8 | ML |  | 2022-02-02 | 1.9774 | 1.0089 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220203 | 0.63 | 0.76 | -0.43 | 0.48 | 1.13 | -0.35 | 0.4 | 1.99 | 8.4 | 9.2 | 7.4 | 9 | 119.8 | 119.9 | 0.39 | 0.56 | -0.14 | 0.33 | 0.75 | -0.13 | 0.26 | 2.009 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2022-02-03 | 1.9774 | 1.0064 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220204 | 0.68 | 0.71 | -0.46 | 0.34 | 1.16 | -0.36 | 0.34 | 1.997 | 8.6 | 9 | 8.3 | 8.6 | 119.9 | 120 | 0.33 | 0.41 | -0.14 | 0.18 | 0.67 | -0.16 | 0.14 | 2.015 | 11.6 | 11.5 | 11.6 | 11.3 | 119.5 | 119.8 | ML |  | 2022-02-04 | 1.9774 | 1.0099 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220207 | 0.72 | 0.77 | -0.39 | 0.48 | 1.15 | -0.33 | 0.4 | 1.988 | 8.9 | 8.8 | 8.3 | 8.4 | 119.9 | 120 | 0.33 | 0.49 | -0.1 | 0.25 | 0.78 | -0.15 | 0.2 | 2.011 | 11.5 | 11.6 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2022-02-07 | 1.9774 | 1.0054 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220209 | 0.69 | 0.75 | -0.42 | 0.48 | 1.16 | -0.38 | 0.39 | 1.992 | 8.2 | 9.4 | 8.2 | 8.5 | 119.8 | 120.1 | 0.42 | 0.52 | -0.15 | 0.27 | 0.75 | -0.18 | 0.19 | 2.01 | 11.7 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | GB |  | 2022-02-09 | 1.9774 | 1.0074 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220214 | 0.73 | 0.67 | -0.41 | 0.37 | 1.21 | -0.34 | 0.36 | 1.987 | 9.1 | 8.5 | 8 | 8.8 | 119.9 | 120 | 0.41 | 0.39 | -0.17 | 0.2 | 0.75 | -0.17 | 0.17 | 2.002 | 11.6 | 11.7 | 11.8 | 11.4 | 119.5 | 119.8 | ML |  | 2022-02-14 | 1.9774 | 1.0049 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220217 | 0.8 | 0.67 | -0.35 | 0.38 | 1.17 | -0.33 | 0.36 | 1.988 | 9.2 | 8.4 | 8.2 | 8.4 | 119.9 | 120.1 | 0.36 | 0.5 | -0.12 | 0.3 | 0.78 | -0.14 | 0.23 | 2.007 | 11.5 | 11.6 | 11.5 | 11.3 | 119.6 | 119.8 | GB |  | 2022-02-17 | 1.9774 | 1.0054 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220218 | 0.69 | 0.72 | -0.4 | 0.36 | 1.13 | -0.32 | 0.36 | 1.986 | 8.8 | 8.8 | 7.9 | 8.8 | 119.9 | 120 | 0.37 | 0.46 | -0.13 | 0.24 | 0.78 | -0.19 | 0.18 | 2.001 | 11.6 | 11.6 | 11.5 | 11.5 | 119.6 | 119.8 | ML |  | 2022-02-18 | 1.9774 | 1.0043 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220221 | 0.83 | 0.7 | -0.45 | 0.38 | 1.3 | -0.36 | 0.37 | 1.986 | 9.3 | 8.2 | 8.4 | 8.6 | 119.9 | 120 | 0.35 | 0.44 | -0.14 | 0.22 | 0.78 | -0.19 | 0.16 | 2.004 | 11.7 | 11.9 | 11.8 | 11.5 | 119.5 | 119.8 | ML |  | 2022-02-21 | 1.9774 | 1.0043 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220223 | 0.7 | 0.67 | -0.37 | 0.34 | 1.16 | -0.37 | 0.32 | 1.991 | 8.9 | 8.8 | 8.3 | 8.5 | 119.9 | 120 | 0.36 | 0.39 | -0.12 | 0.19 | 0.75 | -0.14 | 0.16 | 2.008 | 11.6 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2022-02-23 | 1.9774 | 1.0069 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220224 | 0.81 | 0.75 | -0.43 | 0.34 | 1.18 | -0.37 | 0.34 | 1.991 | 8.8 | 8.9 | 8.5 | 8.3 | 119.9 | 120 | 0.39 | 0.41 | -0.1 | 0.2 | 0.71 | -0.16 | 0.17 | 2.008 | 11.5 | 11.4 | 11.4 | 11.4 | 119.6 | 119.8 | GB |  | 2022-02-24 | 1.9774 | 1.0069 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220228 | 0.67 | 0.75 | -0.32 | 0.31 | 1.1 | -0.31 | 0.29 | 1.989 | 9.1 | 8.5 | 8.4 | 8.4 | 119.9 | 120.1 | 0.36 | 0.41 | -0.15 | 0.16 | 0.69 | -0.17 | 0.13 | 2.002 | 11.6 | 11.8 | 11.7 | 11.4 | 119.6 | 119.8 | ML |  | 2022-02-28 | 1.9774 | 1.0059 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220302 | 0.84 | 0.63 | -0.41 | 0.36 | 1.18 | -0.32 | 0.36 | 1.993 | 9.1 | 8.6 | 8.2 | 8.6 | 119.9 | 120.1 | 0.42 | 0.41 | -0.16 | 0.19 | 0.76 | -0.19 | 0.18 | 2.01 | 11.6 | 11.8 | 11.8 | 11.4 | 119.6 | 119.8 | ML |  | 2022-03-02 | 1.9774 | 1.0079 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220303 | 0.78 | 0.74 | -0.36 | 0.39 | 1.14 | -0.32 | 0.36 | 1.981 | 8.8 | 8.9 | 8.5 | 8.3 | 119.9 | 120.1 | 0.38 | 0.5 | -0.14 | 0.28 | 0.79 | -0.16 | 0.23 | 1.997 | 11.5 | 11.5 | 11.6 | 11.3 | 119.6 | 119.9 | GB |  | 2022-03-03 | 1.9774 | 1.0018 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220304 | 0.86 | 0.64 | -0.4 | 0.36 | 1.24 | -0.36 | 0.33 | 1.981 | 9.3 | 8.3 | 8.4 | 8.3 | 119.9 | 120.1 | 0.44 | 0.43 | -0.2 | 0.26 | 0.76 | -0.21 | 0.21 | 1.999 | 11.6 | 11.8 | 11.6 | 11.3 | 119.6 | 119.8 | GB |  | 2022-03-04 | 1.9774 | 1.0018 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220307 | 0.71 | 0.69 | -0.43 | 0.41 | 1.13 | -0.34 | 0.36 | 1.989 | 9 | 8.7 | 8.2 | 8.6 | 119.9 | 120 | 0.46 | 0.42 | -0.2 | 0.22 | 0.75 | -0.17 | 0.21 | 2.002 | 11.6 | 11.7 | 11.6 | 11.3 | 119.6 | 119.9 | CVD |  | 2022-03-07 | 1.9774 | 1.0059 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220310 | 0.78 | 0.79 | -0.34 | 0.48 | 1.22 | -0.32 | 0.41 | 1.986 | 8.6 | 9 | 8.3 | 8.5 | 119.9 | 120 | 0.31 | 0.34 | -0.13 | 0.16 | 0.62 | -0.15 | 0.15 | 2 | 11.5 | 11.4 | 11.6 | 11.3 | 119.6 | 119.8 | CVD |  | 2022-03-10 | 1.9774 | 1.0043 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220315 | 0.75 | 0.79 | -0.41 | 0.45 | 1.17 | -0.35 | 0.36 | 1.994 | 8.4 | 9.2 | 8.4 | 8.4 | 119.9 | 120 | 0.38 | 0.45 | -0.17 | 0.24 | 0.72 | -0.19 | 0.21 | 2.008 | 11.7 | 11.5 | 11.7 | 11.4 | 119.6 | 119.8 | GB |  | 2022-03-15 | 1.9774 | 1.0084 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220318 | 0.8 | 0.83 | -0.43 | 0.43 | 1.25 | -0.38 | 0.37 | 1.991 | 8.4 | 9.1 | 8.5 | 8.4 | 119.9 | 120 | 0.39 | 0.47 | -0.14 | 0.27 | 0.74 | -0.17 | 0.21 | 2.011 | 11.7 | 11.5 | 11.6 | 11.4 | 119.6 | 119.8 | GB |  | 2022-03-18 | 1.9774 | 1.0069 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220321 | 0.73 | 0.84 | -0.38 | 0.45 | 1.23 | -0.34 | 0.4 | 1.992 | 8.9 | 8.7 | 8.4 | 8.5 | 119.9 | 120 | 0.39 | 0.5 | -0.13 | 0.27 | 0.76 | -0.15 | 0.18 | 2.009 | 11.6 | 11.5 | 11.7 | 11.4 | 119.6 | 119.9 | ML |  | 2022-03-21 | 1.9774 | 1.0074 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220322 | 0.75 | 0.8 | -0.4 | 0.49 | 1.31 | -0.36 | 0.44 | 1.989 | 8.9 | 8.8 | 8 | 8.6 | 119.9 | 120.1 | 0.33 | 0.3 | -0.14 | 0.11 | 0.6 | -0.16 | 0.11 | 1.999 | 11.5 | 11.4 | 11.7 | 11.5 | 119.6 | 119.8 | CVD |  | 2022-03-22 | 1.9774 | 1.0059 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220324 | 0.68 | 0.77 | -0.45 | 0.39 | 1.11 | -0.34 | 0.34 | 1.987 | 8.9 | 8.7 | 8.1 | 8.6 | 119.9 | 120.1 | 0.37 | 0.42 | -0.15 | 0.18 | 0.78 | -0.18 | 0.14 | 2.005 | 11.6 | 11.7 | 11.9 | 11.5 | 119.6 | 119.8 | GB |  | 2022-03-24 | 1.9774 | 1.0049 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220328 | 0.7 | 0.64 | -0.44 | 0.35 | 1.09 | -0.36 | 0.34 | 1.977 | 9 | 8.7 | 7.8 | 8.7 | 119.9 | 120.1 | 0.44 | 0.42 | -0.19 | 0.23 | 0.74 | -0.19 | 0.19 | 1.995 | 11.6 | 11.6 | 11.9 | 11.5 | 119.6 | 119.8 | GB |  | 2022-03-28 | 1.9774 | 0.9998 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220330 | 0.79 | 0.7 | -0.41 | 0.42 | 1.15 | -0.36 | 0.36 | 1.998 | 9 | 8.7 | 8.3 | 8.4 | 119.9 | 120.1 | 0.41 | 0.45 | -0.15 | 0.25 | 0.78 | -0.19 | 0.2 | 2.017 | 11.5 | 11.4 | 11.4 | 11.2 | 119.6 | 119.8 | ML |  | 2022-03-30 | 1.9774 | 1.0104 | 1.9991 | 1.009 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220331 | 0.71 | 0.75 | -0.38 | 0.41 | 1.18 | -0.37 | 0.35 | 1.991 | 9 | 8.7 | 8.2 | 8.5 | 119.9 | 120.1 | 0.4 | 0.42 | -0.14 | 0.22 | 0.77 | -0.16 | 0.18 | 2.009 | 11.5 | 11.5 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2022-03-31 | 1.9774 | 1.0069 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220401 | 0.78 | 0.73 | -0.42 | 0.44 | 1.23 | -0.32 | 0.42 | 1.984 | 9 | 8.6 | 8 | 8.7 | 119.9 | 120 | 0.37 | 0.51 | -0.13 | 0.27 | 0.76 | -0.12 | 0.23 | 2.008 | 11.5 | 11.5 | 11.7 | 11.4 | 119.5 | 119.8 | ML |  | 2022-04-01 | 1.9774 | 1.0033 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220404 | 0.7 | 0.8 | -0.38 | 0.45 | 1.24 | -0.34 | 0.4 | 1.986 | 8.7 | 9 | 8.5 | 8.4 | 119.9 | 120 | 0.44 | 0.46 | -0.16 | 0.25 | 0.79 | -0.16 | 0.21 | 2.003 | 11.6 | 11.6 | 11.5 | 11.3 | 119.5 | 119.8 | ML |  | 2022-04-04 | 1.9774 | 1.0043 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220406 | 0.73 | 0.64 | -0.35 | 0.37 | 1.18 | -0.34 | 0.34 | 1.981 | 9.2 | 8.5 | 8.2 | 8.4 | 119.9 | 120.1 | 0.45 | 0.44 | -0.19 | 0.23 | 0.78 | -0.2 | 0.19 | 1.999 | 11.6 | 11.7 | 11.6 | 11.3 | 119.5 | 119.8 | ML |  | 2022-04-06 | 1.9774 | 1.0018 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220407 | 0.78 | 0.71 | -0.46 | 0.38 | 1.15 | -0.36 | 0.36 | 1.989 | 8.8 | 8.9 | 8.1 | 8.6 | 119.8 | 120 | 0.35 | 0.5 | -0.15 | 0.28 | 0.78 | -0.16 | 0.23 | 2.007 | 11.5 | 11.6 | 11.5 | 11.5 | 119.6 | 119.9 | GB |  | 2022-04-07 | 1.9774 | 1.0059 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220408 | 0.86 | 0.76 | -0.44 | 0.47 | 1.39 | -0.38 | 0.42 | 1.976 | 8.5 | 9.1 | 8.5 | 8.3 | 119.9 | 120.1 | 0.47 | 0.39 | -0.22 | 0.19 | 0.78 | -0.2 | 0.2 | 1.995 | 11.6 | 11.5 | 11.5 | 11.3 | 119.6 | 119.9 | GB |  | 2022-04-08 | 1.9774 | 0.99929 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220411 | 0.78 | 0.67 | -0.39 | 0.31 | 1.35 | -0.4 | 0.3 | 1.994 | 8.6 | 8.9 | 8.7 | 8 | 120 | 120.1 | 0.39 | 0.44 | -0.14 | 0.19 | 0.73 | -0.16 | 0.16 | 2.019 | 11.5 | 11.4 | 11.4 | 11.4 | 119.7 | 119.8 | LL |  | 2022-04-11 | 1.9774 | 1.0084 | 1.9991 | 1.01 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220414 | 0.67 | 0.61 | -0.42 | 0.23 | 1.08 | -0.36 | 0.2 | 1.992 | 8.5 | 9.1 | 8.6 | 8.2 | 119.9 | 120 | 0.33 | 0.45 | -0.11 | 0.22 | 0.72 | -0.14 | 0.18 | 2.009 | 11.6 | 11.5 | 11.5 | 11.4 | 119.6 | 119.9 | CVD |  | 2022-04-14 | 1.9774 | 1.0074 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220416 | [14 empty] | 0.26 | 0.35 | -0.11 | 0.15 | 0.62 | -0.12 | 0.15 | 1.993 | 11.5 | 11.6 | 11.8 | 11.4 | 119.6 |  | CVD |  | 2022-04-16 | 1.9774 | 0 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220419 | 0.67 | 0.61 | -0.35 | 0.41 | 1.16 | -0.32 | 0.39 | 1.978 | 8.9 | 8.6 | 7.8 | 8.7 | 119.9 | 120.1 | [13 empty] | 119.8 | AD |  | 2022-04-19 | 1.9774 | 1.0003 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220421 | 0.86 | 0.79 | -0.4 | 0.45 | 1.27 | -0.35 | 0.4 | 1.996 | 8.7 | 9 | 8.3 | 8.3 | 119.9 | 120 | 0.41 | 0.37 | -0.19 | 0.19 | 0.7 | -0.2 | 0.18 | 2.007 | 11.5 | 11.5 | 11.5 | 11.6 | 119.6 | 119.8 | GB |  | 2022-04-21 | 1.9774 | 1.0094 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220422 | 0.81 | 0.76 | -0.37 | 0.44 | 1.15 | -0.33 | 0.36 | 1.988 | 8.9 | 8.7 | 8.5 | 8.3 | 119.9 | 120 | 0.37 | 0.48 | -0.14 | 0.26 | 0.74 | -0.15 | 0.23 | 2.005 | 11.6 | 11.6 | 11.5 | 11.3 | 119.5 | 119.8 | GB |  | 2022-04-22 | 1.9774 | 1.0054 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220425 | 0.68 | 0.73 | -0.32 | 0.39 | 1.16 | -0.31 | 0.34 | 1.988 | 9.2 | 8.4 | 8.5 | 8.3 | 119.9 | 120 | 0.4 | 0.48 | -0.16 | 0.26 | 0.82 | -0.19 | 0.21 | 2.009 | 11.6 | 11.7 | 11.5 | 11.3 | 119.5 | 119.9 | CVD |  | 2022-04-25 | 1.9774 | 1.0054 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220427 | 0.79 | 0.7 | -0.43 | 0.4 | 1.25 | -0.35 | 0.38 | 1.986 | 9.2 | 8.4 | 8.2 | 8.6 | 119.9 | 120 | 0.39 | 0.49 | -0.13 | 0.26 | 0.79 | -0.17 | 0.19 | 2.002 | 11.6 | 11.8 | 11.5 | 11.6 | 119.6 | 119.8 | GB |  | 2022-04-27 | 1.9774 | 1.0043 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220504 | 0.7 | 0.7 | -0.34 | 0.3 | 1.2 | -0.33 | 0.25 | 2.004 | 9 | 8.7 | 8.6 | 8.2 | 119.9 | 120 | 0.41 | 0.55 | -0.13 | 0.3 | 0.82 | -0.16 | 0.23 | 2.022 | 11.5 | 11.5 | 11.4 | 11.3 | 119.6 | 119.9 | LL |  | 2022-05-04 | 1.9774 | 1.0135 | 1.9991 | 1.0115 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220505 | 0.8 | 0.66 | -0.37 | 0.33 | 1.24 | -0.34 | 0.34 | 1.986 | 9 | 8.5 | 8.4 | 8.4 | 119.9 | 120.1 | 0.41 | 0.45 | -0.15 | 0.23 | 0.73 | -0.18 | 0.18 | 2.006 | 11.6 | 11.6 | 11.6 | 11.3 | 119.6 | 120 | ML |  | 2022-05-05 | 1.9774 | 1.0043 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220509 | 0.76 | 0.68 | -0.37 | 0.41 | 1.14 | -0.32 | 0.36 | 1.985 | 8.9 | 8.7 | 8.1 | 8.3 | 119.9 | 120.2 | 0.41 | 0.46 | -0.15 | 0.23 | 0.75 | -0.17 | 0.19 | 2.005 | 11.5 | 11.5 | 11.6 | 11.3 | 119.6 | 119.9 | LL |  | 2022-05-09 | 1.9774 | 1.0038 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220511 | 0.78 | 0.73 | -0.44 | 0.44 | 1.16 | -0.35 | 0.38 | 1.983 | 8.6 | 9.1 | 8.3 | 8.3 | 119.9 | 120.1 | 0.44 | 0.39 | -0.22 | 0.22 | 0.76 | -0.17 | 0.22 | 2.002 | 11.6 | 11.5 | 11.5 | 11.2 | 119.6 | 120 | ML |  | 2022-05-11 | 1.9774 | 1.0028 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220512 | 0.62 | 0.72 | -0.38 | 0.34 | 1.2 | -0.39 | 0.29 | 1.988 | 8.6 | 9 | 8.2 | 8.4 | 119.9 | 120.1 | 0.4 | 0.43 | -0.17 | 0.23 | 0.77 | -0.2 | 0.2 | 2.002 | 11.5 | 11.4 | 11.4 | 11.6 | 119.6 | 120 | AD |  | 2022-05-12 | 1.9774 | 1.0054 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220513 | 0.79 | 0.79 | -0.35 | 0.48 | 1.11 | -0.31 | 0.39 | 1.981 | 8.8 | 8.9 | 8 | 8.4 | 119.9 | 120.1 | 0.35 | 0.39 | -0.15 | 0.2 | 0.68 | -0.14 | 0.21 | 2.003 | 11.4 | 11.4 | 11.4 | 11.4 | 119.7 | 119.8 | LL |  | 2022-05-13 | 1.9774 | 1.0018 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220516 | 0.71 | 0.64 | -0.36 | 0.4 | 1.18 | -0.34 | 0.37 | 1.988 | 8.8 | 8.9 | 8.1 | 8.4 | 119.9 | 120.1 | 0.37 | 0.4 | -0.16 | 0.16 | 0.77 | -0.17 | 0.17 | 2.008 | 11.5 | 11.4 | 11.6 | 11.4 | 119.5 | 119.9 | CVD |  | 2022-05-16 | 1.9774 | 1.0054 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220518 | 0.77 | 0.7 | -0.41 | 0.43 | 1.32 | -0.37 | 0.39 | 1.981 | 8.6 | 9 | 8 | 8.5 | 119.9 | 120.1 | 0.31 | 0.43 | -0.13 | -0.06 | 0.6 | -0.14 | -0.04 | 1.996 | 11.4 | 11.4 | 11.5 | 11.4 | 119.6 | 119.9 | ML |  | 2022-05-18 | 1.9774 | 1.0018 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220519 | 0.63 | 0.71 | -0.34 | 0.39 | 1.07 | -0.33 | 0.3 | 1.993 | 8.6 | 9 | 8.4 | 8.2 | 119.9 | 120.1 | 0.36 | 0.46 | -0.14 | 0.21 | 0.71 | -0.16 | 0.15 | 2.015 | 11.4 | 11.4 | 11.4 | 11.2 | 119.6 | 119.9 | GB |  | 2022-05-19 | 1.9774 | 1.0079 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220520 | 0.75 | 0.77 | -0.31 | 0.46 | 1.17 | -0.34 | 0.38 | 1.986 | 8.7 | 8.9 | 8.1 | 8.4 | 119.9 | 120.1 | 0.35 | 0.38 | -0.12 | 0.21 | 0.76 | -0.14 | 0.19 | 2.004 | 11.4 | 11.4 | 11.6 | 11.3 | 119.6 | 120.1 | CVD |  | 2022-05-20 | 1.9774 | 1.0043 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220523 | 0.87 | 0.68 | -0.39 | 0.37 | 1.27 | -0.36 | 0.39 | 1.981 | 9 | 8.5 | 7.9 | 8.4 | 119.9 | 120.2 | 0.37 | 0.39 | -0.17 | -0.02 | 0.59 | -0.17 | -0.03 | 1.99 | 11.5 | 11.6 | 11.5 | 11.3 | 119.7 | 119.9 | ML |  | 2022-05-23 | 1.9774 | 1.0018 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220525 | 0.86 | 0.75 | -0.37 | 0.41 | 1.32 | -0.34 | 0.4 | 1.994 | 9 | 8.6 | 8.2 | 8.4 | 119.9 | 120.1 | 0.34 | 0.46 | -0.13 | 0.2 | 0.72 | -0.16 | 0.19 | 2.014 | 11.5 | 11.4 | 11.5 | 11.2 | 119.6 | 119.9 | ML |  | 2022-05-25 | 1.9774 | 1.0084 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220527 | 0.83 | 0.77 | -0.49 | 0.38 | 1.16 | -0.39 | 0.32 | 1.995 | 8.9 | 8.7 | 8 | 8.6 | 119.9 | 120.1 | 0.43 | 0.5 | -0.14 | 0.19 | 0.79 | -0.19 | 0.14 | 2.018 | 11.5 | 11.5 | 11.5 | 11.3 | 119.5 | 119.9 | ML |  | 2022-05-27 | 1.9774 | 1.0089 | 1.9991 | 1.0095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220530 | 0.76 | 0.77 | -0.44 | 0.45 | 1.17 | -0.4 | 0.38 | 2.003 | 8.6 | 9 | 7.9 | 8.5 | 119.9 | 120.1 | 0.41 | 0.45 | -0.16 | 0.25 | 0.76 | -0.2 | 0.18 | 2.028 | 11.3 | 11.4 | 11.4 | 11.2 | 119.5 | 120 | GB |  | 2022-05-30 | 1.9774 | 1.0129 | 1.9991 | 1.0145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220603 | 0.78 | 0.81 | -0.42 | 0.52 | 1.29 | -0.37 | 0.46 | 1.989 | 8.5 | 9 | 7.7 | 8.5 | 119.9 | 120.2 | 0.39 | 0.28 | -0.19 | 0.1 | 0.72 | -0.19 | 0.13 | 2.009 | 11.6 | 11.5 | 11.6 | 11.3 | 119.6 | 119.9 | GB |  | 2022-06-03 | 1.9774 | 1.0059 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220607 | 0.7 | 0.95 | -0.41 | 0.57 | 1.28 | -0.37 | 0.45 | 1.993 | 8.5 | 9.1 | 8.1 | 8.5 | 119.9 | 120.1 | 0.49 | 0.49 | -0.24 | 0.31 | 0.87 | -0.22 | 0.28 | 2.008 | 11.6 | 11.5 | 11.6 | 11.3 | 119.6 | 119.8 | YM |  | 2022-06-07 | 1.9774 | 1.0079 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220609 | 0.88 | 0.79 | -0.56 | 0.4 | 1.4 | -0.42 | 0.35 | 1.995 | 9 | 8.7 | 8.1 | 8.8 | 119.9 | 120 | 0.43 | 0.45 | -0.13 | 0.21 | 0.74 | -0.18 | 0.14 | 2.015 | 11.5 | 11.5 | 11.6 | 11.3 | 119.5 | 119.9 | GB |  | 2022-06-09 | 1.9774 | 1.0089 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220610 | 0.7 | 0.71 | -0.35 | 0.41 | 1.15 | -0.36 | 0.36 | 1.992 | 9 | 8.7 | 8.2 | 8.5 | 119.9 | 120.1 | 0.39 | 0.42 | -0.16 | 0.21 | 0.77 | -0.19 | 0.19 | 2.015 | 11.5 | 11.5 | 11.5 | 11.3 | 119.5 | 119.9 | GB |  | 2022-06-10 | 1.9774 | 1.0074 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220613 | 0.76 | 0.71 | -0.39 | 0.43 | 1.14 | -0.36 | 0.37 | 2.007 | 8.8 | 8.7 | 8.1 | 8.3 | 119.9 | 120.2 | 0.38 | 0.5 | -0.13 | 0.28 | 0.77 | -0.14 | 0.24 | 2.033 | 11.3 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | GB |  | 2022-06-13 | 1.9774 | 1.015 | 1.9991 | 1.017 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220614 | 0.67 | 0.71 | -0.29 | 0.42 | 1.07 | -0.28 | 0.37 | 1.991 | 8.8 | 8.7 | 8.1 | 8.5 | 119.9 | 120.1 | 0.3 | 0.34 | -0.11 | 0.17 | 0.59 | -0.13 | 0.18 | 2.008 | 11.4 | 11.4 | 11.6 | 11.3 | 119.6 | 120 | GB |  | 2022-06-14 | 1.9774 | 1.0069 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220616 | 0.77 | 0.81 | -0.38 | 0.48 | 1.23 | -0.32 | 0.42 | 1.996 | 8.7 | 9 | 8.3 | 8.2 | 119.9 | 120.2 | 0.31 | 0.54 | -0.08 | 0.31 | 0.73 | -0.11 | 0.27 | 2.016 | 11.5 | 11.4 | 11.4 | 11.3 | 119.6 | 119.9 | ML |  | 2022-06-16 | 1.9774 | 1.0094 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220620 | 0.75 | 0.75 | -0.37 | 0.39 | 1.2 | -0.36 | 0.34 | 1.986 | 8.9 | 8.7 | 8.3 | 8.4 | 119.9 | 120.1 | 0.35 | 0.36 | -0.13 | 0.11 | 0.67 | -0.16 | 0.12 | 2.005 | 11.5 | 11.6 | 11.9 | 11.5 | 119.7 | 120.1 | ML |  | 2022-06-20 | 1.9774 | 1.0043 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220623 | 0.67 | 0.87 | -0.32 | 0.53 | 1.29 | -0.32 | 0.45 | 1.985 | 8.8 | 8.8 | 8 | 8.4 | 119.9 | 120.2 | 0.41 | 0.47 | -0.15 | 0.25 | 0.8 | -0.18 | 0.21 | 2.006 | 11.5 | 11.4 | 11.8 | 11.3 | 119.7 | 119.8 | AD |  | 2022-06-23 | 1.9774 | 1.0038 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220630 | 0.75 | 0.84 | -0.39 | 0.46 | 1.38 | -0.41 | 0.38 | 1.999 | 8.8 | 8.9 | 8.6 | 8.3 | 119.9 | 119.9 | 0.42 | 0.49 | -0.17 | 0.27 | 0.89 | -0.25 | 0.22 | 2.02 | 11.4 | 11.3 | 11.4 | 11.2 | 119.6 | 119.4 | ML |  | 2022-06-30 | 1.9774 | 1.0109 | 1.9991 | 1.0105 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220701 | 0.76 | 0.73 | -0.38 | 0.39 | 1.2 | -0.36 | 0.35 | 2.02 | 9 | 8.7 | 8.8 | 8.6 | 119.9 | 119.6 | 0.4 | 0.49 | -0.14 | 0.25 | 0.8 | -0.18 | 0.23 | 2.039 | 11.4 | 11.3 | 11.5 | 11.3 | 119.6 | 120.1 | ML |  | 2022-07-01 | 1.9774 | 1.0215 | 1.9991 | 1.02 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220706 | 0.69 | 0.82 | -0.34 | 0.46 | 1.28 | -0.36 | 0.39 | 2.005 | 8.6 | 8.9 | 8.2 | 8.1 | 119.9 | 120.3 | 0.31 | 0.54 | -0.08 | 0.31 | 0.74 | -0.15 | 0.25 | 2.024 | 11.3 | 11.3 | 11.5 | 11.2 | 119.7 | 120.1 | ML |  | 2022-07-06 | 1.9774 | 1.014 | 1.9991 | 1.0125 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220707 | 0.86 | 0.6 | -0.4 | 0.41 | 1.25 | -0.38 | 0.38 | 2.001 | 9 | 8.6 | 7.6 | 8.4 | 119.9 | 120.3 | 0.36 | 0.36 | -0.15 | 0.16 | 0.77 | -0.2 | 0.18 | 2.02 | 11.4 | 11.2 | 11.6 | 11.2 | 119.6 | 119.9 | ML |  | 2022-07-07 | 1.9774 | 1.0119 | 1.9991 | 1.0105 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220708 | 0.8 | 0.78 | -0.37 | 0.48 | 1.27 | -0.33 | 0.46 | 2.009 | 9.1 | 8.5 | 6.9 | 9.6 | 119.8 | 119.9 | 0.33 | 0.4 | -0.12 | 0.2 | 0.77 | -0.19 | 0.16 | 2.028 | 11.4 | 11.3 | 11.9 | 11.6 | 119.6 | 120.1 | CVD |  | 2022-07-08 | 1.9774 | 1.016 | 1.9991 | 1.0145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220712 | 0.82 | 0.91 | -0.46 | 0.63 | 1.47 | -0.37 | 0.56 | 1.976 | 8.6 | 9.1 | 8.4 | 8.2 | 119.9 | 120.3 | 0.32 | 0.34 | -0.13 | 0.15 | 0.67 | -0.15 | 0.17 | 1.989 | 12.1 | 11.8 | 12.1 | 11.9 | 119.7 | 120.2 | ML |  | 2022-07-12 | 1.9774 | 0.99929 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220714 | 0.69 | 0.9 | -0.4 | 0.56 | 1.41 | -0.38 | 0.5 | 1.972 | 8.4 | 9.1 | 8.1 | 8.3 | 119.9 | 120.2 | 0.57 | 0.43 | -0.29 | 0.24 | 0.86 | -0.29 | 0.23 | 1.976 | 11.4 | 11.4 | 11.7 | 11.1 | 119.7 | 120.1 | ML |  | 2022-07-14 | 1.9774 | 0.99727 | 1.9991 | 0.98844 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220715 | 0.75 | 0.84 | -0.4 | 0.52 | 1.39 | -0.4 | 0.48 | 1.976 | 8.2 | 9.5 | 8.4 | 8.3 | 119.8 | 120.2 | 0.37 | 0.46 | -0.14 | 0.23 | 0.77 | -0.18 | 0.24 | 1.992 | 12.2 | 11.7 | 12.2 | 11.8 | 119.6 | 120.1 | ef |  | 2022-07-15 | 1.9774 | 0.99929 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220718 | 0.75 | 0.87 | -0.39 | 0.56 | 1.42 | -0.37 | 0.51 | 1.983 | 9.2 | 8.6 | 7.2 | 9.1 | 119.9 | 120.1 | 0.36 | 0.3 | -0.14 | 0.1 | 0.65 | -0.17 | 0.12 | 1.989 | 11.7 | 12.1 | 12.4 | 11.7 | 119.6 | 120.1 | YM |  | 2022-07-18 | 1.9774 | 1.0028 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220720 | 0.72 | 0.97 | -0.37 | 0.62 | 1.45 | -0.34 | 0.54 | 1.975 | 8.9 | 8.9 | 8.2 | 8.5 | 119.9 | 120.2 | 0.41 | 0.44 | -0.21 | 0.25 | 0.8 | -0.2 | 0.26 | 1.986 | 11.8 | 12.1 | 12.3 | 11.6 | 119.6 | 120.1 | ML |  | 2022-07-20 | 1.9774 | 0.99879 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220721 | 0.81 | 0.86 | -0.38 | 0.61 | 1.52 | -0.38 | 0.55 | 1.971 | 8.9 | 8.9 | 8.1 | 8.6 | 119.9 | 120.2 | 0.49 | 0.38 | -0.24 | 0.17 | 0.84 | -0.24 | 0.2 | 1.979 | 11.8 | 12.2 | 12.1 | 12.1 | 119.7 | 120.2 | ML |  | 2022-07-21 | 1.9774 | 0.99676 | 1.9991 | 0.98995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220722 | 0.79 | 0.92 | -0.42 | 0.62 | 1.53 | -0.38 | 0.56 | 1.979 | 9 | 8.9 | 8.3 | 8.6 | 119.9 | 120.2 | 0.42 | 0.28 | -0.23 | 0.1 | 0.73 | -0.23 | 0.13 | 1.972 | 11.9 | 12.5 | 11.9 | 12.2 | 119.7 | 120.1 | ML |  | 2022-07-22 | 1.9774 | 1.0008 | 1.9991 | 0.98644 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220725 | 0.86 | 0.9 | -0.46 | 0.58 | 1.53 | -0.39 | 0.53 | 1.981 | 8.3 | 9.3 | 8 | 8.5 | 119.9 | 120.2 | 0.21 | 0.41 | -0.06 | 0.19 | 0.71 | -0.12 | 0.21 | 1.983 | 12.2 | 11.7 | 12 | 11.7 | 119.7 | 120.1 | ML |  | 2022-07-25 | 1.9774 | 1.0018 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220727 | 0.74 | 0.96 | -0.37 | 0.64 | 1.62 | -0.37 | 0.58 | 1.972 | 9.1 | 8.6 | 8.2 | 8.6 | 119.9 | 120.2 | 0.33 | 0.43 | -0.14 | 0.26 | 0.82 | -0.16 | 0.28 | 1.977 | 11.8 | 12.2 | 12.2 | 11.6 | 119.7 | 120.1 | AC |  | 2022-07-27 | 1.9774 | 0.99727 | 1.9991 | 0.98895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220801 | 0.76 | 0.99 | -0.38 | 0.63 | 1.42 | -0.38 | 0.5 | 1.979 | 8.9 | 8.8 | 8.4 | 8.4 | 119.9 | 120.2 | 0.57 | 0.47 | -0.31 | 0.31 | 0.88 | -0.3 | 0.29 | 1.983 | 11.8 | 12 | 12.2 | 11.7 | 119.6 | 120.1 | AC |  | 2022-08-01 | 1.9774 | 1.0008 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220803 | 0.59 | 0.98 | -0.33 | 0.52 | 1.42 | -0.36 | 0.44 | 1.984 | 8.9 | 8.9 | 8.4 | 8.4 | 119.9 | 120.2 | 0.41 | 0.48 | -0.18 | 0.28 | 0.91 | -0.2 | 0.29 | 1.988 | 11.9 | 12 | 12.2 | 11.8 | 119.6 | 120.1 | cvd |  | 2022-08-03 | 1.9774 | 1.0033 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220805 | 0.75 | 0.94 | -0.37 | 0.58 | 1.48 | -0.36 | 0.53 | 1.977 | 9.1 | 8.7 | 8.3 | 8.4 | 119.9 | 120.2 | 0.36 | 0.43 | -0.17 | 0.23 | 0.77 | -0.17 | 0.24 | 1.977 | 11.8 | 12.2 | 12.2 | 11.7 | 119.6 | 120.1 | GB |  | 2022-08-05 | 1.9774 | 0.9998 | 1.9991 | 0.98895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220808 | 0.58 | 0.74 | -0.27 | 0.51 | 1.31 | -0.31 | 0.45 | 1.979 | 9 | 8.8 | 7.9 | 8.6 | 119.9 | 120.2 | 0.35 | 0.27 | -0.15 | 0.13 | 0.69 | -0.19 | 0.17 | 1.984 | 11.7 | 12.1 | 12.2 | 11.6 | 119.6 | 120 | AC, Matrixx2 |  | 2022-08-08 | 1.9774 | 1.0008 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220809 | 0.41 | 0.96 | -0.21 | 0.48 | 1.29 | -0.25 | 0.39 | 1.97 | 9.2 | 8.7 | 8.4 | 8.4 | 119.8 | 120.2 | 0.34 | 0.51 | -0.04 | 0.2 | 0.88 | -0.09 | 0.12 | 1.981 | 11.8 | 12.1 | 12.3 | 11.5 | 119.5 | 120.2 | AC |  | 2022-08-09 | 1.9774 | 0.99626 | 1.9991 | 0.99095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220810 | 0.86 | 1.09 | -0.42 | 0.65 | 1.46 | -0.36 | 0.56 | 1.999 | 9.1 | 8.7 | 8.2 | 8.8 | 119.9 | 120.1 | 0.4 | 0.42 | -0.15 | 0.25 | 0.85 | -0.2 | 0.26 | 1.999 | 11.8 | 12.1 | 12 | 11.9 | 119.7 | 120.1 | AC |  | 2022-08-10 | 1.9774 | 1.0109 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220815 | 0.82 | 0.99 | -0.42 | 0.62 | 1.47 | -0.41 | 0.54 | 1.98 | 8.7 | 9.1 | 8.3 | 8.5 | 119.9 | 120.2 | 0.52 | 0.43 | -0.27 | 0.25 | 0.84 | -0.26 | 0.23 | 1.985 | 12 | 11.9 | 12.2 | 11.5 | 119.6 | 120.1 | GB |  | 2022-08-15 | 1.9774 | 1.0013 | 1.9991 | 0.99295 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220818 | 0.91 | 0.89 | -0.46 | 0.6 | 1.55 | -0.43 | 0.54 | 1.976 | 9.1 | 8.7 | 8.1 | 8.6 | 119.9 | 120.2 | 0.47 | 0.35 | -0.22 | 0.19 | 0.79 | -0.23 | 0.21 | 1.982 | 11.8 | 12.1 | 12.2 | 11.5 | 119.6 | 120.2 | AC |  | 2022-08-18 | 1.9774 | 0.99929 | 1.9991 | 0.99145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220819 | 0.71 | 1.06 | -0.36 | 0.71 | 1.53 | -0.36 | 0.59 | 1.973 | 8.6 | 9.1 | 8.5 | 8.3 | 119.9 | 120.2 | 0.37 | 0.46 | -0.2 | 0.26 | 0.89 | -0.21 | 0.28 | 1.975 | 11.8 | 12 | 12.2 | 11.8 | 119.7 | 120.1 | AC |  | 2022-08-19 | 1.9774 | 0.99777 | 1.9991 | 0.98794 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220822 | 0.87 | 1.02 | -0.43 | 0.68 | 1.56 | -0.4 | 0.59 | 1.961 | 8.6 | 9 | 8 | 8.6 | 119.9 | 120.2 | 0.36 | 0.36 | -0.18 | 0.17 | 0.8 | -0.21 | 0.21 | 1.973 | 11.5 | 11.4 | 11.8 | 11.3 | 119.7 | 120.1 | GB |  | 2022-08-22 | 1.9774 | 0.99171 | 1.9991 | 0.98694 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220824 | 0.82 | 1.12 | -0.39 | 0.68 | 1.93 | -0.45 | 0.54 | 1.97 | 8.6 | 9 | 8 | 8.6 | 119.9 | 120.2 | 0.46 | 0.3 | -0.2 | 0.02 | 0.72 | -0.24 | 0.03 | 1.98 | 11.4 | 11.4 | 11.6 | 11.4 | 119.6 | 120.3 | YM |  | 2022-08-24 | 1.9774 | 0.99626 | 1.9991 | 0.99045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220825 | 0.78 | 1 | -0.24 | 0.66 | 1.65 | -0.36 | 0.55 | 1.975 | 8.9 | 8.8 | 7.8 | 8.7 | 120 | 120.3 | 0.26 | 0.45 | -0.12 | 0.21 | 0.74 | -0.15 | 0.2 | 1.973 | 11.2 | 11.4 | 11.8 | 11.2 | 119.8 | 120.2 | AC |  | 2022-08-25 | 1.9774 | 0.99879 | 1.9991 | 0.98694 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220826 | 0.87 | 0.86 | -0.4 | 0.53 | 1.58 | -0.35 | 0.52 | 1.968 | 9 | 8.5 | 7.7 | 8.4 | 119.9 | 120.2 | 0.45 | 0.5 | -0.19 | 0.29 | 0.85 | -0.24 | 0.26 | 1.982 | 11.5 | 11.7 | 11.4 | 11.4 | 119.7 | 120.1 | YM |  | 2022-08-26 | 1.9774 | 0.99525 | 1.9991 | 0.99145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220829 | 0.82 | 0.93 | -0.35 | 0.59 | 1.52 | -0.38 | 0.51 | 1.961 | 9.1 | 8.4 | 8.1 | 8.3 | 119.9 | 120.2 | 0.37 | 0.37 | -0.17 | 0.19 | 0.8 | -0.22 | 0.21 | 1.981 | 11.5 | 11.6 | 11.7 | 11.2 | 119.7 | 120.1 | YM |  | 2022-08-29 | 1.9774 | 0.99171 | 1.9991 | 0.99095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220831 | 0.64 | 0.99 | -0.45 | 0.56 | 1.28 | -0.37 | 0.46 | 1.976 | 8.5 | 9 | 8 | 8.6 | 119.9 | 120.2 | 0.39 | 0.35 | -0.19 | 0.18 | 0.73 | -0.2 | 0.21 | 1.982 | 11.4 | 11.4 | 11.9 | 11.3 | 119.7 | 120.1 | ML, Matrixx 1 |  | 2022-08-31 | 1.9774 | 0.99929 | 1.9991 | 0.99145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220901 | 0.82 | 0.87 | -0.38 | 0.58 | 1.45 | -0.37 | 0.52 | 1.988 | 8.7 | 9.1 | 8 | 8.7 | 119.9 | 120.2 | 0.45 | 0.42 | -0.2 | 0.24 | 0.84 | -0.24 | 0.24 | 1.994 | 11.9 | 11.8 | 12.3 | 11.6 | 119.6 | 120.1 | GB |  | 2022-09-01 | 1.9774 | 1.0054 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220902 | 0.8 | 0.89 | -0.42 | 0.6 | 1.48 | -0.39 | 0.51 | 1.985 | 8.2 | 9.4 | 7.9 | 8.6 | 119.8 | 120.2 | 0.48 | 0.49 | -0.25 | 0.27 | 0.89 | -0.26 | 0.27 | 1.989 | 12.1 | 11.7 | 12.3 | 11.6 | 119.6 | 120.1 | YM |  | 2022-09-02 | 1.9774 | 1.0038 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220905 | 0.73 | 0.9 | -0.35 | 0.65 | 1.48 | -0.38 | 0.56 | 1.991 | 8.8 | 8.9 | 8.2 | 8.5 | 119.9 | 120.2 | 0.44 | 0.45 | -0.25 | 0.23 | 0.83 | -0.24 | 0.22 | 1.989 | 11.8 | 12.1 | 11.8 | 12.2 | 119.7 | 120.2 | CVD KW |  | 2022-09-05 | 1.9774 | 1.0069 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220907 | 1 | 0.88 | -0.49 | 0.6 | 1.67 | -0.43 | 0.55 | 1.968 | 8.8 | 8.8 | 8 | 8.5 | 119.9 | 120.2 | 0.28 | 0.41 | -0.11 | 0.19 | 0.66 | -0.14 | 0.18 | 1.986 | 11.8 | 12.1 | 12 | 12.1 | 119.7 | 120.1 | ML |  | 2022-09-07 | 1.9774 | 0.99525 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220908 | 0.87 | 0.96 | -0.57 | 0.65 | 1.84 | -0.46 | 0.6 | 1.988 | 8.9 | 9.1 | 8.2 | 8.7 | 119.9 | 120.2 | 0.36 | 0.44 | -0.19 | 0.22 | 0.8 | -0.22 | 0.22 | 1.988 | 11.9 | 11.9 | 12.2 | 11.7 | 119.6 | 120.1 | CVD |  | 2022-09-08 | 1.9774 | 1.0054 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220909 | 0.95 | 1.11 | -0.52 | 0.66 | 1.73 | -0.44 | 0.58 | 1.968 | 8.6 | 9 | 8.4 | 8.4 | 119.9 | 120.1 | 0.39 | 0.32 | -0.15 | 0.12 | 0.7 | -0.17 | 0.15 | 1.986 | 11.4 | 11.4 | 11.6 | 11.2 | 119.7 | 120.2 | YM |  | 2022-09-09 | 1.9774 | 0.99525 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220912 | 0.84 | 0.89 | -0.47 | 0.58 | 1.53 | -0.41 | 0.53 | 1.965 | 8.6 | 9 | 8.4 | 8.1 | 119.9 | 120.2 | 0.51 | 0.37 | -0.25 | 0.18 | 0.82 | -0.26 | 0.19 | 1.979 | 11.5 | 11.6 | 11.3 | 11.5 | 119.7 | 120.1 | ML |  | 2022-09-12 | 1.9774 | 0.99373 | 1.9991 | 0.98995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220914 | 0.99 | 0.84 | -0.47 | 0.54 | 1.53 | -0.41 | 0.53 | 1.981 | 9.3 | 8.5 | 7.8 | 8.7 | 119.9 | 120.2 | 0.44 | 0.31 | -0.21 | 0.12 | 0.78 | -0.25 | 0.14 | 1.983 | 11.5 | 11.7 | 11.7 | 11.3 | 119.6 | 120.2 | GB |  | 2022-09-14 | 1.9774 | 1.0018 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220915 | 0.75 | 0.93 | -0.43 | 0.66 | 1.56 | -0.4 | 0.58 | 1.986 | 8.6 | 9.2 | 8.2 | 8.5 | 119.9 | 120.2 | 0.48 | 0.49 | -0.25 | 0.25 | 0.85 | -0.25 | 0.25 | 1.992 | 11.8 | 11.7 | 11.8 | 11.8 | 119.7 | 120.1 | GB |  | 2022-09-15 | 1.9774 | 1.0043 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220916 | 0.82 | 0.81 | -0.44 | 0.62 | 1.43 | -0.4 | 0.51 | 1.989 | 8.5 | 9.3 | 8.1 | 8.6 | 119.9 | 120.2 | 0.55 | 0.54 | -0.28 | 0.32 | 0.97 | -0.29 | 0.31 | 1.984 | 11.5 | 11.4 | 11.4 | 11.5 | 119.7 | 120.2 | YM |  | 2022-09-16 | 1.9774 | 1.0059 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220919 | 0.91 | 0.84 | -0.49 | 0.59 | 1.57 | -0.44 | 0.53 | 1.987 | 8.7 | 9.1 | 8.4 | 8.4 | 119.9 | 120.2 | 0.45 | 0.27 | -0.22 | 0.1 | 0.68 | -0.25 | 0.11 | 1.981 | 11.4 | 11.4 | 11.6 | 11.3 | 119.7 | 120.1 | ML |  | 2022-09-19 | 1.9774 | 1.0049 | 1.9991 | 0.99095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220921 | 0.72 | 0.83 | -0.36 | 0.51 | 1.2 | -0.35 | 0.44 | 1.981 | 8.8 | 9 | 8 | 8.5 | 119.9 | 120.2 | 0.38 | 0.52 | -0.15 | 0.3 | 0.87 | -0.19 | 0.29 | 1.986 | 11.4 | 11.4 | 11.7 | 11.2 | 119.7 | 120 | GB |  | 2022-09-21 | 1.9774 | 1.0018 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220922 | 0.8 | 0.93 | -0.42 | 0.6 | 1.55 | -0.38 | 0.54 | 1.984 | 8.8 | 9 | 7.9 | 8.8 | 119.9 | 120.2 | 0.54 | 0.51 | -0.26 | 0.33 | 0.96 | -0.26 | 0.32 | 1.992 | 11.9 | 12 | 12.4 | 11.7 | 119.6 | 120.2 | GB |  | 2022-09-22 | 1.9774 | 1.0033 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220923 | 0.73 | 0.79 | -0.46 | 0.52 | 1.41 | -0.4 | 0.48 | 1.989 | 8.5 | 9.2 | 8.2 | 8.6 | 119.9 | 120.2 | 0.33 | 0.41 | -0.15 | 0.2 | 0.69 | -0.16 | 0.17 | 1.979 | 11.8 | 12.1 | 11.9 | 12 | 119.7 | 120.1 | YM |  | 2022-09-23 | 1.9774 | 1.0059 | 1.9991 | 0.98995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220926 | 0.94 | 0.77 | -0.51 | 0.56 | 1.44 | -0.42 | 0.5 | 1.978 | 8.8 | 9 | 8.4 | 8.3 | 119.9 | 120.2 | 0.45 | 0.3 | -0.21 | 0.08 | 0.66 | -0.23 | 0.09 | 1.979 | 11.9 | 11.9 | 12 | 11.8 | 119.7 | 120.2 | YM |  | 2022-09-26 | 1.9774 | 1.0003 | 1.9991 | 0.98995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20220928 | 0.97 | 0.82 | -0.46 | 0.57 | 1.61 | -0.43 | 0.54 | 1.967 | 9 | 8.6 | 8.2 | 8.1 | 119.9 | 120.3 | 0.46 | 0.29 | -0.21 | 0.04 | 0.76 | -0.24 | 0.05 | 1.983 | 11.4 | 11.4 | 11.5 | 11.3 | 119.7 | 119.9 | CVD, MatriXX_2 |  | 2022-09-28 | 1.9774 | 0.99474 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221009 | 0.46 | 0.84 | -0.28 | 0.44 | 1.39 | -0.26 | 0.47 | 1.997 | 9 | 8.8 | 8.4 | 8.2 | 119.8 | 120.1 | 0.29 | 0.54 | -0.09 | 0.19 | 0.93 | -0.07 | 0.22 | 2.022 | 11.4 | 11.4 | 11.5 | 11.2 | 119.5 | 119.8 | AC, MatriXX2, afternoon measurements |  | 2022-10-09 | 1.9774 | 1.0099 | 1.9991 | 1.0115 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221011 | 0.39 | 0.84 | -0.23 | 0.44 | 1.28 | -0.26 | 0.48 | 2 | 8.1 | 9.5 | 8.4 | 8.3 | 119.7 | 120.1 | 0.33 | 0.51 | -0.1 | 0.17 | 0.94 | -0.06 | 0.19 | 2.023 | 11.7 | 11.6 | 11.8 | 11.3 | 119.4 | 120 | AC |  | 2022-10-11 | 1.9774 | 1.0114 | 1.9991 | 1.012 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221012 | 0.76 | 0.8 | -0.39 | 0.52 | 1.38 | -0.36 | 0.5 | 1.98 | 8.4 | 9 | 8.3 | 7.7 | 119.9 | 120.3 | 0.6 | 0.62 | -0.25 | 0.06 | 1.19 | -0.23 | 0.07 | 2.008 | 10.9 | 11.1 | 10.9 | 11 | 119.6 | 120 | ML,MatriXX_1 |  | 2022-10-12 | 1.9774 | 1.0013 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221013 | 0.81 | 0.77 | -0.43 | 0.51 | 1.27 | -0.35 | 0.47 | 1.986 | 9 | 8.7 | 8.1 | 8.2 | 119.9 | 120.2 | 0.35 | 0.46 | -0.13 | 0.22 | 0.78 | -0.15 | 0.25 | 2.008 | 11.5 | 11.5 | 11.5 | 11.2 | 119.5 | 120.1 | YM |  | 2022-10-13 | 1.9774 | 1.0043 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221014 | 0.8 | 0.79 | -0.42 | 0.54 | 1.29 | -0.36 | 0.48 | 1.975 | 8.4 | 9 | 8.1 | 8 | 119.9 | 120.2 | 0.46 | 0.49 | -0.21 | 0.16 | 0.88 | -0.21 | 0.15 | 1.998 | 11.2 | 11.1 | 11.2 | 11 | 119.6 | 120 | CVD Matrixx1 |  | 2022-10-14 | 1.9774 | 0.99879 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221021 | 0.85 | 0.71 | -0.37 | 0.46 | 1.25 | -0.34 | 0.46 | 1.984 | 9.1 | 8.4 | 7.8 | 8.4 | 119.9 | 120.2 | 0.41 | 0.43 | -0.18 | 0.24 | 0.73 | -0.21 | 0.22 | 2.005 | 11.5 | 11.6 | 11.3 | 11.5 | 119.6 | 120.1 | YM |  | 2022-10-21 | 1.9774 | 1.0033 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221024 | 0.66 | 0.7 | -0.34 | 0.5 | 1.15 | -0.3 | 0.43 | 2.004 | 8.4 | 9.1 | 8.5 | 7.7 | 119.9 | 120.2 | 0.37 | 0.57 | -0.17 | 0.17 | 0.97 | -0.18 | 0.13 | 2.016 | 11.3 | 11.3 | 11.1 | 11.1 | 119.6 | 119.7 | ML, MatriXX_2 |  | 2022-10-24 | 1.9774 | 1.0135 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221024 | 0.5 | 0.93 | -0.3 | 0.51 | 1.36 | -0.28 | 0.5 | 1.985 | 9 | 8.9 | 8.3 | 8.5 | 119.8 | 120.1 | 0.31 | 0.56 | -0.13 | 0.22 | 0.96 | -0.11 | 0.23 | 2.01 | 11.5 | 11.4 | 11.8 | 11.4 | 119.4 | 120 | YM |  | 2022-10-24 | 1.9774 | 1.0038 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221027 | 0.94 | 0.82 | -0.5 | 0.6 | 1.35 | -0.4 | 0.52 | 2.006 | 8.7 | 8.9 | 8.2 | 7.9 | 119.9 | 120.2 | 0.38 | 0.52 | -0.18 | 0.14 | 0.95 | -0.18 | 0.12 | 2.022 | 11.2 | 11.3 | 11 | 11.1 | 119.6 | 120 | ML |  | 2022-10-27 | 1.9774 | 1.0145 | 1.9991 | 1.0115 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221031 | 0.63 | 0.84 | -0.27 | 0.54 | 1.16 | -0.27 | 0.47 | 1.979 | 8.7 | 8.9 | 8.2 | 8.1 | 119.9 | 120.2 | 0.25 | 0.31 | -0.08 | 0.04 | 0.56 | -0.11 | 0.06 | 2.001 | 11.3 | 11.4 | 11.4 | 11.3 | 119.6 | 120.1 | ML |  | 2022-10-31 | 1.9774 | 1.0008 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221102 | 0.72 | 0.97 | -0.36 | 0.62 | 1.34 | -0.37 | 0.5 | 1.982 | 8.6 | 9 | 8.2 | 8.2 | 119.9 | 120.2 | 0.42 | 0.25 | -0.22 | 0.09 | 0.68 | -0.24 | 0.12 | 1.995 | 11.4 | 11.4 | 11.4 | 11.7 | 119.7 | 120.1 | KA |  | 2022-11-02 | 1.9774 | 1.0023 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221103 | 0.74 | 0.76 | -0.29 | 0.53 | 1.26 | -0.28 | 0.49 | 1.991 | 9 | 8.5 | 7.9 | 8.4 | 119.9 | 120.2 | 0.29 | 0.43 | -0.08 | 0.21 | 0.7 | -0.16 | 0.2 | 2.015 | 11.4 | 11.6 | 11.3 | 11.4 | 119.6 | 120 | AC |  | 2022-11-03 | 1.9774 | 1.0069 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221107 | 0.71 | 0.84 | -0.35 | 0.53 | 1.31 | -0.39 | 0.41 | 1.985 | 8.8 | 8.8 | 8.3 | 8 | 119.9 | 120.2 | 0.48 | 0.47 | -0.21 | 0.3 | 0.91 | -0.26 | 0.27 | 2.001 | 11.5 | 11.4 | 11.5 | 11.6 | 119.6 | 120.1 | ML |  | 2022-11-07 | 1.9774 | 1.0038 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221109 | 0.79 | 0.74 | -0.35 | 0.5 | 1.21 | -0.32 | 0.45 | 1.989 | 9.1 | 8.6 | 7.9 | 8.3 | 119.8 | 120.2 | 0.49 | 0.34 | -0.26 | 0.12 | 0.72 | -0.23 | 0.12 | 2 | 11.4 | 11.3 | 11.2 | 11.4 | 119.6 | 120.2 | ML ? |  | 2022-11-09 | 1.9774 | 1.0059 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221110 | 0.62 | 0.77 | -0.27 | 0.49 | 1.14 | -0.25 | 0.49 | 1.972 | 8.8 | 8.5 | 7.9 | 8.1 | 120 | 120.3 | 0.32 | 0.3 | -0.12 | 0.1 | 0.54 | -0.13 | 0.1 | 1.99 | 11.4 | 11.6 | 11.2 | 11.6 | 119.7 | 120.1 | ML |  | 2022-11-10 | 1.9774 | 0.99727 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221111 | 0.82 | 0.94 | -0.45 | 0.63 | 1.41 | -0.41 | 0.53 | 1.983 | 9 | 8.6 | 8.1 | 8.2 | 119.9 | 120.2 | 0.48 | 0.4 | -0.25 | 0.15 | 0.84 | -0.22 | 0.13 | 1.998 | 11.4 | 11.3 | 11.4 | 11.1 | 119.5 | 120 | CVD, Matrix 1 |  | 2022-11-11 | 1.9774 | 1.0028 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.77 | 0.8 | -0.39 | 0.51 | 1.17 | -0.31 | 0.44 | 1.979 | 8.8 | 8.7 | 8.2 | 7.9 | 119.9 | 120.3 | 0.27 | 0.31 | -0.12 | 0.06 | 0.46 | -0.12 | 0.07 | 1.997 | 11.3 | 11.3 | 11.2 | 11.4 | 119.6 | 120.1 | YM |  | 2022-11-18 | 1.9774 | 1.0008 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.86 | 0.82 | -0.42 | 0.57 | 1.31 | -0.36 | 0.48 | 1.985 | 8.7 | 9 | 8 | 8.1 | 119.8 | 120.2 | 0.3 | 0.49 | -0.11 | 0.23 | 0.73 | -0.14 | 0.18 | 2.007 | 11.4 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | ML, REFERENCE IMAGES | MatriXX_1 | 2022-11-18 | 1.9774 | 1.0038 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.81 | 0.8 | -0.41 | 0.51 | 1.23 | -0.36 | 0.47 | 1.982 | 8.7 | 9 | 8.1 | 8.1 | 119.8 | 120.2 | 0.3 | 0.47 | -0.12 | 0.12 | 0.74 | -0.12 | 0.13 | 2.008 | 11.3 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | ML | MatriXX_1 | 2022-11-18 | 1.9774 | 1.0023 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.84 | 0.75 | -0.45 | 0.53 | 1.24 | -0.35 | 0.45 | 1.985 | 8.7 | 9 | 8 | 8.2 | 119.8 | 120.2 | 0.35 | 0.44 | -0.12 | 0.17 | 0.7 | -0.12 | 0.15 | 2.006 | 11.4 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | ML | MatriXX_1 | 2022-11-18 | 1.9774 | 1.0038 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.5 | 0.91 | -0.29 | 0.45 | 1.36 | -0.27 | 0.45 | 1.987 | 8.9 | 8.9 | 8.2 | 8.3 | 119.8 | 120.2 | 0.33 | 0.57 | -0.1 | 0.21 | 0.88 | -0.09 | 0.22 | 2.009 | 11.4 | 11.4 | 11.5 | 11.2 | 119.5 | 119.9 | ML, REFERENCE IMAGES | MatriXX_2 | 2022-11-18 | 1.9774 | 1.0049 | 1.9991 | 1.005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.49 | 0.85 | -0.26 | 0.46 | 1.26 | -0.26 | 0.45 | 1.987 | 8.9 | 8.9 | 8.1 | 8.3 | 119.8 | 120.2 | 0.37 | 0.56 | -0.1 | 0.22 | 0.92 | -0.09 | 0.24 | 2.01 | 11.4 | 11.4 | 11.5 | 11.2 | 119.5 | 119.9 | ML | MatriXX_2 | 2022-11-18 | 1.9774 | 1.0049 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.46 | 0.85 | -0.24 | 0.4 | 1.24 | -0.24 | 0.42 | 1.987 | 9 | 8.8 | 8.2 | 8.3 | 119.8 | 120.2 | 0.47 | 0.53 | -0.13 | 0.16 | 0.95 | -0.1 | 0.19 | 2.01 | 11.4 | 11.4 | 11.5 | 11.2 | 119.5 | 119.9 | ML | MatriXX_2 | 2022-11-18 | 1.9774 | 1.0049 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.51 | 0.58 | -0.3 | 0.18 | 1.1 | -0.32 | 0.18 | 1.987 | 8.7 | 9.2 | 8.1 | 8.3 | 119.8 | 120.2 | 0.36 | 0.39 | -0.17 | 0.03 | 0.71 | -0.16 | 0 | 2.007 | 11.5 | 11.5 | 11.3 | 11.2 | 119.4 | 119.9 | ML, REFERENCE IMAGES | MatriXX_3 | 2022-11-18 | 1.9774 | 1.0049 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.53 | 0.54 | -0.33 | 0.14 | 1.05 | -0.3 | 0.17 | 1.983 | 8.7 | 9.2 | 8.2 | 8.2 | 119.7 | 120.2 | 0.32 | 0.33 | -0.14 | 0.02 | 0.65 | -0.16 | -0.03 | 2.006 | 11.6 | 11.5 | 11.3 | 11.2 | 119.4 | 119.9 | ML | MatriXX_3 | 2022-11-18 | 1.9774 | 1.0028 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221118 | 0.55 | 0.61 | -0.36 | 0.21 | 1.05 | -0.31 | 0.21 | 1.988 | 8.6 | 9.3 | 8.1 | 8.3 | 119.7 | 120.2 | 0.38 | 0.4 | -0.14 | 0.1 | 0.69 | -0.13 | 0 | 2.008 | 11.5 | 11.5 | 11.3 | 11.2 | 119.4 | 119.9 | ML | MatriXX_3 | 2022-11-18 | 1.9774 | 1.0054 | 1.9991 | 1.0045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221121 | 0.83 | 0.74 | -0.44 | 0.52 | 1.31 | -0.37 | 0.41 | 1.991 | 8.3 | 9.2 | 8 | 8.1 | 119.8 | 120.3 | 0.37 | 0.42 | -0.15 | 0.02 | 0.58 | -0.17 | 0 | 2.004 | 11.7 | 11.5 | 11.2 | 11.4 | 119.6 | 120 | ML |  | 2022-11-21 | 1.9774 | 1.0069 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221123 | 0.94 | 0.81 | -0.47 | 0.53 | 1.33 | -0.39 | 0.45 | 1.983 | 8.9 | 8.8 | 8 | 8.3 | 119.9 | 120.2 | 0.31 | 0.29 | -0.14 | 0.08 | 0.6 | -0.17 | 0.05 | 1.997 | 11.4 | 11.4 | 11.5 | 11.3 | 119.6 | 120.1 | YM |  | 2022-11-23 | 1.9774 | 1.0028 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221124 | 0.87 | 0.61 | -0.41 | 0.43 | 1.39 | -0.39 | 0.44 | 1.963 | 9.2 | 8 | 7.4 | 8.5 | 119.9 | 120.2 | 0.4 | 0.5 | -0.03 | -0.01 | 0.8 | -0.08 | 0.07 | 1.984 | 11.3 | 11.2 | 11.1 | 11.2 | 119.6 | 120.1 | CVD |  | 2022-11-24 | 1.9774 | 0.99272 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221125 | 0.81 | 0.79 | -0.41 | 0.5 | 1.31 | -0.35 | 0.42 | 1.994 | 8.1 | 9.4 | 8.4 | 7.9 | 119.8 | 120.2 | 0.28 | 0.39 | -0.12 | -0.07 | 0.62 | -0.13 | -0.06 | 1.996 | 11.6 | 11.5 | 11.1 | 11.3 | 119.6 | 120.1 | GB |  | 2022-11-25 | 1.9774 | 1.0084 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221128 | 0.93 | 0.83 | -0.52 | 0.51 | 1.25 | -0.4 | 0.42 | 1.992 | 8.3 | 9.2 | 8 | 8.1 | 119.9 | 120.3 | 0.38 | 0.44 | -0.1 | 0 | 0.75 | -0.16 | 0.02 | 2.005 | 11.6 | 11.4 | 11.1 | 11.2 | 119.7 | 120.1 | CVD, Matlab R2018b, new Citrix portal |  | 2022-11-28 | 1.9774 | 1.0074 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221130 | 0.93 | 0.85 | -0.46 | 0.55 | 1.37 | -0.39 | 0.42 | 1.987 | 8.6 | 9 | 8.2 | 8 | 119.9 | 120.3 | 0.28 | 0.43 | -0.11 | -0.1 | 0.55 | -0.11 | -0.06 | 1.994 | 11.4 | 11.3 | 11.4 | 11.3 | 119.7 | 120 | CVD, Matlab R2018b, new Citrix portal |  | 2022-11-30 | 1.9774 | 1.0049 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221202 | 0.83 | 0.65 | -0.43 | 0.44 | 1.25 | -0.4 | 0.36 | 1.989 | 8.7 | 9 | 7.7 | 8.6 | 119.9 | 120.2 | 0.28 | 0.36 | -0.12 | -0.01 | 0.62 | -0.15 | -0.05 | 1.998 | 11.4 | 11.3 | 11.4 | 11.3 | 119.6 | 120 | CVD |  | 2022-12-02 | 1.9774 | 1.0059 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221205 | 0.94 | 0.9 | -0.48 | 0.59 | 1.33 | -0.39 | 0.49 | 1.954 | 8.7 | 8.6 | 7.9 | 8.1 | 119.9 | 120.2 | 0.33 | 0.34 | -0.13 | 0.04 | 0.59 | -0.17 | 0.02 | 1.997 | 11.4 | 11.4 | 11.3 | 11.3 | 119.6 | 120.1 | ML |  | 2022-12-05 | 1.9774 | 0.98817 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221207 | 0.83 | 0.83 | -0.44 | 0.47 | 1.23 | -0.39 | 0.38 | 2.005 | 8.2 | 9.3 | 7.9 | 8 | 119.8 | 120.3 | 0.34 | 0.45 | -0.14 | -0.02 | 0.78 | -0.09 | -0.04 | 2.016 | 11.4 | 11.4 | 11.1 | 11.1 | 119.7 | 120 | ML |  | 2022-12-07 | 1.9774 | 1.014 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221208 | 0.92 | 0.85 | -0.46 | 0.55 | 1.39 | -0.39 | 0.45 | 1.981 | 8.6 | 9 | 8.5 | 7.8 | 119.9 | 120.2 | 0.33 | 0.34 | -0.16 | 0.04 | 0.57 | -0.16 | 0.03 | 1.994 | 11.3 | 11.4 | 11.4 | 11.9 | 119.7 | 120.1 | ML |  | 2022-12-08 | 1.9774 | 1.0018 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221209 | 0.85 | 0.62 | -0.35 | 0.39 | 1.26 | -0.32 | 0.43 | 1.987 | 9.2 | 8.3 | 8 | 8.3 | 119.9 | 120.2 | 0.3 | 0.34 | -0.13 | -0.02 | 0.52 | -0.12 | -0.02 | 2.001 | 11.4 | 11.6 | 11.2 | 11.2 | 119.6 | 120.1 | ML |  | 2022-12-09 | 1.9774 | 1.0049 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221213 | 0.91 | 0.73 | -0.41 | 0.44 | 1.31 | -0.37 | 0.45 | 1.981 | 9 | 8.6 | 7.9 | 8.2 | 119.9 | 120.2 | 0.39 | 0.41 | -0.14 | 0.08 | 0.83 | -0.15 | 0.11 | 1.997 | 11.3 | 11.3 | 11.1 | 11.2 | 119.6 | 120 | AC |  | 2022-12-13 | 1.9774 | 1.0018 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221213 | 0.91 | 0.73 | -0.41 | 0.44 | 1.31 | -0.37 | 0.45 | 1.981 | 9 | 8.6 | 7.9 | 8.2 | 119.9 | 120.2 | 0.39 | 0.41 | -0.14 | 0.08 | 0.83 | -0.15 | 0.11 | 1.997 | 11.3 | 11.3 | 11.1 | 11.2 | 119.6 | 120.1 | ML | MatriXX_1 | 2022-12-13 | 1.9774 | 1.0018 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221215 | 0.89 | 0.89 | -0.45 | 0.5 | 1.3 | -0.41 | 0.42 | 1.978 | 8.4 | 9.2 | 8 | 8.2 | 119.9 | 120.2 | 0.3 | 0.38 | -0.13 | 0.01 | 0.56 | -0.15 | 0 | 1.994 | 11.6 | 11.4 | 11.3 | 11.3 | 119.6 | 119.9 | AC, MatriXX_2 |  | 2022-12-15 | 1.9774 | 1.0003 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221219 | 0.51 | 1.01 | -0.3 | 0.55 | 1.46 | -0.28 | 0.5 | 1.994 | 9.4 | 8.3 | 8.1 | 8.5 | 119.8 | 120.1 | 0.32 | 0.61 | -0.16 | 0.25 | 1 | -0.15 | 0.26 | 2.016 | 11.6 | 11.7 | 11.5 | 11.2 | 119.5 | 119.9 | AC, MatriXX_2 |  | 2022-12-19 | 1.9774 | 1.0084 | 1.9991 | 1.0085 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221221 | 0.59 | 0.79 | -0.31 | 0.33 | 1.27 | -0.29 | 0.32 | 1.995 | 8.8 | 9.1 | 8.9 | 7.5 | 119.8 | 120.2 | 0.33 | 0.6 | -0.13 | 0.23 | 0.96 | -0.12 | 0.24 | 2.014 | 11.4 | 11.4 | 11.7 | 11.6 | 119.6 | 119.9 | AC, MatriXX_2 |  | 2022-12-21 | 1.9774 | 1.0089 | 1.9991 | 1.0075 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221222 | 0.57 | 0.88 | -0.23 | 0.46 | 1.38 | -0.25 | 0.46 | 1.995 | 9.4 | 8.3 | 8.3 | 8.2 | 119.8 | 120.1 | 0.36 | 0.58 | -0.13 | 0.25 | 0.96 | -0.11 | 0.28 | 2.015 | 11.5 | 11.6 | 11.4 | 11.1 | 119.4 | 119.9 | AC, MatriXX_2 |  | 2022-12-22 | 1.9774 | 1.0089 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221223 | 0.42 | 0.87 | -0.24 | 0.41 | 1.4 | -0.25 | 0.47 | 1.99 | 8.6 | 9.2 | 8.5 | 8 | 119.8 | 120.2 | 0.29 | 0.51 | -0.08 | 0.21 | 0.92 | -0.07 | 0.25 | 2.01 | 11.4 | 11.4 | 11.6 | 11.5 | 119.5 | 119.9 | AC, MatriXX_2 |  | 2022-12-23 | 1.9774 | 1.0064 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221227 | 0.54 | 0.92 | -0.22 | 0.48 | 1.41 | -0.25 | 0.46 | 1.987 | 9.4 | 8.4 | 8.4 | 8.2 | 119.8 | 120.2 | 0.32 | 0.5 | -0.11 | 0.2 | 0.88 | -0.1 | 0.22 | 2.011 | 11.6 | 11.5 | 11.6 | 11.2 | 119.5 | 120 | AC, MatriXX_2 |  | 2022-12-27 | 1.9774 | 1.0049 | 1.9991 | 1.006 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221229 | 0.4 | 0.83 | -0.23 | 0.47 | 1.32 | -0.24 | 0.46 | 2 | 8.9 | 8.8 | 8.2 | 8.4 | 119.8 | 120.2 | 0.24 | 0.68 | -0.07 | 0.33 | 0.89 | -0.1 | 0.27 | 2.019 | 11.4 | 11.3 | 11.5 | 11.4 | 119.5 | 119.9 | ML, MatriXX_2 |  | 2022-12-29 | 1.9774 | 1.0114 | 1.9991 | 1.01 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20221230 | 0.52 | 0.89 | -0.23 | 0.43 | 1.38 | -0.28 | 0.41 | 1.986 | 9.4 | 8.2 | 8.5 | 8 | 119.8 | 120.2 | 0.38 | 0.53 | -0.14 | 0.24 | 1 | -0.14 | 0.27 | 2.007 | 11.6 | 11.7 | 11.6 | 11.5 | 119.5 | 119.9 | ML, MatriXX_2 |  | 2022-12-30 | 1.9774 | 1.0043 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230102 | 0.62 | 0.72 | -0.31 | 0.34 | 1.25 | -0.28 | 0.31 | 1.98 | 9 | 8.9 | 8.8 | 7.9 | 119.8 | 120.2 | 0.37 | 0.54 | -0.09 | 0.23 | 0.9 | -0.1 | 0.2 | 1.994 | 11.5 | 11.5 | 11.7 | 11.3 | 119.5 | 119.9 | ML, MatriXX_2 |  | 2023-01-02 | 1.9774 | 1.0013 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230110 | 0.53 | 0.96 | -0.33 | 0.55 | 1.45 | -0.31 | 0.55 | 1.974 | 8.9 | 8.9 | 8.3 | 8.2 | 119.8 | 120.2 | 0.34 | 0.56 | -0.13 | 0.21 | 1.04 | -0.13 | 0.24 | 1.996 | 11.5 | 11.4 | 11.8 | 11.3 | 119.5 | 120 | ML, MatriXX_2 |  | 2023-01-10 | 1.9774 | 0.99828 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230113 | 0.61 | 0.98 | -0.37 | 0.48 | 1.46 | -0.33 | 0.51 | 1.976 | 8.9 | 9 | 8.4 | 8.1 | 119.8 | 120.2 | 0.34 | 0.63 | -0.14 | 0.25 | 1 | -0.13 | 0.28 | 1.996 | 11.5 | 11.4 | 11.6 | 11.5 | 119.6 |  | ML, MatriXX 2 |  | 2023-01-13 | 1.9774 | 0.99929 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230116 | 0.45 | 0.79 | -0.25 | 0.4 | 1.22 | -0.24 | 0.39 | 1.976 | 8.7 | 9.1 | 8.8 | 7.9 | 119.8 | 120.2 | 0.35 | 0.57 | -0.13 | 0.25 | 0.88 | -0.11 | 0.24 | 1.998 | 11.5 | 11.5 | 11.8 | 11.3 | 119.5 | 119.9 | ML, MatriXX 2 |  | 2023-01-16 | 1.9774 | 0.99929 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230118 | 0.42 | 0.87 | -0.22 | 0.45 | 1.28 | -0.19 | 0.47 | 1.976 | 8.8 | 9 | 8.4 | 8.2 | 119.9 | 120.2 | 0.3 | 0.55 | -0.08 | 0.23 | 0.93 | -0.08 | 0.23 | 1.998 | 11.5 | 11.4 | 11.6 | 11.5 | 119.6 | 120 | ML, MatriXX 2 |  | 2023-01-18 | 1.9774 | 0.99929 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230120 | 0.51 | 0.83 | -0.27 | 0.38 | 1.18 | -0.24 | 0.35 | 1.976 | 8.8 | 9 | 8.9 | 7.7 | 119.9 | 120.2 | 0.35 | 0.6 | -0.14 | 0.23 | 0.93 | -0.12 | 0.21 | 1.999 | 11.4 | 11.4 | 11.6 | 11.2 | 119.5 | 119.9 | ML, MatriXX 2 |  | 2023-01-20 | 1.9774 | 0.99929 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230124 | 0.42 | 0.95 | -0.23 | 0.49 | 1.31 | -0.24 | 0.47 | 1.981 | 9 | 8.7 | 8.3 | 8.3 | 119.9 | 120.2 | 0.34 | 0.54 | -0.14 | 0.2 | 0.88 | -0.12 | 0.2 | 2.002 | 11.5 | 11.5 | 11.8 | 11.3 | 119.5 | 119.9 | ML, MatriXX 2 |  | 2023-01-24 | 1.9774 | 1.0018 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230125 | 0.41 | 0.89 | -0.21 | 0.46 | 1.35 | -0.21 | 0.46 | 1.982 | 9.6 | 7.7 | 8.1 | 8.4 | 119.7 | 120.2 | 0.36 | 0.55 | -0.11 | 0.2 | 0.86 | -0.1 | 0.2 | 2.001 | 11.7 | 12 | 11.5 | 11.4 | 119.5 | 120 | ML, MatriXX 2 |  | 2023-01-25 | 1.9774 | 1.0023 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230130 | 0.54 | 0.99 | -0.27 | 0.53 | 1.49 | -0.24 | 0.5 | 1.977 | 9.3 | 8.4 | 8.4 | 8.4 | 119.8 | 120.2 | 0.32 | 0.52 | -0.11 | 0.2 | 0.9 | -0.11 | 0.22 | 2 | 11.6 | 11.8 | 11.9 | 11.4 | 119.5 | 119.9 | ML, MatriXX 2 |  | 2023-01-30 | 1.9774 | 0.9998 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230209 | 0.83 | 0.64 | -0.44 | 0.4 | 1.17 | -0.36 | 0.34 | 1.97 | 8.4 | 9.1 | 8 | 8.2 | 119.9 | 120.2 | 0.44 | 0.31 | -0.21 | 0.14 | 0.71 | -0.18 | 0.15 | 1.993 | 11.5 | 11.4 | 11.7 | 11.4 | 119.6 | 120 | ML, MatriXX1 |  | 2023-02-09 | 1.9774 | 0.99626 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230210 | 0.71 | 0.72 | -0.37 | 0.46 | 1.21 | -0.32 | 0.4 | 1.972 | 9 | 8.7 | 7.8 | 8.5 | 119.9 | 120.2 | 0.37 | 0.26 | -0.13 | 0.07 | 0.6 | -0.13 | 0.06 | 1.991 | 11.4 | 11.4 | 11.7 | 11.4 | 119.6 | 120 | ML, MatriXX1 |  | 2023-02-10 | 1.9774 | 0.99727 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230213 | 0.87 | 0.76 | -0.45 | 0.45 | 1.22 | -0.37 | 0.37 | 1.961 | 8.5 | 9.1 | 7.9 | 8.3 | 119.9 | 120.3 | 0.27 | 0.4 | -0.08 | -0.05 | 0.6 | -0.08 | -0.06 | 1.978 | 11.4 | 11.5 | 11.8 | 11.5 | 119.6 | 120 | ML, MatriXX1 |  | 2023-02-13 | 1.9774 | 0.99171 | 1.9991 | 0.98945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230214 | 0.86 | 0.75 | -0.42 | 0.49 | 1.2 | -0.32 | 0.43 | 1.962 | 9 | 8.7 | 7.9 | 8.4 | 119.9 | 120.2 | 0.29 | 0.36 | -0.11 | -0.08 | 0.52 | -0.09 | -0.06 | 1.974 | 11.4 | 11.6 | 11.8 | 11.5 | 119.6 | 120.1 | ML, MatriXX1 |  | 2023-02-14 | 1.9774 | 0.99221 | 1.9991 | 0.98744 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230216 | 0.76 | 0.87 | -0.37 | 0.55 | 1.27 | -0.31 | 0.48 | 1.971 | 8.7 | 9 | 7.9 | 8.5 | 119.8 | 120.2 | 0.32 | 0.39 | -0.13 | -0.03 | 0.53 | -0.11 | -0.01 | 1.988 | 11.3 | 11.4 | 11.6 | 11.5 | 119.6 | 120 | ML, MatriXX1 |  | 2023-02-16 | 1.9774 | 0.99676 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230216 | 0.76 | 0.87 | -0.37 | 0.55 | 1.27 | -0.31 | 0.48 | 1.971 | 8.7 | 9 | 7.9 | 8.5 | 119.8 | 120.2 | 0.32 | 0.39 | -0.13 | -0.03 | 0.53 | -0.11 | -0.01 | 1.988 | 11.3 | 11.4 | 11.6 | 11.5 | 119.6 | 120 | JR | MatriXX_1 | 2023-02-16 | 1.9774 | 0.99676 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230220 | 0.81 | 0.77 | -0.41 | 0.5 | 1.2 | -0.3 | 0.44 | 1.968 | 8.3 | 9.2 | 8 | 8.3 | 119.9 | 120.2 | 0.35 | 0.42 | -0.14 | 0.02 | 0.58 | -0.11 | 0.01 | 1.984 | 11.6 | 11.5 | 11.8 | 11.5 | 119.6 | 120 | ML, MatriXX1 |  | 2023-02-20 | 1.9774 | 0.99525 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[29 empty] | Intermission _ Matrixx from DailyQA |  | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 2
20230220 | 0.81 | 0.77 | -0.41 | 0.5 | 1.2 | -0.3 | 0.44 | 1.968 | 8.3 | 9.2 | 8 | 8.3 | 119.9 | 120.2 | 0.35 | 0.42 | -0.14 | 0.02 | 0.58 | -0.11 | 0.01 | 1.984 | 11.6 | 11.5 | 11.8 | 11.5 | 119.6 | 120 | AJ | MatriXX_1 | 2023-02-20 | 1.9774 | 0.99525 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230221 | 0.84 | 0.75 | -0.42 | 0.47 | 1.3 | -0.34 | 0.41 | 1.956 | 8.5 | 9.1 | 8.2 | 8.1 | 119.9 | 120.2 | 0.28 | 0.36 | -0.07 | 0.11 | 0.58 | -0.09 | 0.09 | 1.977 | 11.6 | 11.5 | 11.8 | 11.6 | 119.6 | 120 | AE, Training session, beam not tunned | MatriXX_1 | 2023-02-21 | 1.9774 | 0.98918 | 1.9991 | 0.98895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230221 | 0.86 | 0.79 | -0.45 | 0.47 | 1.2 | -0.36 | 0.4 | 1.955 | 7.9 | 9.5 | 8.3 | 8.2 | 119.8 | 120.2 | 0.37 | 0.39 | -0.16 | -0.01 | 0.54 | -0.13 | -0.02 | 1.973 | 11.8 | 11.6 | 11.7 | 11.5 | 119.6 | 120 | AR, Training session, beam not tunned | MatriXX_1 | 2023-02-21 | 1.9774 | 0.98867 | 1.9991 | 0.98694 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230224 | 0.97 | 0.73 | -0.46 | 0.43 | 1.24 | -0.35 | 0.38 | 1.957 | 8.6 | 9.1 | 8 | 8.3 | 119.8 | 120.2 | 0.4 | 0.45 | -0.19 | -0.04 | 0.61 | -0.14 | 0.01 | 1.975 | 11.4 | 11.4 | 11.7 | 11.6 | 119.6 | 119.9 | OJ | MatriXX_1 | 2023-02-24 | 1.9774 | 0.98968 | 1.9991 | 0.98794 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230227 | 1.04 | 0.77 | -0.57 | 0.45 | 1.29 | -0.38 | 0.33 | 1.952 | 8.6 | 9 | 8 | 8.3 | 119.9 | 120.2 | 0.2 | 0.41 | -0.06 | -0.09 | 0.51 | -0.03 | -0.03 | 1.963 | 11.5 | 11.4 | 11.7 | 11.5 | 119.6 | 120 | GK,AJ,AR,OJ,EH | MatriXX_1 | 2023-02-27 | 1.9774 | 0.98715 | 1.9991 | 0.98194 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230228 | 0.89 | 0.77 | -0.47 | 0.53 | 1.26 | -0.36 | 0.43 | 1.965 | 8.5 | 9.1 | 8.3 | 7.9 | 119.9 | 120.3 | 0.55 | 0.53 | -0.25 | -0.13 | 0.7 | -0.14 | -0.09 | 1.984 | 11.4 | 11.4 | 11.6 | 11.6 | 119.5 | 119.9 | GF | MatriXX_1 | 2023-02-28 | 1.9774 | 0.99373 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230301 | 0.84 | 0.72 | -0.42 | 0.45 | 1.23 | -0.33 | 0.4 | 1.966 | 8.8 | 8.9 | 7.7 | 8.4 | 119.9 | 120.3 | 0.28 | 0.39 | -0.12 | -0.05 | 0.65 | -0.12 | -0.06 | 1.97 | 11.3 | 11.3 | 11.5 | 11.3 | 119.7 | 120.1 | SM, AE | MatriXX_1 | 2023-03-01 | 1.9774 | 0.99423 | 1.9991 | 0.98544 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230302 | 0.82 | 0.83 | -0.36 | 0.53 | 1.19 | -0.26 | 0.46 | 1.966 | 8.7 | 8.9 | 7.8 | 8.4 | 119.8 | 120.2 | 0.36 | 0.36 | -0.12 | 0.16 | 0.56 | -0.09 | 0.16 | 1.982 | 11.3 | 11.3 | 11.6 | 11.4 | 119.6 | 120 | JR, Remeber to record the GI vaules/ML | MatriXX_1 | 2023-03-02 | 1.9774 | 0.99423 | 1.9991 | 0.99145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230303 | 0.86 | 0.87 | -0.46 | 0.57 | 1.34 | -0.35 | 0.45 | 1.964 | 8.6 | 9 | 7.8 | 8.4 | 119.9 | 120.2 | 0.29 | 0.39 | -0.11 | -0.07 | 0.59 | -0.1 | -0.07 | 1.97 | 11.3 | 11.3 | 11.7 | 11.5 | 119.6 | 120 | SM, AR | MatriXX_1 | 2023-03-03 | 1.9774 | 0.99322 | 1.9991 | 0.98544 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230306 | 0.78 | 0.74 | -0.38 | 0.43 | 1.19 | -0.32 | 0.4 | 1.989 | 8.4 | 9.1 | 8 | 8 | 119.9 | 120.3 | 0.27 | 0.37 | -0.09 | 0.02 | 0.51 | -0.08 | 0.02 | 1.981 | 11.5 | 11.5 | 11.7 | 11.5 | 119.6 | 120 | ML | MatriXX_1 | 2023-03-06 | 1.9774 | 1.0059 | 1.9991 | 0.99095 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230306 | [14 empty] | 0.3 | 0.34 | -0.12 | 0.04 | 0.55 | -0.09 | 0.03 | 1.98 | 11.5 | 11.5 | 11.7 | 11.4 | 119.5 | 120 | ML, only large cube was repeated | MatriXX_1 | 2023-03-06 | 1.9774 | 0 | 1.9991 | 0.99045 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230307 | 0.89 | 0.89 | -0.48 | 0.56 | 1.29 | -0.38 | 0.48 | 1.964 | 9 | 8.7 | 7.8 | 8.4 | 119.9 | 120.2 | 0.46 | 0.45 | 0.08 | 0.07 | 0.79 | 0.02 | 0.14 | 1.976 | 11.1 | 11.1 | 11 | 11.2 | 119.7 | 120 | ML | MatriXX_1 | 2023-03-07 | 1.9774 | 0.99322 | 1.9991 | 0.98844 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230307 | [14 empty] | 0.45 | 0.43 | 0.08 | 0.1 | 0.82 | 0.03 | 0.16 | 1.976 | 11.1 | 11.1 | 11 | 11.2 | 119.7 | 120 | ML, only large cube was repeated | MatriXX_1 | 2023-03-07 | 1.9774 | 0 | 1.9991 | 0.98844 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230308 | 0.78 | 0.75 | -0.39 | 0.49 | 1.12 | -0.32 | 0.41 | 1.963 | 8.6 | 9 | 8 | 8.2 | 119.9 | 120.2 | 0.18 | 0.41 | -0.03 | -0.09 | 0.56 | -0.05 | -0.06 | 1.97 | 11.4 | 11.4 | 11.6 | 11.5 | 119.6 | 120 | ML | MatriXX_1 | 2023-03-08 | 1.9774 | 0.99272 | 1.9991 | 0.98544 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230308 | 0.68 | 0.82 | -0.34 | 0.46 | 1.09 | -0.3 | 0.39 | 1.962 | 8.5 | 9.1 | 7.7 | 8.4 | 119.8 | 120.3 | 0.27 | 0.38 | -0.1 | -0.04 | 0.51 | -0.08 | -0.03 | 1.978 | 11.3 | 11.4 | 11.6 | 11.4 | 119.6 | 120 | ML, Large cube: region approx. 3x3cm (right, down corner of the field) where dose deviates from reference in av. -2.5% | MatriXX_1 | 2023-03-08 | 1.9774 | 0.99221 | 1.9991 | 0.98945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230309 | 0.85 | 0.69 | -0.43 | 0.46 | 1.18 | -0.32 | 0.39 | 1.971 | 8.5 | 9.1 | 8.1 | 8.1 | 119.9 | 120.2 | 0.28 | 0.33 | -0.07 | 0.05 | 0.48 | -0.08 | 0.02 | 1.989 | 11.4 | 11.4 | 11.6 | 11.5 | 119.6 | 119.9 | EH | MatriXX_1 | 2023-03-09 | 1.9774 | 0.99676 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230309 | 0.45 | 0.9 | -0.18 | 0.45 | 1.32 | -0.18 | 0.44 | 1.979 | 9.3 | 8.4 | 8.6 | 8 | 119.8 | 120.2 | 0.3 | 0.53 | -0.08 | 0.21 | 0.89 | -0.08 | 0.18 | 1.999 | 11.6 | 11.7 | 11.8 | 11.3 | 119.5 | 119.9 | ML | MatriXX_2 | 2023-03-09 | 1.9774 | 1.0008 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230310 | 0.81 | 0.63 | -0.41 | 0.38 | 1.21 | -0.32 | 0.41 | 1.97 | 8.9 | 8.5 | 8.2 | 7.9 | 119.9 | 120.3 | 0.27 | 0.49 | -0.13 | -0.22 | 0.74 | -0.12 | -0.16 | 1.983 | 11.3 | 11.4 | 11.6 | 11.4 | 119.6 | 120 | AR, GK | MatriXX_1 | 2023-03-10 | 1.9774 | 0.99626 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230313 | 0.41 | 0.62 | -0.23 | 0.25 | 1.09 | -0.28 | 0.19 | 1.959 | 9 | 8.8 | 8.1 | 8.5 | 119.8 | 120.1 | 0.34 | 0.35 | -0.17 | 0.03 | 0.72 | -0.15 | -0.06 | 1.986 | 11.6 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | AJ | MatriXX_3 | 2023-03-13 | 1.9774 | 0.99069 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230314 | 0.65 | 0.6 | -0.31 | 0.23 | 1.09 | -0.29 | 0.24 | 1.969 | 8.8 | 9.1 | 7.9 | 8.6 | 119.8 | 120.1 | 0.32 | 0.42 | -0.14 | -0.01 | 0.73 | -0.12 | -0.07 | 1.991 | 11.6 | 11.6 | 12 | 11.7 | 119.4 | 119.8 | SM, EH | MatriXX_3 | 2023-03-14 | 1.9774 | 0.99575 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230315 | 0.65 | 0.71 | -0.32 | 0.27 | 1.07 | -0.3 | 0.23 | 1.967 | 8.9 | 9 | 7.9 | 8.7 | 119.8 | 120.1 | 0.28 | 0.38 | -0.1 | -0.01 | 0.71 | -0.11 | -0.09 | 1.987 | 11.5 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | GF | MatriXX_3 | 2023-03-15 | 1.9774 | 0.99474 | 1.9991 | 0.99395 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230316 | 0.52 | 0.66 | -0.28 | 0.33 | 1.12 | -0.29 | 0.24 | 1.97 | 8.7 | 9.2 | 8 | 8.5 | 119.8 | 120.1 | 0.31 | 0.4 | -0.15 | 0.05 | 0.7 | -0.16 | -0.02 | 1.992 | 11.7 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | EH | MatriXX_3 | 2023-03-16 | 1.9774 | 0.99626 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230320 | 0.68 | 0.61 | -0.4 | 0.23 | 1.09 | -0.32 | 0.23 | 1.973 | 8.5 | 9.3 | 8.3 | 8.6 | 119.8 | 120.1 | 0.39 | 0.37 | -0.15 | 0.03 | 0.72 | -0.15 | -0.04 | 1.989 | 11.8 | 11.7 | 11.9 | 11.6 | 119.3 | 119.8 | AC | MatriXX_3 | 2023-03-20 | 1.9774 | 0.99777 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230320 | 0.47 | 0.64 | -0.26 | 0.17 | 1.16 | -0.28 | 0.18 | 1.968 | 8.6 | 9.3 | 8.2 | 8.5 | 119.7 | 120.1 | 0.31 | 0.38 | -0.15 | -0.01 | 0.75 | -0.16 | -0.06 | 1.987 | 11.8 | 11.7 | 11.9 | 11.5 | 119.3 | 119.8 | AC | MatriXX_3 | 2023-03-20 | 1.9774 | 0.99525 | 1.9991 | 0.99395 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230320 | 0.54 | 0.59 | -0.3 | 0.15 | 1.06 | -0.3 | 0.17 | 1.971 | 8.6 | 9.3 | 8.2 | 8.5 | 119.7 | 120.1 | 0.32 | 0.37 | -0.14 | 0 | 0.74 | -0.15 | -0.05 | 1.988 | 11.8 | 11.7 | 11.9 | 11.6 | 119.3 | 119.8 | AC | MatriXX_3 | 2023-03-20 | 1.9774 | 0.99676 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230321 | 0.68 | 0.65 | -0.34 | 0.27 | 1.05 | -0.29 | 0.23 | 1.974 | 8.4 | 9.4 | 8.1 | 8.6 | 119.7 | 120.1 | 0.33 | 0.37 | -0.11 | 0.05 | 0.73 | -0.12 | -0.04 | 1.989 | 11.7 | 11.6 | 11.8 | 11.5 | 119.4 | 119.9 | GF,AE | MatriXX_3 | 2023-03-21 | 1.9774 | 0.99828 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230322 | 0.51 | 0.63 | -0.29 | 0.21 | 1.12 | -0.29 | 0.18 | 1.969 | 8.4 | 9.3 | 8.2 | 8.4 | 119.8 | 120.2 | 0.28 | 0.34 | -0.1 | 0 | 0.67 | -0.13 | -0.05 | 1.989 | 11.9 | 11.7 | 11.9 | 11.6 | 119.4 | 119.8 | SF | MatriXX_3 | 2023-03-22 | 1.9774 | 0.99575 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230323 | 0.64 | 0.71 | -0.31 | 0.28 | 1.07 | -0.28 | 0.26 | 1.972 | 9.2 | 8.8 | 7.6 | 8.9 | 119.7 | 120.1 | 0.33 | 0.39 | -0.12 | 0 | 0.65 | -0.12 | -0.05 | 1.986 | 11.7 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | SM, OJ | MatriXX_3 | 2023-03-23 | 1.9774 | 0.99727 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230324 | 0.43 | 0.59 | -0.22 | 0.3 | 0.98 | -0.24 | 0.23 | 1.965 | 8.6 | 9.2 | 7.8 | 8.6 | 119.8 | 120.2 | 0.28 | 0.43 | -0.07 | -0.04 | 0.69 | -0.1 | -0.09 | 1.986 | 11.6 | 11.6 | 12.1 | 11.7 | 119.4 | 119.8 | SM, EH | MatriXX_3 | 2023-03-24 | 1.9774 | 0.99373 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230327 | 0.61 | 0.54 | -0.29 | 0.24 | 1.04 | -0.26 | 0.21 | 1.963 | 8.8 | 9.2 | 7.8 | 8.7 | 119.8 | 120.1 | 0.3 | 0.38 | -0.09 | 0.01 | 0.68 | -0.13 | -0.07 | 1.982 | 11.8 | 11.7 | 12.1 | 11.7 | 119.4 | 119.8 | AJ, GK | MatriXX_3 | 2023-03-27 | 1.9774 | 0.99272 | 1.9991 | 0.99145 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230328 | 0.47 | 0.64 | -0.26 | 0.33 | 1.07 | -0.27 | 0.25 | 1.963 | 9 | 8.9 | 7.6 | 8.8 | 119.8 | 120.1 | 0.29 | 0.34 | -0.12 | 0.02 | 0.72 | -0.13 | -0.06 | 1.976 | 11.4 | 11.4 | 11.7 | 11.5 | 119.4 | 119.8 | AE, GF | MatriXX_3 | 2023-03-28 | 1.9774 | 0.99272 | 1.9991 | 0.98844 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230329 | 0.51 | 0.59 | -0.32 | 0.31 | 1.04 | -0.28 | 0.23 | 1.969 | 8.9 | 9 | 7.9 | 8.6 | 119.8 | 120.1 | 0.29 | 0.35 | -0.12 | -0.02 | 0.72 | -0.14 | -0.08 | 1.984 | 11.7 | 11.6 | 11.7 | 11.5 | 119.5 | 119.9 | GK,ON,EH | MatriXX_3 | 2023-03-29 | 1.9774 | 0.99575 | 1.9991 | 0.99245 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230330 | 0.4 | 0.66 | -0.23 | 0.18 | 1.01 | -0.25 | 0.18 | 1.964 | 8.8 | 9.1 | 8 | 8.5 | 119.8 | 120.2 | 0.34 | 0.48 | -0.14 | -0.06 | 0.74 | -0.12 | -0.07 | 1.986 | 11.6 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | JR,ON | MatriXX_3 | 2023-03-30 | 1.9774 | 0.99322 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230331 | 0.54 | 0.64 | -0.25 | 0.32 | 1.08 | -0.26 | 0.28 | 1.963 | 8.7 | 9.1 | 8 | 8.6 | 119.8 | 120.1 | 0.34 | 0.34 | -0.12 | 0.02 | 0.69 | -0.13 | -0.03 | 1.985 | 11.6 | 11.7 | 12.1 | 11.6 | 119.4 | 119.8 | AR, SM | MatriXX_3 | 2023-03-31 | 1.9774 | 0.99272 | 1.9991 | 0.99295 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230403 | 0.5 | 0.63 | -0.27 | 0.18 | 1.07 | -0.26 | 0.15 | 1.975 | 9 | 9 | 8.1 | 8.5 | 119.8 | 120.2 | 0.34 | 0.38 | -0.14 | -0.01 | 0.73 | -0.14 | -0.04 | 1.996 | 11.6 | 11.6 | 12 | 11.6 | 119.3 | 119.8 | SM, AR | MatriXX_3 | 2023-04-03 | 1.9774 | 0.99879 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230404 | 0.75 | 0.59 | -0.37 | 0.23 | 1.04 | -0.35 | 0.2 | 1.97 | 8.4 | 9.5 | 8.1 | 8.5 | 119.7 | 120.1 | 0.29 | 0.36 | -0.1 | 0.01 | 0.74 | -0.15 | -0.09 | 1.988 | 12 | 11.8 | 12 | 11.6 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-04-04 | 1.9774 | 0.99626 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230405 | 0.6 | 0.65 | -0.24 | 0.24 | 1.1 | -0.24 | 0.24 | 1.973 | 9.3 | 8.5 | 7.5 | 8.8 | 119.8 | 120.1 | 0.27 | 0.32 | -0.12 | 0.03 | 0.66 | -0.14 | -0.05 | 1.993 | 11.6 | 11.7 | 11.9 | 11.5 | 119.4 | 119.8 | AE, OJ | MatriXX_3 | 2023-04-05 | 1.9774 | 0.99777 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230406 | 0.62 | 0.76 | -0.26 | 0.32 | 1.17 | -0.28 | 0.29 | 1.965 | 9.5 | 8 | 6.9 | 9.4 | 119.8 | 120.1 | 0.32 | 0.38 | -0.15 | -0.02 | 0.73 | -0.17 | -0.08 | 1.988 | 11.8 | 12.1 | 12.2 | 11.8 | 119.4 | 119.8 | EH | MatriXX_3 | 2023-04-06 | 1.9774 | 0.99373 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230407 | 0.53 | 0.62 | -0.29 | 0.2 | 1.06 | -0.28 | 0.2 | 1.97 | 8.5 | 9.3 | 8 | 8.5 | 119.8 | 120.2 | 0.26 | 0.4 | -0.09 | -0.05 | 0.73 | -0.12 | -0.07 | 1.991 | 11.7 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | SM, JR | MatriXX_3 | 2023-04-07 | 1.9774 | 0.99626 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230411 | 0.55 | 0.61 | -0.26 | 0.26 | 1.06 | -0.25 | 0.23 | 1.97 | 8.8 | 9.1 | 7.9 | 8.7 | 119.8 | 120.1 | 0.33 | 0.41 | -0.13 | -0.03 | 0.72 | -0.12 | -0.07 | 1.988 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | GF, EH, ON | MatriXX_3 | 2023-04-11 | 1.9774 | 0.99626 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230412 | 0.75 | 0.69 | -0.37 | 0.29 | 1.12 | -0.32 | 0.25 | 1.972 | 8.4 | 9.5 | 8 | 8.7 | 119.7 | 120.1 | 0.34 | 0.38 | -0.1 | 0.06 | 0.72 | -0.12 | -0.02 | 1.991 | 11.9 | 11.7 | 11.9 | 11.5 | 119.4 | 119.8 | JR,GF | MatriXX_3 | 2023-04-12 | 1.9774 | 0.99727 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230413 | 0.57 | 0.64 | -0.29 | 0.28 | 1.2 | -0.27 | 0.25 | 1.968 | 8.7 | 9.3 | 7.9 | 8.7 | 119.8 | 120.1 | 0.3 | 0.39 | -0.08 | 0.02 | 0.74 | -0.11 | -0.04 | 1.99 | 11.8 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | SM, AE | MatriXX_3 | 2023-04-13 | 1.9774 | 0.99525 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230414 | 0.53 | 0.64 | -0.28 | 0.29 | 1.04 | -0.29 | 0.25 | 1.963 | 8.8 | 9.1 | 7.8 | 8.7 | 119.8 | 120.1 | 0.31 | 0.36 | -0.11 | 0.03 | 0.72 | -0.13 | -0.04 | 1.986 | 11.7 | 11.7 | 11.9 | 11.6 | 119.4 | 119.9 | OJ, SM | MatriXX_3 | 2023-04-14 | 1.9774 | 0.99272 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230417 | 0.52 | 0.62 | -0.31 | 0.13 | 1.02 | -0.26 | 0.14 | 1.971 | 8.6 | 9.3 | 8.1 | 8.5 | 119.7 | 120.1 | 0.29 | 0.39 | -0.07 | 0 | 0.75 | -0.11 | -0.08 | 1.99 | 11.8 | 11.7 | 11.8 | 11.5 | 119.4 | 119.8 | AJ, OJ | MatriXX_3 | 2023-04-17 | 1.9774 | 0.99676 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230418 | 0.6 | 0.63 | -0.3 | 0.28 | 1.02 | -0.3 | 0.23 | 1.968 | 7.9 | 9.6 | 7.6 | 8.8 | 119.7 | 120.1 | 0.29 | 0.38 | -0.08 | -0.01 | 0.72 | -0.12 | -0.08 | 1.991 | 12.2 | 11.8 | 12.3 | 11.7 | 119.4 | 119.8 | GK,GF | MatriXX_3 | 2023-04-18 | 1.9774 | 0.99525 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230419 | 0.64 | 0.58 | -0.34 | 0.25 | 1.13 | -0.34 | 0.22 | 1.966 | 9.1 | 8.9 | 7.9 | 8.7 | 119.8 | 120.1 | 0.34 | 0.37 | -0.15 | 0 | 0.71 | -0.15 | -0.04 | 1.989 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | OJ, AR | MatriXX_3 | 2023-04-19 | 1.9774 | 0.99423 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230420 | 0.52 | 0.56 | -0.3 | 0.16 | 1.16 | -0.26 | 0.19 | 1.971 | 8.9 | 9.1 | 8.1 | 8.5 | 119.8 | 120.2 | 0.34 | 0.37 | -0.14 | -0.01 | 0.69 | -0.15 | -0.04 | 1.993 | 11.5 | 11.5 | 11.9 | 11.5 | 119.4 | 119.8 | ON, EH | MatriXX_3 | 2023-04-20 | 1.9774 | 0.99676 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230421 | 0.45 | 0.6 | -0.26 | 0.21 | 1.03 | -0.28 | 0.15 | 1.967 | 8.8 | 9.2 | 8.1 | 8.5 | 119.8 | 120.2 | 0.33 | 0.4 | -0.13 | -0.01 | 0.71 | -0.11 | -0.05 | 1.988 | 11.7 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | SM, ON | MatriXX_3 | 2023-04-21 | 1.9774 | 0.99474 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230424 | 0.48 | 0.62 | -0.28 | 0.25 | 1.14 | -0.3 | 0.18 | 1.966 | 8.9 | 9 | 8.1 | 8.5 | 119.8 | 120.2 | 0.3 | 0.38 | -0.08 | 0.01 | 0.69 | -0.11 | -0.07 | 1.987 | 11.5 | 11.5 | 11.9 | 11.6 | 119.4 | 119.9 | GK,AE | MatriXX_3 | 2023-04-24 | 1.9774 | 0.99423 | 1.9991 | 0.99395 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230425 | 0.49 | 0.54 | -0.26 | 0.19 | 0.99 | -0.28 | 0.18 | 1.964 | 8.6 | 9.2 | 8 | 8.4 | 119.8 | 120.2 | 0.27 | 0.4 | -0.1 | -0.03 | 0.69 | -0.1 | -0.04 | 1.988 | 11.7 | 11.6 | 11.7 | 11.4 | 119.4 | 119.9 | AE, GF | MatriXX_3 | 2023-04-25 | 1.9774 | 0.99322 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230426 | 0.66 | 0.52 | -0.33 | 0.22 | 1.04 | -0.28 | 0.21 | 1.992 | 8.8 | 9.2 | 7.6 | 8.6 | 119.8 | 120.2 | 0.36 | 0.3 | -0.16 | 0.02 | 0.71 | -0.16 | -0.05 | 2.001 | 11.3 | 11.4 | 11.8 | 11.4 | 119.4 | 119.9 | AR, GK | MatriXX_3 | 2023-04-26 | 1.9774 | 1.0074 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230427 | 0.51 | 0.57 | -0.26 | 0.27 | 1.12 | -0.27 | 0.24 | 1.98 | 8.8 | 9.2 | 7.9 | 8.5 | 119.8 | 120.2 | 0.33 | 0.33 | -0.17 | -0.01 | 0.72 | -0.15 | -0.04 | 1.998 | 11.5 | 11.6 | 11.8 | 11.5 | 119.4 | 119.9 | GK,OJ | MatriXX_3 | 2023-04-27 | 1.9774 | 1.0013 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230428 | 0.59 | 0.62 | -0.28 | 0.3 | 1.05 | -0.26 | 0.25 | 1.969 | 9 | 8.9 | 7.8 | 8.6 | 119.8 | 120.2 | 0.32 | 0.4 | -0.11 | -0.01 | 0.72 | -0.13 | -0.09 | 1.992 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | GK,JR | MatriXX_3 | 2023-04-28 | 1.9774 | 0.99575 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230502 | 0.49 | 0.57 | -0.28 | 0.2 | 1.1 | -0.27 | 0.17 | 1.971 | 8.7 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.3 | 0.39 | -0.1 | 0 | 0.71 | -0.13 | -0.07 | 1.993 | 11.5 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | AE, AR, EF | MatriXX_3 | 2023-05-02 | 1.9774 | 0.99676 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230503 | 0.59 | 0.63 | -0.32 | 0.16 | 1.08 | -0.28 | 0.14 | 1.97 | 8.8 | 9.2 | 8.2 | 8.3 | 119.8 | 120.2 | 0.3 | 0.33 | -0.14 | 0 | 0.73 | -0.15 | -0.05 | 1.99 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-05-03 | 1.9774 | 0.99626 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230504 | 0.75 | 0.56 | -0.39 | 0.25 | 1.04 | -0.31 | 0.22 | 1.972 | 8.4 | 9.4 | 7.8 | 8.6 | 119.7 | 120.1 | 0.38 | 0.35 | -0.16 | -0.01 | 0.71 | -0.15 | -0.06 | 1.995 | 11.9 | 11.7 | 11.9 | 11.6 | 119.4 | 119.8 | JR | MatriXX_3 | 2023-05-04 | 1.9774 | 0.99727 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230504 | 0.49 | 0.59 | -0.29 | 0.14 | 1.09 | -0.26 | 0.15 | 1.975 | 8.8 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.36 | 0.36 | -0.13 | -0.01 | 0.73 | -0.14 | -0.06 | 1.995 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | ON, EF | MatriXX_3 | 2023-05-04 | 1.9774 | 0.99879 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230505 | 0.43 | 0.6 | -0.24 | 0.21 | 0.97 | -0.24 | 0.17 | 1.971 | 8.7 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.3 | 0.33 | -0.12 | 0.02 | 0.74 | -0.14 | -0.06 | 1.992 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | SM, GK | MatriXX_3 | 2023-05-05 | 1.9774 | 0.99676 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230508 | 0.58 | 0.6 | -0.33 | 0.29 | 1.15 | -0.31 | 0.24 | 1.97 | 8.9 | 9.1 | 7.9 | 8.6 | 119.8 | 120.2 | 0.38 | 0.36 | -0.17 | -0.01 | 0.73 | -0.16 | -0.07 | 1.988 | 11.5 | 11.5 | 11.9 | 11.5 | 119.4 | 119.9 | OJ, SF | MatriXX_3 | 2023-05-08 | 1.9774 | 0.99626 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230509 | 0.65 | 0.63 | -0.36 | 0.35 | 1.11 | -0.32 | 0.28 | 1.967 | 8.7 | 9.2 | 7.1 | 8.9 | 119.8 | 120.2 | 0.37 | 0.35 | -0.16 | 0.01 | 0.7 | -0.15 | -0.03 | 1.983 | 11.6 | 11.6 | 11.7 | 11.4 | 119.5 | 120 | GF | MatriXX_3 | 2023-05-09 | 1.9774 | 0.99474 | 1.9991 | 0.99195 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230510 | 0.56 | 0.58 | -0.31 | 0.18 | 1.05 | -0.26 | 0.17 | 1.968 | 8.7 | 9.2 | 7.9 | 8.4 | 119.8 | 120.2 | 0.33 | 0.36 | -0.1 | 0.01 | 0.69 | -0.13 | -0.04 | 1.987 | 11.7 | 11.6 | 11.7 | 11.4 | 119.4 | 119.9 | ON | MatriXX_3 | 2023-05-10 | 1.9774 | 0.99525 | 1.9991 | 0.99395 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230511 | 0.81 | 0.59 | -0.39 | 0.26 | 1.09 | -0.33 | 0.22 | 1.973 | 8.4 | 9.4 | 7.7 | 8.7 | 119.7 | 120.1 | 0.35 | 0.4 | -0.16 | -0.02 | 0.78 | -0.17 | -0.06 | 1.991 | 11.9 | 11.7 | 11.9 | 11.5 | 119.4 | 119.9 | AR | MatriXX_3 | 2023-05-11 | 1.9774 | 0.99777 | 1.9991 | 0.99595 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230512 | 0.54 | 0.55 | -0.29 | 0.13 | 1.05 | -0.28 | 0.12 | 1.976 | 8.7 | 9.3 | 8.1 | 8.3 | 119.8 | 120.2 | 0.35 | 0.35 | -0.12 | 0.01 | 0.7 | -0.14 | -0.05 | 1.999 | 11.8 | 11.7 | 11.8 | 11.5 | 119.4 | 119.9 | SF, JR | MatriXX_3 | 2023-05-12 | 1.9774 | 0.99929 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230515 | 0.58 | 0.58 | -0.32 | 0.29 | 1.14 | -0.28 | 0.25 | 1.966 | 8.7 | 9.3 | 7.8 | 8.7 | 119.7 | 120.1 | 0.29 | 0.35 | -0.11 | 0.01 | 0.75 | -0.14 | -0.05 | 1.985 | 11.7 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | OJ, ON, EN | MatriXX_3 | 2023-05-15 | 1.9774 | 0.99423 | 1.9991 | 0.99295 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230516 | 0.67 | 0.62 | -0.36 | 0.13 | 1.18 | -0.34 | 0.13 | 1.967 | 8.6 | 9.4 | 8.1 | 8.5 | 119.8 | 120.2 | 0.4 | 0.36 | -0.16 | 0.01 | 0.76 | -0.17 | -0.04 | 1.992 | 11.8 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | EN,AJ,AR | MatriXX_3 | 2023-05-16 | 1.9774 | 0.99474 | 1.9991 | 0.99645 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230517 | 0.61 | 0.61 | -0.37 | 0.26 | 1.14 | -0.34 | 0.21 | 1.977 | 8.8 | 9.2 | 7.9 | 8.6 | 119.7 | 120.1 | 0.31 | 0.35 | -0.14 | 0.01 | 0.73 | -0.15 | -0.04 | 1.993 | 11.6 | 11.6 | 11.9 | 11.5 | 119.4 | 119.8 | AE, OJ, EF | MatriXX_3 | 2023-05-17 | 1.9774 | 0.9998 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230519 | 0.64 | 0.57 | -0.32 | 0.23 | 1.04 | -0.26 | 0.2 | 1.979 | 8.7 | 9.3 | 7.7 | 8.7 | 119.7 | 120.1 | 0.32 | 0.43 | -0.11 | -0.02 | 0.69 | -0.15 | -0.06 | 2.001 | 11.5 | 11.5 | 11.9 | 11.5 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-05-19 | 1.9774 | 1.0008 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230522 | 0.56 | 0.6 | -0.3 | 0.12 | 1.16 | -0.3 | 0.14 | 1.969 | 8.9 | 9.1 | 8.3 | 8.4 | 119.8 | 120.1 | 0.31 | 0.39 | -0.13 | -0.02 | 0.7 | -0.14 | -0.05 | 1.989 | 11.5 | 11.5 | 11.8 | 11.5 | 119.4 | 119.8 | AE | MatriXX_3 | 2023-05-22 | 1.9774 | 0.99575 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230523 | 0.71 | 0.63 | -0.36 | 0.24 | 1.1 | -0.3 | 0.22 | 1.981 | 9.2 | 9 | 7.6 | 9.1 | 119.7 | 120 | 0.32 | 0.31 | -0.12 | 0.02 | 0.69 | -0.11 | -0.06 | 1.99 | 11.6 | 11.6 | 11.8 | 11.6 | 119.4 | 119.8 | GF, ON | MatriXX_3 | 2023-05-23 | 1.9774 | 1.0018 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230524 | 0.54 | 0.63 | -0.28 | 0.29 | 1.03 | -0.26 | 0.24 | 1.968 | 8.7 | 9.2 | 7.8 | 8.7 | 119.8 | 120.1 | 0.36 | 0.34 | -0.15 | 0.04 | 0.69 | -0.15 | -0.03 | 1.988 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | AE, SM | MatriXX_3 | 2023-05-24 | 1.9774 | 0.99525 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230525 | 0.6 | 0.58 | -0.32 | 0.12 | 1.07 | -0.3 | 0.12 | 1.969 | 8.6 | 9.4 | 8.1 | 8.5 | 119.7 | 120.1 | 0.32 | 0.37 | -0.17 | -0.04 | 0.72 | -0.17 | -0.07 | 1.989 | 11.8 | 11.7 | 11.9 | 11.6 | 119.4 | 119.8 | AR, SM | MatriXX_3 | 2023-05-25 | 1.9774 | 0.99575 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230526 | 0.74 | 0.58 | -0.32 | 0.21 | 1.06 | -0.28 | 0.17 | 1.969 | 9 | 9.1 | 7.9 | 8.7 | 119.7 | 120.1 | 0.29 | 0.33 | -0.12 | 0.01 | 0.78 | -0.16 | -0.05 | 1.989 | 11.5 | 11.5 | 11.8 | 11.5 | 119.4 | 119.8 | OJ, EN | MatriXX_3 | 2023-05-26 | 1.9774 | 0.99575 | 1.9991 | 0.99495 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230529 | 0.52 | 0.63 | -0.26 | 0.28 | 1.04 | -0.26 | 0.25 | 1.968 | 9 | 9.1 | 8 | 8.7 | 119.7 | 120.1 | 0.37 | 0.33 | -0.16 | 0.01 | 0.75 | -0.17 | -0.04 | 1.988 | 11.6 | 11.5 | 11.8 | 11.5 | 119.4 | 119.8 | OJ | MatriXX_3 | 2023-05-29 | 1.9774 | 0.99525 | 1.9991 | 0.99445 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230530 | 0.74 | 0.58 | -0.36 | 0.19 | 1.12 | -0.28 | 0.16 | 1.975 | 8.7 | 9.3 | 8 | 8.7 | 119.7 | 120.1 | 0.37 | 0.35 | -0.17 | 0.03 | 0.72 | -0.16 | -0.03 | 1.993 | 11.7 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | AE | MatriXX_3 | 2023-05-30 | 1.9774 | 0.99879 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230531 | 0.66 | 0.65 | -0.34 | 0.27 | 1.12 | -0.3 | 0.22 | 1.97 | 8.9 | 9.1 | 7.8 | 8.8 | 119.8 | 120.1 | 0.39 | 0.43 | -0.17 | -0.02 | 0.76 | -0.17 | -0.07 | 1.99 | 11.6 | 11.6 | 12.1 | 11.7 | 119.4 | 119.7 | JR | MatriXX_3 | 2023-05-31 | 1.9774 | 0.99626 | 1.9991 | 0.99545 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230601 | 0.43 | 0.58 | -0.26 | 0.15 | 1.09 | -0.26 | 0.17 | 1.977 | 8.7 | 9.2 | 8.1 | 8.5 | 119.8 | 120.1 | 0.26 | 0.38 | -0.1 | 0.01 | 0.66 | -0.1 | -0.05 | 1.996 | 11.7 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | EF | MatriXX_3 | 2023-06-01 | 1.9774 | 0.9998 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230602 | 0.41 | 0.58 | -0.25 | 0.15 | 1.04 | -0.25 | 0.17 | 1.966 | 9 | 8.9 | 8.2 | 8.4 | 119.8 | 120.1 | 0.28 | 0.41 | -0.09 | -0.01 | 0.74 | -0.12 | -0.06 | 1.987 | 11.6 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | AJ | MatriXX_3 | 2023-06-02 | 1.9774 | 0.99423 | 1.9991 | 0.99395 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230605 | 0.68 | 0.61 | -0.35 | 0.24 | 1.12 | -0.31 | 0.22 | 1.977 | 8.7 | 9.3 | 8.1 | 8.6 | 119.7 | 120.1 | 0.41 | 0.41 | -0.14 | 0.02 | 0.77 | -0.16 | -0.02 | 1.997 | 11.6 | 11.6 | 11.8 | 11.5 | 119.3 | 119.7 | EF | MatriXX_3 | 2023-06-05 | 1.9774 | 0.9998 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230607 | 0.5 | 0.57 | -0.27 | 0.14 | 1.07 | -0.26 | 0.14 | 1.981 | 8.8 | 9.1 | 8.2 | 8.4 | 119.8 | 120.1 | 0.37 | 0.36 | -0.14 | -0.01 | 0.71 | -0.14 | -0.07 | 2.006 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-06-07 | 1.9774 | 1.0018 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230608 | 0.74 | 0.56 | -0.33 | 0.21 | 1.15 | -0.3 | 0.18 | 1.985 | 9 | 9 | 8 | 8.6 | 119.7 | 120.1 | 0.35 | 0.38 | -0.17 | -0.03 | 0.72 | -0.17 | -0.06 | 2.007 | 11.5 | 11.6 | 11.8 | 11.5 | 119.4 | 119.7 | EN | MatriXX_3 | 2023-06-08 | 1.9774 | 1.0038 | 1.9991 | 1.004 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230609 | 0.72 | 0.54 | -0.37 | 0.23 | 1.02 | -0.3 | 0.21 | 1.986 | 8.8 | 9.2 | 7.9 | 8.6 | 119.7 | 120.1 | 0.29 | 0.36 | -0.11 | 0.01 | 0.66 | -0.11 | -0.02 | 2.006 | 11.3 | 11.4 | 11.5 | 11.4 | 119.4 | 119.8 | AR | MatriXX_3 | 2023-06-09 | 1.9774 | 1.0043 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230612 | 0.49 | 0.6 | -0.24 | 0.28 | 1.08 | -0.28 | 0.22 | 1.976 | 8.9 | 9 | 7.9 | 8.6 | 119.8 | 120.1 | 0.28 | 0.33 | -0.09 | 0.05 | 0.68 | -0.13 | -0.05 | 1.996 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | AC | MatriXX_3 | 2023-06-12 | 1.9774 | 0.99929 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230613 | 0.59 | 0.63 | -0.31 | 0.28 | 1.13 | -0.32 | 0.24 | 1.978 | 8.7 | 9.3 | 8 | 8.6 | 119.7 | 120.1 | 0.25 | 0.37 | -0.09 | -0.02 | 0.64 | -0.11 | -0.09 | 1.997 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | OJ | MatriXX_3 | 2023-06-13 | 1.9774 | 1.0003 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230614 | 0.72 | 0.59 | -0.37 | 0.26 | 1.12 | -0.3 | 0.22 | 1.981 | 9.1 | 9 | 7.7 | 8.8 | 119.7 | 120.1 | 0.39 | 0.38 | -0.18 | 0.04 | 0.77 | -0.19 | -0.02 | 2.004 | 11.5 | 11.5 | 11.4 | 11.3 | 119.4 | 119.8 | EF | MatriXX_3 | 2023-06-14 | 1.9774 | 1.0018 | 1.9991 | 1.0025 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230615 | 0.68 | 0.63 | -0.39 | 0.31 | 1.15 | -0.34 | 0.24 | 1.976 | 9 | 9 | 7.8 | 8.7 | 119.7 | 120.1 | 0.33 | 0.34 | -0.15 | 0.02 | 0.73 | -0.14 | -0.04 | 1.998 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EF | MatriXX_3 | 2023-06-15 | 1.9774 | 0.99929 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230616 | 0.44 | 0.6 | -0.25 | 0.18 | 1.09 | -0.28 | 0.17 | 1.975 | 8.8 | 9.1 | 8.4 | 8.3 | 119.8 | 120.1 | 0.34 | 0.35 | -0.15 | 0.01 | 0.71 | -0.13 | -0.04 | 1.998 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-06-16 | 1.9774 | 0.99879 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230619 | 0.5 | 0.64 | -0.23 | 0.26 | 1 | -0.27 | 0.22 | 2.005 | 8.9 | 9 | 8 | 8.6 | 119.7 | 120.1 | 0.35 | 0.35 | -0.14 | 0.04 | 0.7 | -0.15 | -0.03 | 2.025 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | EN, After LLRF calibration/ML | MatriXX_3 | 2023-06-19 | 1.9774 | 1.014 | 1.9991 | 1.013 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230620 | 0.62 | 0.75 | -0.25 | 0.22 | 1.08 | -0.26 | 0.18 | 1.976 | 9.5 | 8.4 | 8 | 8.7 | 119.7 | 120.1 | 0.27 | 0.41 | -0.11 | -0.01 | 0.73 | -0.13 | -0.07 | 1.995 | 11.7 | 12 | 11.7 | 11.5 | 119.4 | 119.8 | GK | MatriXX_3 | 2023-06-20 | 1.9774 | 0.99929 | 1.9991 | 0.99795 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230621 | 0.49 | 0.58 | -0.22 | 0.27 | 1.02 | -0.26 | 0.2 | 1.98 | 8.7 | 9.2 | 7.9 | 8.6 | 119.8 | 120.1 | 0.26 | 0.37 | -0.09 | 0.01 | 0.67 | -0.1 | -0.04 | 2 | 11.5 | 11.5 | 11.4 | 11.4 | 119.4 | 119.8 | OJ | MatriXX_3 | 2023-06-21 | 1.9774 | 1.0013 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230622 | 0.47 | 0.63 | -0.28 | 0.24 | 1.08 | -0.29 | 0.14 | 1.981 | 9 | 8.9 | 8 | 8.4 | 119.8 | 120.1 | 0.36 | 0.31 | -0.17 | 0.02 | 0.65 | -0.15 | -0.04 | 2 | 11.5 | 11.5 | 11.3 | 11.3 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-06-22 | 1.9774 | 1.0018 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230626 | 0.57 | 0.49 | -0.28 | 0.1 | 0.96 | -0.3 | 0.08 | 1.985 | 8.4 | 9.4 | 8 | 8.5 | 119.7 | 120.1 | 0.27 | 0.38 | -0.09 | -0.02 | 0.68 | -0.11 | -0.04 | 2.01 | 11.7 | 11.7 | 11.7 | 11.5 | 119.3 | 119.8 | AR | MatriXX_3 | 2023-06-26 | 1.9774 | 1.0038 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230627 | 0.73 | 0.61 | -0.38 | 0.25 | 1.01 | -0.33 | 0.21 | 1.983 | 8.5 | 9.3 | 7.8 | 8.6 | 119.7 | 120.1 | 0.29 | 0.42 | -0.1 | -0.04 | 0.68 | -0.11 | -0.08 | 2.005 | 11.6 | 11.6 | 11.6 | 11.5 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-06-27 | 1.9774 | 1.0028 | 1.9991 | 1.003 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230628 | 0.69 | 0.58 | -0.34 | 0.23 | 1.08 | -0.31 | 0.19 | 1.981 | 9.2 | 8.8 | 7.8 | 8.6 | 119.7 | 120.1 | 0.37 | 0.39 | -0.16 | 0.06 | 0.71 | -0.17 | 0 | 2.002 | 11.6 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | AR | MatriXX_3 | 2023-06-28 | 1.9774 | 1.0018 | 1.9991 | 1.0015 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230629 | 0.61 | 0.61 | -0.29 | 0.15 | 0.96 | -0.29 | 0.14 | 1.983 | 8.5 | 9.4 | 8.1 | 8.4 | 119.7 | 120.1 | 0.27 | 0.34 | -0.1 | 0.02 | 0.63 | -0.13 | -0.05 | 2 | 11.7 | 11.6 | 11.4 | 11.3 | 119.4 | 119.9 | ON | MatriXX_3 | 2023-06-29 | 1.9774 | 1.0028 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230630 | 0.58 | 0.58 | -0.32 | 0.26 | 1.08 | -0.3 | 0.22 | 1.978 | 8.7 | 9.2 | 7.9 | 8.6 | 119.7 | 120.1 | 0.31 | 0.37 | -0.12 | 0 | 0.65 | -0.13 | -0.05 | 1.998 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-06-30 | 1.9774 | 1.0003 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230703 | 0.59 | 0.56 | -0.31 | 0.24 | 1.06 | -0.28 | 0.21 | 1.976 | 8.8 | 9.2 | 7.8 | 8.5 | 119.7 | 120.1 | 0.28 | 0.39 | -0.07 | -0.02 | 0.69 | -0.1 | -0.07 | 1.997 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | AR | MatriXX_3 | 2023-07-03 | 1.9774 | 0.99929 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230704 | 0.53 | 0.55 | -0.27 | 0.13 | 1.09 | -0.27 | 0.13 | 1.975 | 8.8 | 9.2 | 8 | 8.5 | 119.7 | 120.1 | 0.34 | 0.38 | -0.12 | -0.01 | 0.75 | -0.12 | -0.05 | 1.998 | 11.5 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-07-04 | 1.9774 | 0.99879 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230705 | 0.64 | 0.59 | -0.3 | 0.24 | 1.08 | -0.28 | 0.23 | 1.974 | 8.8 | 9.2 | 7.7 | 8.7 | 119.7 | 120.1 | 0.28 | 0.35 | -0.12 | -0.01 | 0.69 | -0.12 | -0.05 | 1.996 | 11.7 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | AR | MatriXX_3 | 2023-07-05 | 1.9774 | 0.99828 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230706 | 0.53 | 0.62 | -0.28 | 0.25 | 1.02 | -0.27 | 0.21 | 1.969 | 9 | 9 | 7.9 | 8.6 | 119.8 | 120.1 | 0.34 | 0.38 | -0.11 | 0.01 | 0.73 | -0.13 | -0.03 | 1.996 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | OJ | MatriXX_3 | 2023-07-06 | 1.9774 | 0.99575 | 1.9991 | 0.99845 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230707 | 0.46 | 0.66 | -0.28 | 0.14 | 1.03 | -0.28 | 0.15 | 1.973 | 8.7 | 9.2 | 8.1 | 8.4 | 119.8 | 120.1 | 0.35 | 0.33 | -0.15 | 0.02 | 0.72 | -0.17 | -0.03 | 1.993 | 11.5 | 11.6 | 11.7 | 11.4 | 119.4 | 119.8 | AJ, OJ | MatriXX_3 | 2023-07-07 | 1.9774 | 0.99777 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230710 | 0.44 | 0.56 | -0.28 | 0.17 | 1.11 | -0.29 | 0.17 | 1.974 | 8.7 | 9.2 | 8.1 | 8.5 | 119.8 | 120.1 | 0.31 | 0.42 | -0.14 | -0.01 | 0.74 | -0.17 | -0.04 | 1.994 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-07-10 | 1.9774 | 0.99828 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230711 | 0.48 | 0.59 | -0.29 | 0.13 | 1.05 | -0.28 | 0.13 | 1.973 | 8.6 | 9.2 | 8.4 | 8.1 | 119.8 | 120.2 | 0.26 | 0.41 | -0.11 | -0.06 | 0.71 | -0.12 | -0.07 | 1.993 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | OJ | MatriXX_3 | 2023-07-11 | 1.9774 | 0.99777 | 1.9991 | 0.99695 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230712 | 0.43 | 0.56 | -0.27 | 0.16 | 1.05 | -0.28 | 0.17 | 1.977 | 8.8 | 9.1 | 8 | 8.3 | 119.8 | 120.2 | 0.28 | 0.36 | -0.07 | -0.02 | 0.68 | -0.11 | -0.07 | 1.999 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.9 | ON | MatriXX_3 | 2023-07-12 | 1.9774 | 0.9998 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230713 | 0.53 | 0.62 | -0.3 | 0.2 | 1.05 | -0.31 | 0.18 | 1.975 | 8.4 | 9.4 | 8.1 | 8.3 | 119.7 | 120.2 | 0.28 | 0.38 | -0.06 | 0 | 0.7 | -0.08 | -0.04 | 1.998 | 11.7 | 11.6 | 11.5 | 11.4 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-07-13 | 1.9774 | 0.99879 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230714 | 0.51 | 0.53 | -0.3 | 0.13 | 1.05 | -0.28 | 0.15 | 1.991 | 8.7 | 9.2 | 7.9 | 8.4 | 119.7 | 120.2 | 0.36 | 0.39 | -0.16 | -0.06 | 0.73 | -0.13 | -0.06 | 2.015 | 11.3 | 11.4 | 11.4 | 11.3 | 119.4 | 119.9 | ON | MatriXX_3 | 2023-07-14 | 1.9774 | 1.0069 | 1.9991 | 1.008 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230717 | 0.6 | 0.61 | -0.32 | 0.12 | 0.99 | -0.29 | 0.13 | 1.992 | 8.4 | 9.5 | 8.2 | 8.2 | 119.8 | 120.2 | 0.38 | 0.39 | -0.14 | -0.02 | 0.71 | -0.13 | -0.01 | 2.01 | 12 | 11.8 | 11.6 | 11.4 | 119.4 | 119.9 | ON | MatriXX_3 | 2023-07-17 | 1.9774 | 1.0074 | 1.9991 | 1.0055 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230718 | 0.68 | 0.61 | -0.36 | 0.08 | 1.12 | -0.3 | 0.07 | 1.969 | 8.5 | 9.4 | 8.3 | 8.3 | 119.7 | 120.1 | 0.32 | 0.4 | -0.09 | -0.03 | 0.68 | -0.11 | -0.04 | 1.994 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.8 | AC | MatriXX_3 | 2023-07-18 | 1.9774 | 0.99575 | 1.9991 | 0.99745 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230719 | 0.47 | 0.72 | -0.15 | 0.28 | 1.05 | -0.22 | 0.24 | 1.98 | 9.4 | 8.4 | 7.7 | 8.6 | 119.7 | 120.1 | 0.36 | 0.42 | -0.15 | -0.05 | 0.74 | -0.14 | -0.05 | 2.003 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-07-19 | 1.9774 | 1.0013 | 1.9991 | 1.002 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230720 | 0.51 | 0.63 | -0.17 | 0.25 | 1.13 | -0.24 | 0.25 | 1.978 | 9.2 | 8.6 | 7.6 | 8.8 | 119.8 | 120.1 | 0.31 | 0.37 | -0.13 | -0.02 | 0.69 | -0.14 | -0.07 | 2 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-07-20 | 1.9774 | 1.0003 | 1.9991 | 1.0005 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230721 | 0.42 | 0.67 | -0.27 | 0.18 | 1.05 | -0.27 | 0.19 | 1.986 | 8.8 | 9.2 | 7.9 | 8.5 | 119.7 | 120.1 | 0.33 | 0.34 | -0.13 | 0.05 | 0.72 | -0.14 | -0.01 | 2.006 | 11.5 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | AR | MatriXX_3 | 2023-07-21 | 1.9774 | 1.0043 | 1.9991 | 1.0035 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230724 | 0.41 | 0.51 | -0.26 | 0.11 | 0.98 | -0.27 | 0.13 | 1.974 | 8.7 | 9.2 | 8.2 | 8.2 | 119.7 | 120.2 | 0.3 | 0.35 | -0.13 | 0.01 | 0.72 | -0.15 | -0.04 | 1.999 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | EN | MatriXX_3 | 2023-07-24 | 1.9774 | 0.99828 | 1.9991 | 0.99995 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230725 | 0.66 | 0.67 | -0.32 | 0.26 | 1.11 | -0.29 | 0.24 | 1.971 | 8.5 | 9.3 | 7.9 | 8.7 | 119.7 | 120.1 | 0.38 | 0.4 | -0.16 | -0.01 | 0.74 | -0.14 | -0.01 | 1.986 | 11.8 | 11.7 | 11.9 | 11.5 | 119.4 | 119.9 | EN | MatriXX_3 | 2023-07-25 | 1.9774 | 0.99676 | 1.9991 | 0.99345 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230726 | 0.44 | 0.58 | -0.26 | 0.2 | 1.13 | -0.29 | 0.16 | 1.976 | 9 | 8.9 | 8.2 | 8.2 | 119.8 | 120.2 | 0.28 | 0.4 | -0.07 | -0.03 | 0.73 | -0.11 | -0.06 | 2.001 | 11.5 | 11.5 | 11.7 | 11.4 | 119.4 | 119.9 | AR | MatriXX_3 | 2023-07-26 | 1.9774 | 0.99929 | 1.9991 | 1.001 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230727 | 0.48 | 0.62 | -0.28 | 0.17 | 1.14 | -0.25 | 0.18 | 1.973 | 8.8 | 9.2 | 8 | 8.4 | 119.8 | 120.2 | 0.36 | 0.38 | -0.15 | 0.02 | 0.73 | -0.14 | -0.02 | 1.997 | 11.6 | 11.6 | 11.9 | 11.5 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-07-27 | 1.9774 | 0.99777 | 1.9991 | 0.99895 | 1.01 | 0.99 | 1.02 | 0.98 | 1
20230728 | 0.44 | 0.63 | -0.24 | 0.21 | 1.04 | -0.27 | 0.17 | 1.977 | 8.6 | 9.3 | 8 | 8.3 | 119.8 | 120.2 | 0.2 | 0.35 | -0.04 | 0.01 | 0.63 | -0.05 | -0.06 | 1.998 | 11.8 | 11.7 | 11.8 | 11.5 | 119.4 | 119.8 | ON | MatriXX_3 | 2023-07-28 | 1.9774 | 0.9998 | 1.9991 | 0.99945 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 0
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1
[31 empty] | #VALUE! | 1.9774 | 0 | 1.9991 | 0 | 1.01 | 0.99 | 1.02 | 0.98 | 1

## Summary
GTR1 | [7 empty] | GTR2
| Reference | R15_SOBP10 | R31_SOBP10 | [5 empty] | Reference | R15_SOBP10 | R31_SOBP10
| Average for 2019 | 2.0026 | 2.0002 | [5 empty] | Average for 2019 | 1.9831 | 2.0103
| STD | 0.0055015 | 0.0066466 | [5 empty] | STD | 0.00635 | 0.010985
| RSTD | 0.2747% | 0.3323% | [5 empty] | RSTD | 0.3202% | 0.5465%
| Date/Plan ID | R15_SOBP10 | R31_SOBP10 |  | Date/Plan ID | R15_SOBP10 | R31_SOBP10 |  | Date/Plan ID | R15_SOBP10 | R31_SOBP10 |  | Date/Plan ID | R15_SOBP10 | R31_SOBP10
20190206 | 2019-02-06 | 2 | 1.999 | LS | 2019-02-06 | -0.1298% | -0.05999% | 20190220 | 2019-02-20 | 1.98 | 2.005 | LS | 2019-02-20 | -0.1563% | -0.2636%
20190220 | 2019-02-20 | 1.998 | 1.995 | LS | 2019-02-20 | -0.2297% | -0.26% | 20190313 | 2019-03-13 | 1.988 | 2.016 | LS | 2019-03-13 | 0.2471% | 0.2835%
20190305 | 2019-03-05 | 2.014 | 2.008 | LS | 2019-03-05 | 0.5693% | 0.39% | 20190409 | 2019-04-09 | 1.994 | 2.014 | LS | 2019-04-09 | 0.5496% | 0.1841%
20190312 | 2019-03-12 | 1.994 | 2.003 | LS | 2019-03-12 | -0.4294% | 0.14% | 20190424 | 2019-04-24 | 1.979 | 2 | LS | 2019-04-24 | -0.2067% | -0.5124%
20190317 | 2019-03-17 | 2.007 | 2.006 | LS | 2019-03-17 | 0.2197% | 0.29% | 20190508 | 2019-05-08 | 1.977 | 2.037 | LS | 2019-05-08 | -0.3076% | 1.328%
20190320 | 2019-03-20 | 2.006 | 1.995 | LS | 2019-03-20 | 0.1698% | -0.26% | 20190515 | 2019-05-15 | 1.982 | 2.005 | LS | 2019-05-15 | -0.05547% | -0.2636%
20190326 | 2019-03-26 | 2 | 1.997 | LS | 2019-03-26 | -0.1298% | -0.16% | 20190528 | 2019-05-28 | 1.986 | 2.005 | LS | 2019-05-28 | 0.1462% | -0.2636%
20190422 | 2019-04-22 | 2.001 | 1.993 | LS | 2019-04-22 | -0.0799% | -0.36% | 20190618 | 2019-06-18 | 1.987 | 2.013 | LS | 2019-06-18 | 0.1967% | 0.1343%
20190507 | 2019-05-07 | 2.003 | 1.994 | LS | 2019-05-07 | 0.01997% | -0.31% | 20190625 | 2019-06-25 | 1.986 | 2.009 | LS | 2019-06-25 | 0.1462% | -0.06467%
20190514 | 2019-05-14 | 2.003 | 2.012 | LS | 2019-05-14 | 0.01997% | 0.5899% | 20190702 | 2019-07-02 | 1.972 | 1.999 | LS | 2019-07-02 | -0.5597% | -0.5621%
20190523 | 2019-05-23 | 2.02 | 2.015 | LS | 2019-05-23 | 0.8689% | 0.7399% | 20190717 | 2019-07-17 | 1.982 | 2.005 | LS | 2019-07-17 | -0.05547% | -0.2636%
20190619 | 2019-06-19 | 1.999 | 1.995 | LS | 2019-06-19 | -0.1798% | -0.26% | 20190724 | 2019-07-24 | 1.983 | 2.006 | LS | 2019-07-24 | -0.005043% | -0.2139%
20190702 | 2019-07-02 | 1.994 | 1.995 | LS | 2019-07-02 | -0.4294% | -0.26% | 20190730 | 2019-07-30 | 1.981 | 2.01 | LS | 2019-07-30 | -0.1059% | -0.01492%
20190709 | 2019-07-09 | 1.998 | 1.999 | LS | 2019-07-09 | -0.2297% | -0.05999% | 20190809 | 2019-08-09 | 1.991 | 2.015 | LS | 2019-08-09 | 0.3984% | 0.2338%
20190723 | 2019-07-23 | 2.001 | 1.997 | LS | 2019-07-23 | -0.0799% | -0.16% | 20190819 | 2019-08-19 | 1.99 | 0 | LS | 2019-08-19 | 0.3479%
20190801 | 2019-08-01 | 2.004 | 2.003 | LS | 2019-08-01 | 0.06991% | 0.14% | 20190905 | 2019-09-05 | 1.989 | 2.01 | LS | 2019-09-05 | 0.2975% | -0.01492%
20190807 | 2019-08-07 | 2.019 | 0 | LS | 2019-08-07 | 0.8189% |  | 20190912 | 2019-09-12 | 1.991 | 0 | LS | 2019-09-12 | 0.3984%
20190814 | 2019-08-14 | 2.003 | 2.006 | LS | 2019-08-14 | 0.01997% | 0.29% | 20190920 | 2019-09-20 | 1.986 | 2.01 | LS | 2019-09-20 | 0.1462% | -0.01492%
20190820 | 2019-08-20 | 2.004 | 2.001 | LS | 2019-08-20 | 0.06991% | 0.04% | 20191002 | 2019-10-02 | 1.976 | 1.998 | LS | 2019-10-02 | -0.358% | -0.6118%
20190828 | 2019-08-28 | 2.004 | 2 | LS | 2019-08-28 | 0.06991% | -0.009999% | 20191008 | 2019-10-08 | 1.986 | 2.01 | LS | 2019-10-08 | 0.1462% | -0.01492%
20190911 | 2019-09-11 | 2 | 2.003 | LS | 2019-09-11 | -0.1298% | 0.14% | 20191014 | 2019-10-14 | 1.987 | 2.004 | LS | 2019-10-14 | 0.1967% | -0.3134%
20190918 | 2019-09-18 | 2.002 | 2.004 | LS | 2019-09-18 | -0.02996% | 0.19% | 20191017 | 2019-10-17 | 1.982 | 0 | LS | 2019-10-17 | -0.05547%
20191009 | 2019-10-09 | 2.005 | 2.002 | LS | 2019-10-09 | 0.1198% | 0.08999% | 20191023 | 2019-10-23 | 1.987 | 2.008 | LS | 2019-10-23 | 0.1967% | -0.1144%
20191013 | 2019-10-13 | 1.997 | 1.997 | LS | 2019-10-13 | -0.2796% | -0.16% | 20191024 | 2019-10-24 | 1.964 | 1.99 | LS | 2019-10-24 | -0.9631% | -1.01%
20191016 | 2019-10-16 | 1.993 | 1.986 | LS | 2019-10-16 | -0.4794% | -0.7099% | 20191031 | 2019-10-31 | 1.981 | 2.01 | LS | 2019-10-31 | -0.1059% | -0.01492%
20191030 | 2019-10-30 | 2.005 | 2.002 | LS | 2019-10-30 | 0.1198% | 0.08999% | 20191108 | 2019-11-08 | 1.977 | 2.001 | LS | 2019-11-08 | -0.3076% | -0.4626%
20191107 | 2019-11-07 | 2 | 1.997 | LS | 2019-11-07 | -0.1298% | -0.16% | 20191114 | 2019-11-14 | 1.959 | 1.99 | LS | 2019-11-14 | -1.215% | -1.01%
20191114 | 2019-11-14 | 1.999 | 1.997 | LS | 2019-11-14 | -0.1798% | -0.16% | 20191126 | 2019-11-26 | 1.992 | 2.016 | LS | 2019-11-26 | 0.4488% | 0.2835%
20191119 | 2019-11-19 | 2.007 | 2.006 | LS | 2019-11-19 | 0.2197% | 0.29% | 20191205 | 2019-12-05 | 1.983 | 2.007 | LS | 2019-12-05 | -0.005043% | -0.1642%
20191122 | 2019-11-22 | 2.004 | 2.006 | LS | 2019-11-22 | 0.06991% | 0.29% | 20191210 | 2019-12-10 | 1.966 | 1.995 | LS | 2019-12-10 | -0.8623% | -0.7611%
20191128 | 2019-11-28 | 1.993 | 1.99 | LS | 2019-11-28 | -0.4794% | -0.5099% | 20191220 | 2019-12-20 | 1.983 | 2.004 | LS | 2019-12-20 | -0.005043% | -0.3134%
20191204 | 2019-12-04 | 2.002 | 2 | LS | 2019-12-04 | -0.02996% | -0.009999% | 20191229 | 2019-12-29 | 1.978 | 1.997 | LS | 2019-12-29 | -0.2572% | -0.6616%
20191211 | 2019-12-11 | 1.987 | 1.987 | LS | 2019-12-11 | -0.779% | -0.6599% | 20200104 | 2020-01-04 | 1.983 | 2.009 | LS | 2020-01-04 | -0.005043% | -0.06467%
20191219 | 2019-12-19 | 2.004 | 2.003 | LS | 2019-12-19 | 0.06991% | 0.14% | 20200115 | 2020-01-15 | 1.978 | 2 | LS | 2020-01-15 | -0.2572% | -0.5124%
20191228 | 2019-12-28 | 2.008 | 2.009 | LS | 2019-12-28 | 0.2696% | 0.44% | 20200122 | 2020-01-22 | 1.975 | 1.998 | LS, after AI upgrade and iso change | 2020-01-22 | -0.4085% | -0.6118%
20191231 | 2019-12-31 | 2.008 | 2.009 | LS | 2019-12-31 | 0.2696% | 0.44% | 20200129 | 2020-01-29 | 1.982 | 2.007 | LS | 2020-01-29 | -0.05547% | -0.1642%
20200107 | 2020-01-07 | 2.004 | 1.998 | LS | 2020-01-07 | 0.06991% | -0.11% | 20200203 | 2020-02-03 | 1.991 | 2.015 | LS | 2020-02-03 | 0.3984% | 0.2338%
20200107 | 2020-01-07 | 1.999 | 1.991 | ML | 2020-01-07 | -0.1798% | -0.46% | 20200213 | 2020-02-13 | 1.986 | 2.007 | LS | 2020-02-13 | 0.1462% | -0.1642%
20200110 | 2020-01-10 | 1.995 | 1.991 | LS | 2020-01-10 | -0.3795% | -0.46% | 20200214 | 2020-02-14 | 1.984 | 2.005 | LS | 2020-02-14 | 0.04538% | -0.2636%
20200124 | 2020-01-24 | 2.009 | 2.001 | LS | 2020-01-24 | 0.3196% | 0.04% | 20200220 | 2020-02-20 | 1.984 | 2.005 | LS | 2020-02-20 | 0.04538% | -0.2636%
20200130 | 2020-01-30 | 2.011 | 2.002 | LS | 2020-01-30 | 0.4195% | 0.08999% | 20200227 | 2020-02-27 | 1.989 | 2.01 | LS | 2020-02-27 | 0.2975% | -0.01492%
20200206 | 2020-02-06 | 2.005 | 1.998 | LS | 2020-02-06 | 0.1198% | -0.11% | 20200311 | 2020-03-11 | 1.982 | 2.005 | LS | 2020-03-11 | -0.05547% | -0.2636%
20200213 | 2020-02-13 | 2.003 | 1.995 | LS | 2020-02-13 | 0.01997% | -0.26% | 20200318 | 2020-03-18 | 1.989 | 2.014 | LS | 2020-03-18 | 0.2975% | 0.1841%
20200228 | 2020-02-28 | 2.01 | 2 | LS | 2020-02-28 | 0.3695% | -0.009999% | 20200326 | 2020-03-26 | 1.984 | 2.009 | LS | 2020-03-26 | 0.04538% | -0.06467%
20200303 | 2020-03-03 | 2.009 | 1.998 | LS | 2020-03-03 | 0.3196% | -0.11% | 20200401 | 2020-04-01 | 1.985 | 2.011 | LS | 2020-04-01 | 0.09581% | 0.03482%
20200306 | 2020-03-06 | 2.001 | 1.992 | LS | 2020-03-06 | -0.0799% | -0.41% | 20200408 | 2020-04-08 | 1.983 | 2.002 | LS | 2020-04-08 | -0.005043% | -0.4129%
20200312 | 2020-03-12 | 1.999 | 1.996 | LS | 2020-03-12 | -0.1798% | -0.21% | 20200417 | 2020-04-17 | 1.981 | 2.004 | LS | 2020-04-17 | -0.1059% | -0.3134%
20200319 | 2020-03-19 | 2.004 | 2.002 | LS | 2020-03-19 | 0.06991% | 0.08999% | 20200429 | 2020-04-29 | 1.971 | 1.997 | LS | 2020-04-29 | -0.6102% | -0.6616%
20200325 | 2020-03-25 | 2.003 | 2 | LS | 2020-03-25 | 0.01997% | -0.009999% | 20200506 | 2020-05-06 | 1.984 | 2.001 | LS | 2020-05-06 | 0.04538% | -0.4626%
20200402 | 2020-04-02 | 2.002 | 1.999 | LS | 2020-04-02 | -0.02996% | -0.05999% | 20200513 | 2020-05-13 | 1.982 | 2.008 | GB | 2020-05-13 | -0.05547% | -0.1144%
20200409 | 2020-04-09 | 1.997 | 1.987 | LS | 2020-04-09 | -0.2796% | -0.6599% | 20200519 | 2020-05-19 | 1.977 | 2 | AA | 2020-05-19 | -0.3076% | -0.5124%
20200416 | 2020-04-16 | 1.996 | 1.992 | LS | 2020-04-16 | -0.3296% | -0.41% | 20200525 | 2020-05-25 | 1.987 | 2.006 | LS | 2020-05-25 | 0.1967% | -0.2139%
20200423 | 2020-04-23 | 1.997 | 1.992 | LS | 2020-04-23 | -0.2796% | -0.41% | 20200604 | 2020-06-04 | 1.985 | 2.006 | AA | 2020-06-04 | 0.09581% | -0.2139%
20200507 | 2020-05-07 | 2.004 | 1.994 | LS | 2020-05-07 | 0.06991% | -0.31% | 20200611 | 2020-06-11 | 1.986 | 2.011 | LS | 2020-06-11 | 0.1462% | 0.03482%
20200515 | 2020-05-15 | 1.996 | 1.992 | AA | 2020-05-15 | -0.3296% | -0.41% | 20200616 | 2020-06-16 | 1.98 | 2.007 | AA | 2020-06-16 | -0.1563% | -0.1642%
20200528 | 2020-05-28 | 2.002 | 1.996 | LS | 2020-05-28 | -0.02996% | -0.21% | 20200618 | 2020-06-18 | 1.983 | 2.008 | AA | 2020-06-18 | -0.005043% | -0.1144%
20200603 | 2020-06-03 | 2.003 | 2.002 | LS | 2020-06-03 | 0.01997% | 0.08999% | 20200625 | 2020-06-25 | 1.983 | 2.009 | LS | 2020-06-25 | -0.005043% | -0.06467%
20200610 | 2020-06-10 | 2.003 | 1.993 | AA | 2020-06-10 | 0.01997% | -0.36% | 20200701 | 2020-07-01 | 1.974 | 2.002 | AA | 2020-07-01 | -0.4589% | -0.4129%
20200624 | 2020-06-24 | 2 | 1.995 | LS after downtime (1) | 2020-06-24 | -0.1298% | -0.26% | 20200706 | 2020-07-06 | 1.982 | 2.006 | LS | 2020-07-06 | -0.05547% | -0.2139%
20200624 | 2020-06-24 | 2.011 | 2.014 | LS after downtime (repetition 1) | 2020-06-24 | 0.4195% | 0.6899% | 20200707 | 2020-07-07 | 2.003 | 0 | LS | 2020-07-07 | 1.003%
20200624 | 2020-06-24 | 0 | 2.013 | LS after downtime (repetition 2) | 2020-06-24 |  | 0.6399% | 20200709 | 2020-07-09 | 1.982 | 2.008 | LS | 2020-07-09 | -0.05547% | -0.1144%
20200624 | 2020-06-24 | 0 | 2.013 | LS after downtime (repetition 3, updated DM) | 2020-06-24 |  | 0.6399% | 20200728 | 2020-07-28 | 1.98 | 2.007 | CVD | 2020-07-28 | -0.1563% | -0.1642%
20200624 | 2020-06-24 | 0 | 2.019 | CVD repetition in the evening | 2020-06-24 |  | 0.9399% | 20200806 | 2020-08-06 | 1.978 | 2.002 | LS | 2020-08-06 | -0.2572% | -0.4129%
20200624 | 2020-06-24 | 0 | 2 | LS | 2020-06-24 |  | -0.009999% | 20200813 | 2020-08-13 | 1.977 | 2 | LS | 2020-08-13 | -0.3076% | -0.5124%
20200702 | 2020-07-02 | 1.996 | 1.992 | AA | 2020-07-02 | -0.3296% | -0.41% | 20200820 | 2020-08-20 | 1.981 | 2.01 | LS | 2020-08-20 | -0.1059% | -0.01492%
20200705 | 2020-07-05 | 2 | 2.021 | LS | 2020-07-05 | -0.1298% | 1.04% | 20200827 | 2020-08-27 | 1.988 | 2.012 | AD | 2020-08-27 | 0.2471% | 0.08456%
20200707 | 2020-07-07 | 2.003 | 1.997 | LS | 2020-07-07 | 0.01997% | -0.16% | 20200904 | 2020-09-04 | 1.984 | 2.011 | AD | 2020-09-04 | 0.04538% | 0.03482%
20200708 | 2020-07-08 | 2 | 1.995 | AA | 2020-07-08 | -0.1298% | -0.26% | 20200911 | 2020-09-11 | 1.988 | 2.014 | AD | 2020-09-11 | 0.2471% | 0.1841%
20200715 | 2020-07-15 | 1.987 | 1.988 | AA | 2020-07-15 | -0.779% | -0.6099% | 20200917 | 2020-09-17 | 1.996 | 2.024 | AD | 2020-09-17 | 0.6505% | 0.6815%
20200727 | 2020-07-27 | 1.999 | 1.994 | AA | 2020-07-27 | -0.1798% | -0.31% | 20200924 | 2020-09-24 | 1.987 | 2.009 | AD | 2020-09-24 | 0.1967% | -0.06467%
20200805 | 2020-08-05 | 1.992 | 1.995 | LS | 2020-08-05 | -0.5293% | -0.26% | 20201001 | 2020-10-01 | 1.987 | 2.017 | AD | 2020-10-01 | 0.1967% | 0.3333%
20200812 | 2020-08-12 | 1.991 | 1.992 | LS | 2020-08-12 | -0.5792% | -0.41% | 20201008 | 2020-10-08 | 1.987 | 2.016 | AD | 2020-10-08 | 0.1967% | 0.2835%
20200819 | 2020-08-19 | 1.992 | 1.993 | AD | 2020-08-19 | -0.5293% | -0.36% | 20201011 | 2020-10-11 | 1.986 | 2.014 | AD | 2020-10-11 | 0.1462% | 0.1841%
20200826 | 2020-08-26 | 1.996 | 1.994 | AD | 2020-08-26 | -0.3296% | -0.31% | 20201022 | 2020-10-22 | 1.989 | 2.01 | ML | 2020-10-22 | 0.2975% | -0.01492%
20200903 | 2020-09-03 | 1.998 | 1.998 | AD | 2020-09-03 | -0.2297% | -0.11% | 20201029 | 2020-10-29 | 1.981 | 1.998 | ML | 2020-10-29 | -0.1059% | -0.6118%
20200909 | 2020-09-09 | 2.004 | 2.002 | AD | 2020-09-09 | 0.06991% | 0.08999% | 20201106 | 2020-11-06 | 1.977 | 2.002 | ML | 2020-11-06 | -0.3076% | -0.4129%
20200916 | 2020-09-16 | 2.006 | 2.004 | AD | 2020-09-16 | 0.1698% | 0.19% | 20201112 | 2020-11-12 | 1.984 | 2.006 | ML | 2020-11-12 | 0.04538% | -0.2139%
20200923 | 2020-09-23 | 2.004 | 2.002 | AD | 2020-09-23 | 0.06991% | 0.08999% | 20201118 | 2020-11-18 | 1.988 | 2.009 | ML | 2020-11-18 | 0.2471% | -0.06467%
20200930 | 2020-09-30 | 2.005 | 2 | AD | 2020-09-30 | 0.1198% | -0.009999% | 20201126 | 2020-11-26 | 1.967 | 1.997 | ML, Matrix_1 | 2020-11-26 | -0.8119% | -0.6616%
20201007 | 2020-10-07 | 1.999 | 1.997 | AD | 2020-10-07 | -0.1798% | -0.16% | 20201204 | 2020-12-04 | 1.968 | 1.993 | ML, Matrix_1 | 2020-12-04 | -0.7614% | -0.8606%
20201012 | 2020-10-12 | 1.998 | 1.992 | ML | 2020-10-12 | -0.2297% | -0.41% | 20201210 | 2020-12-10 | 1.978 | 2.007 | ML, Matrix_1 | 2020-12-10 | -0.2572% | -0.1642%
20201021 | 2020-10-21 | 2 | 1.997 | ML | 2020-10-21 | -0.1298% | -0.16% | 20201217 | 2020-12-17 | 1.978 | 2 | ML, Matrix_1 | 2020-12-17 | -0.2572% | -0.5124%
20201030 | 2020-10-30 | 1.999 | 1.994 | ML | 2020-10-30 | -0.1798% | -0.31% | 20201223 | 2020-12-23 | 1.971 | 1.991 | ML, Matrix_1 | 2020-12-23 | -0.6102% | -0.9601%
20201105 | 2020-11-05 | 1.998 | 1.998 | ML | 2020-11-05 | -0.2297% | -0.11% | 20210103 | 2021-01-03 | 1.976 | 1.999 | ML, Matrix_1 | 2021-01-03 | -0.358% | -0.5621%
20201111 | 2020-11-11 | 2.003 | 1.993 | ML | 2020-11-11 | 0.01997% | -0.36% | 20210114 | 2021-01-14 | 1.966 | 1.998 | ML, Matrix_1 | 2021-01-14 | -0.8623% | -0.6118%
20201119 | 2020-11-19 | 2.004 | 2 | ML, Matrix_1 | 2020-11-19 | 0.06991% | -0.009999% | 20210121 | 2021-01-21 | 1.973 | 1.995 | ML, Matrix_1 | 2021-01-21 | -0.5093% | -0.7611%
20201127 | 2020-11-27 | 2 | 1.996 | ML, Matrix_1 | 2020-11-27 | -0.1298% | -0.21% | 20210128 | 2021-01-28 | 1.974 | 1.998 | ML, Matrix_1 | 2021-01-28 | -0.4589% | -0.6118%
20201202 | 2020-12-02 | 1.997 | 1.99 | ML, Matrix_1 | 2020-12-02 | -0.2796% | -0.5099% | 20210205 | 2021-02-05 | 1.971 | 1.997 | ML, Matrix_1 | 2021-02-05 | -0.6102% | -0.6616%
20201209 | 2020-12-09 | 1.998 | 1.992 | ML, Matrix_1 | 2020-12-09 | -0.2297% | -0.41% | 20210211 | 2021-02-11 | 1.97 | 1.99 | ML, Matrix_1, rot 270 deg | 2021-02-11 | -0.6606% | -1.01%
20201216 | 2020-12-16 | 1.999 | 1.99 | ML, Matrix_1 | 2020-12-16 | -0.1798% | -0.5099% | 20210215 | 2021-02-15 | 1.977 | 1.991 | ML, Matrix_1 | 2021-02-15 | -0.3076% | -0.9601%
20201227 | 2020-12-27 | 2 | 1.997 | ML, Matrix_1 | 2020-12-27 | -0.1298% | -0.16% | 20210226 | 2021-02-26 | 1.974 | 1.996 | ML, Matrix_1 | 2021-02-26 | -0.4589% | -0.7113%
20201230 | 2020-12-30 | 1.994 | 1.988 | ML, Matrix_1 | 2020-12-30 | -0.4294% | -0.6099% | 20210304 | 2021-03-04 | 1.978 | 2 | ML, Matrix_1 | 2021-03-04 | -0.2572% | -0.5124%
20210107 | 2021-01-07 | 1.999 | 1.995 | ML, Matrix_1 | 2021-01-07 | -0.1798% | -0.26% | 20210306 | 2021-03-06 | 1.964 | 1.978 | ML, Matrix_1 | 2021-03-06 | -0.9631% | -1.607%
20210113 | 2021-01-13 | 1.986 | 1.979 | LL, Matrix_1 | 2021-01-13 | -0.8289% | -1.06% | 20210311 | 2021-03-11 | 1.984 | 2.007 | ML, Matrix_1 | 2021-03-11 | 0.04538% | -0.1642%
20210120 | 2021-01-20 | 1.987 | 1.984 | ML, Matrix_1 | 2021-01-20 | -0.779% | -0.8099% | 20210323 | 2021-03-23 | 1.974 | 1.992 | ML, Matrix_1 | 2021-03-23 | -0.4589% | -0.9103%
20210128 | 2021-01-28 | 1.997 | 1.989 | ML, Matrix_1 | 2021-01-28 | -0.2796% | -0.5599% | 20210323 | 2021-03-23 | 1.986 | 2.009 | ML, Matrix_2 | 2021-03-23 | 0.1462% | -0.06467%
20210129 | 2021-01-29 | 1.993 | 1.988 | ML, Matrix_1 | 2021-01-29 | -0.4794% | -0.6099% | 20210401 | 2021-04-01 | 1.972 | 1.995 | ML, MatriXX_1 | 2021-04-01 | -0.5597% | -0.7611%
20210203 | 2021-02-03 | 1.993 | 1.988 | ML, Matrix_1 | 2021-02-03 | -0.4794% | -0.6099% | 20210409 | 2021-04-09 | 1.978 | 1.999 | ML, MatriXX_2 | 2021-04-09 | -0.2572% | -0.5621%
20210211 | 2021-02-11 | 1.996 | 1.981 | ML, Matrix_1 | 2021-02-11 | -0.3296% | -0.9599% | 20210415 | 2021-04-15 | 1.981 | 2.006 | ML, MatriXX_1 | 2021-04-15 | -0.1059% | -0.2139%
20210215 | 2021-02-15 | 1.987 | 1.978 | ML, Matrix_1 | 2021-02-15 | -0.779% | -1.11% | 20210421 | 2021-04-21 | 1.974 | 1.998 | ML, MatriXX_2 | 2021-04-21 | -0.4589% | -0.6118%
20210223 | 2021-02-23 | 2.012 | 2.007 | ML, MatriXX_2 | 2021-02-23 | 0.4694% | 0.34% | 20210430 | 2021-04-30 | 1.993 | 2.014 | ML, MatriXX_1 | 2021-04-30 | 0.4992% | 0.1841%
20210224 | 2021-02-24 | 2.006 | 1.994 | ML, MatriXX_1 | 2021-02-24 | 0.1698% | -0.31% | 20210430 | 2021-04-30 | 1.983 | 2.003 | ML, MatriXX_1, run 1 | 2021-04-30 | -0.005043% | -0.3631%
20210303 | 2021-03-03 | 1.999 | 1.989 | ML, MatriXX_1 | 2021-03-03 | -0.1798% | -0.5599% | 20210430 | 2021-04-30 | 1.98 | 2.003 | ML, MatriXX_1, run 2 | 2021-04-30 | -0.1563% | -0.3631%
20210310 | 2021-03-10 | 1.992 | 1.987 | ML, MatriXX_1 | 2021-03-10 | -0.5293% | -0.6599% | 20210430 | 2021-04-30 | 1.982 | 2.004 | ML, MatriXX_1, run 3 | 2021-04-30 | -0.05547% | -0.3134%
20210324 | 2021-03-24 | 1.998 | 1.992 | ML, MatriXX_2 | 2021-03-24 | -0.2297% | -0.41% | 20210505 | 2021-05-05 | 1.972 | 1.998 | ML, MatriXX_1 | 2021-05-05 | -0.5597% | -0.6118%
20210325 | 2021-03-25 | 1.998 | 1.99 | ML, MatriXX_1 | 2021-03-25 | -0.2297% | -0.5099% | 20210509 | 2021-05-09 | 1.963 | 1.975 | ML, Matrixx 1 | 2021-05-09 | -1.014% | -1.756%
20210331 | 2021-03-31 | 2.002 | 1.995 | ML | 2021-03-31 | -0.02996% | -0.26% | 20210521 | 2021-05-21 | 1.975 | 1.992 | ML, Matrixx 1 | 2021-05-21 | -0.4085% | -0.9103%
20210408 | 2021-04-08 | 1.999 | 1.992 | ML, MatriXX_2 | 2021-04-08 | -0.1798% | -0.41% | 20210604 | 2021-06-04 | 1.995 | 2.019 | ML, Matrixx 2 | 2021-06-04 | 0.6001% | 0.4328%
20210414 | 2021-04-14 | 1.997 | 1.988 | ML, Matrixx_1 | 2021-04-14 | -0.2796% | -0.6099% | 20210607 | 2021-06-07 | 1.975 | 2 | ML, Matrixx_1 | 2021-06-07 | -0.4085% | -0.5124%
20210421 | 2021-04-21 | 2.003 | 1.994 | ML, Matrixx_2 | 2021-04-21 | 0.01997% | -0.31% | 20210609 | 2021-06-09 | 1.967 | 1.993 | CVD, Mx1 | 2021-06-09 | -0.8119% | -0.8606%
20210428 | 2021-04-28 | 2.006 | 2.006 | ML, MatriXX_1 | 2021-04-28 | 0.1698% | 0.29% | 20210610 | 2021-06-10 | 1.976 | 2 | CVD, Mx1 | 2021-06-10 | -0.358% | -0.5124%
20210430 | 2021-04-30 | 2.004 | 1.993 | ML, MatriXX_1, run 1 | 2021-04-30 | 0.06991% | -0.36% | 20210611 | 2021-06-11 | 1.974 | 2 | CVD, Mx1 | 2021-06-11 | -0.4589% | -0.5124%
20210430 | 2021-04-30 | 2.006 | 1.992 | ML, MatriXX_1, run 2 | 2021-04-30 | 0.1698% | -0.41% | 20210616 | 2021-06-16 | 1.962 | 1.977 | CVD, Mx1 | 2021-06-16 | -1.064% | -1.656%
20210430 | 2021-04-30 | 2.005 | 1.991 | ML, MatriXX_1, run 3 | 2021-04-30 | 0.1198% | -0.46% | 20210618 | 2021-06-18 | 1.977 | 2.001 | GB | 2021-06-18 | -0.3076% | -0.4626%
20210507 | 2021-05-07 | 1.992 | 1.983 | ML, MatriXX_1 | 2021-05-07 | -0.5293% | -0.8599% | 20210621 | 2021-06-21 | 1.981 | 2.005 | GB | 2021-06-21 | -0.1059% | -0.2636%
20210509 | 2021-05-09 | 1.995 | 1.973 | ML, Matrixx 1 | 2021-05-09 | -0.3795% | -1.36% | 20210623 | 2021-06-23 | 1.976 | 0 | GB | 2021-06-23 | -0.358%
20210512 | 2021-05-12 | 2.01 | 2.004 | ML, MatriXX_2 | 2021-05-12 | 0.3695% | 0.19% | 20210624 | 2021-06-24 | 1.981 | 1.995 | GB | 2021-06-24 | -0.1059% | -0.7611%
20210520 | 2021-05-20 | 1.997 | 1.987 | ML, MatriXX_1 | 2021-05-20 | -0.2796% | -0.6599% | 20210628 | 2021-06-28 | 1.972 | 1.996 | GB | 2021-06-28 | -0.5597% | -0.7113%
20210526 | 2021-05-26 | 2.006 | 1.993 | ML, MatriXX_2 | 2021-05-26 | 0.1698% | -0.36% | 20210630 | 2021-06-30 | 1.973 | 1.993 | GB | 2021-06-30 | -0.5093% | -0.8606%
20210603 | 2021-06-03 | 2.02 | 2.012 | ML, MatriXX_2 | 2021-06-03 | 0.8689% | 0.5899% | 20210702 | 2021-07-02 | 1.986 | 2.013 | GB | 2021-07-02 | 0.1462% | 0.1343%
20210607 | 2021-06-07 | 2.021 | 2.011 | cvd, Matrix 2 | 2021-06-07 | 0.9188% | 0.5399% | 20210705 | 2021-07-05 | 1.975 | 2.002 | ML | 2021-07-05 | -0.4085% | -0.4129%
20210608 | 2021-06-08 | 2.021 | 2.019 | GB, Matrix 2 | 2021-06-08 | 0.9188% | 0.9399% | 20210707 | 2021-07-07 | 1.976 | 1.996 | ML | 2021-07-07 | -0.358% | -0.7113%
20210609 | 2021-06-09 | 2.017 | 2.008 | CVD, Matrix 2 | 2021-06-09 | 0.7191% | 0.39% | 20210708 | 2021-07-08 | 1.985 | 2.012 | ML | 2021-07-08 | 0.09581% | 0.08456%
20210610 | 2021-06-10 | 2.023 | 2.008 | CVD, Matrix 2 | 2021-06-10 | 1.019% | 0.39% | 20210712 | 2021-07-12 | 1.971 | 1.995 | ML | 2021-07-12 | -0.6102% | -0.7611%
20210611 | 2021-06-11 | 2.016 | 2.005 | CVD, Matrix 2 | 2021-06-11 | 0.6691% | 0.24% | 20210714 | 2021-07-14 | 1.98 | 2.004 | ML | 2021-07-14 | -0.1563% | -0.3134%
20210615 | 2021-06-15 | 2.02 | 2.017 | CVD, Matrix 2 | 2021-06-15 | 0.8689% | 0.8399% | 20210715 | 2021-07-15 | 1.973 | 1.99 | CVD | 2021-07-15 | -0.5093% | -1.01%
20210616 | 2021-06-16 | 2.021 | 2.008 | CVD, Matrix 2 | 2021-06-16 | 0.9188% | 0.39% | 20210716 | 2021-07-16 | 1.974 | 1.998 | ML | 2021-07-16 | -0.4589% | -0.6118%
20210617 | 2021-06-17 | 2.023 | 2.015 | CVD, Matrix 2 | 2021-06-17 | 1.019% | 0.7399% | 20210720 | 2021-07-20 | 1.973 | 1.997 | ML | 2021-07-20 | -0.5093% | -0.6616%
20210618 | 2021-06-18 | 2.016 | 2.006 | GB | 2021-06-18 | 0.6691% | 0.29% | 20210721 | 2021-07-21 | 1.977 | 1.997 | ML | 2021-07-21 | -0.3076% | -0.6616%
20210622 | 2021-06-22 | 2.018 | 2.009 | GB | 2021-06-22 | 0.769% | 0.44% | 20210722 | 2021-07-22 | 1.971 | 1.986 | ML | 2021-07-22 | -0.6102% | -1.209%
20210623 | 2021-06-23 | 2.018 | 2.01 | GB | 2021-06-23 | 0.769% | 0.49% | 20210723 | 2021-07-23 | 1.977 | 1.989 | ML | 2021-07-23 | -0.3076% | -1.06%
20210624 | 2021-06-24 | 2.014 | 2.005 | GB | 2021-06-24 | 0.5693% | 0.24% | 20210726 | 2021-07-26 | 1.978 | 2.001 | ML | 2021-07-26 | -0.2572% | -0.4626%
20210629 | 2021-06-29 | 2.009 | 2.001 | GB | 2021-06-29 | 0.3196% | 0.04% | 20210728 | 2021-07-28 | 1.975 | 1.999 | ML | 2021-07-28 | -0.4085% | -0.5621%
20210630 | 2021-06-30 | 2.01 | 2.002 | GB | 2021-06-30 | 0.3695% | 0.08999% | 20210729 | 2021-07-29 | 1.974 | 1.994 | ML | 2021-07-29 | -0.4589% | -0.8108%
20210701 | 2021-07-01 | 2.019 | 2.012 | ML | 2021-07-01 | 0.8189% | 0.5899% | 20210730 | 2021-07-30 | 1.972 | 1.998 | ML | 2021-07-30 | -0.5597% | -0.6118%
20210706 | 2021-07-06 | 2.001 | 1.998 | ML | 2021-07-06 | -0.0799% | -0.11% | 20210802 | 2021-08-02 | 1.978 | 2.007 | ML | 2021-08-02 | -0.2572% | -0.1642%
20210707 | 2021-07-07 | 2.005 | 2.004 | ML | 2021-07-07 | 0.1198% | 0.19% | 20210804 | 2021-08-04 | 1.986 | 2.001 | ML | 2021-08-04 | 0.1462% | -0.4626%
20210708 | 2021-07-08 | 2.005 | 1.996 | ML | 2021-07-08 | 0.1198% | -0.21% | 20210811 | 2021-08-11 | 1.988 | 2.01 | CVD | 2021-08-11 | 0.2471% | -0.01492%
20210709 | 2021-07-09 | 2.008 | 2.002 | ML | 2021-07-09 | 0.2696% | 0.08999% | 20210812 | 2021-08-12 | 1.985 | 2.007 | ML | 2021-08-12 | 0.09581% | -0.1642%
20210714 | 2021-07-14 | 2.002 | 1.996 | ML | 2021-07-14 | -0.02996% | -0.21% | 20210813 | 2021-08-13 | 1.982 | 2.001 | ML | 2021-08-13 | -0.05547% | -0.4626%
20210715 | 2021-07-15 | 2.016 | 2.004 | CVD | 2021-07-15 | 0.6691% | 0.19% | 20210816 | 2021-08-16 | 1.975 | 1.999 | ML | 2021-08-16 | -0.4085% | -0.5621%
20210716 | 2021-07-16 | 1.998 | 1.991 | ML | 2021-07-16 | -0.2297% | -0.46% | 20210818 | 2021-08-18 | 1.983 | 2.005 | CVD | 2021-08-18 | -0.005043% | -0.2636%
20210719 | 2021-07-19 | 1.998 | 1.987 | ML | 2021-07-19 | -0.2297% | -0.6599% | 20210820 | 2021-08-20 | 1.971 | 1.983 | ML | 2021-08-20 | -0.6102% | -1.358%
20210721 | 2021-07-21 | 2.001 | 1.992 | ML | 2021-07-21 | -0.0799% | -0.41% | 20210820 | 2021-08-20 | 1.971 | 1.982 | ML/drop not confirmed by measurements done with IC | 2021-08-20 | -0.6102% | -1.408%
20210722 | 2021-07-22 | 2.002 | 1.992 | ML | 2021-07-22 | -0.02996% | -0.41% | 20210823 | 2021-08-23 | 1.991 | 2.019 | ML | 2021-08-23 | 0.3984% | 0.4328%
20210723 | 2021-07-23 | 2.003 | 1.99 | ML | 2021-07-23 | 0.01997% | -0.5099% | 20210826 | 2021-08-26 | 1.99 | 2.005 | GB | 2021-08-26 | 0.3479% | -0.2636%
20210727 | 2021-07-27 | 1.998 | 1.992 | ML | 2021-07-27 | -0.2297% | -0.41% | 20210827 | 2021-08-27 | 1.979 | 2.009 | GB | 2021-08-27 | -0.2067% | -0.06467%
20210728 | 2021-07-28 | 2.007 | 2 | ML | 2021-07-28 | 0.2197% | -0.009999% | 20210830 | 2021-08-30 | 1.986 | 2.003 | GB | 2021-08-30 | 0.1462% | -0.3631%
20210729 | 2021-07-29 | 2.003 | 1.998 | ML | 2021-07-29 | 0.01997% | -0.11% | 20210901 | 2021-09-01 | 1.982 | 2.01 | GB | 2021-09-01 | -0.05547% | -0.01492%
20210730 | 2021-07-30 | 1.998 | 1.991 | ML | 2021-07-30 | -0.2297% | -0.46% | 20210902 | 2021-09-02 | 1.979 | 1.992 | GB | 2021-09-02 | -0.2067% | -0.9103%
20210803 | 2021-08-03 | 2.002 | 1.993 | ML | 2021-08-03 | -0.02996% | -0.36% | 20210903 | 2021-09-03 | 1.984 | 2.012 | GB | 2021-09-03 | 0.04538% | 0.08456%
20210804 | 2021-08-04 | 2.006 | 1.996 | ML | 2021-08-04 | 0.1698% | -0.21% | 20210906 | 2021-09-06 | 1.977 | 1.999 | ML, Matrixx 2 | 2021-09-06 | -0.3076% | -0.5621%
20210805 | 2021-08-05 | 2 | 1.994 | ML | 2021-08-05 | -0.1298% | -0.31% | 20210907 | 2021-09-07 | 1.98 | 2.001 | CVD | 2021-09-07 | -0.1563% | -0.4626%
20210806 | 2021-08-06 | 2.002 | 1.993 | ML | 2021-08-06 | -0.02996% | -0.36% | 20210909 | 2021-09-09 | 1.981 | 2.001 | GB | 2021-09-09 | -0.1059% | -0.4626%
20210810 | 2021-08-10 | 2.005 | 1.994 | ML | 2021-08-10 | 0.1198% | -0.31% | 20210910 | 2021-09-10 | 1.974 | 1.998 | GB | 2021-09-10 | -0.4589% | -0.6118%
20210811 | 2021-08-11 | 2.004 | 1.991 | CVD | 2021-08-11 | 0.06991% | -0.46% | 20210914 | 2021-09-14 | 1.985 | 2.01 | ML | 2021-09-14 | 0.09581% | -0.01492%
20210812 | 2021-08-12 | 2.001 | 2 | ML | 2021-08-12 | -0.0799% | -0.009999% | 20210915 | 2021-09-15 | 1.971 | 1.992 | GB, Matrixx 1 | 2021-09-15 | -0.6102% | -0.9103%
20210813 | 2021-08-13 | 1.998 | 1.988 | ML | 2021-08-13 | -0.2297% | -0.6099% | 20210916 | 2021-09-16 | 1.986 | 2.011 | CVD | 2021-09-16 | 0.1462% | 0.03482%
20210817 | 2021-08-17 | 2 | 1.995 | ML | 2021-08-17 | -0.1298% | -0.26% | 20210917 | 2021-09-17 | 1.974 | 1.995 | GB | 2021-09-17 | -0.4589% | -0.7611%
20210818 | 2021-08-18 | 1.999 | 1.987 | CVD | 2021-08-18 | -0.1798% | -0.6599% | 20210920 | 2021-09-20 | 1.983 | 2.009 | ML | 2021-09-20 | -0.005043% | -0.06467%
20210819 | 2021-08-19 | 1.989 | 1.979 | cvd | 2021-08-19 | -0.6791% | -1.06% | 20210922 | 2021-09-22 | 1.98 | 2.001 | GB | 2021-09-22 | -0.1563% | -0.4626%
20210820 | 2021-08-20 | 2 | 1.982 | ML | 2021-08-20 | -0.1298% | -0.9099% | 20210923 | 2021-09-23 | 1.972 | 1.993 | ML | 2021-09-23 | -0.5597% | -0.8606%
20210824 | 2021-08-24 | 1.996 | 1.992 | CVD | 2021-08-24 | -0.3296% | -0.41% | 20210924 | 2021-09-24 | 1.972 | 1.992 | ML | 2021-09-24 | -0.5597% | -0.9103%
20210825 | 2021-08-25 | 2.015 | 2.006 | GB | 2021-08-25 | 0.6192% | 0.29% | 20210927 | 2021-09-27 | 1.977 | 2 | GB | 2021-09-27 | -0.3076% | -0.5124%
20210826 | 2021-08-26 | 1.999 | 1.991 | GB | 2021-08-26 | -0.1798% | -0.46% | 20211006 | 2021-10-06 | 1.976 | 1.997 | GB | 2021-10-06 | -0.358% | -0.6616%
20210831 | 2021-08-31 | 1.984 | 1.982 | CVD, Matrix 2 | 2021-08-31 | -0.9288% | -0.9099% | 20211007 | 2021-10-07 | 1.978 | 2.002 | GB | 2021-10-07 | -0.2572% | -0.4129%
20210901 | 2021-09-01 | 1.986 | 1.979 | GB | 2021-09-01 | -0.8289% | -1.06% | 20211008 | 2021-10-08 | 1.982 | 2.002 | GB | 2021-10-08 | -0.05547% | -0.4129%
20210902 | 2021-09-02 | 1.997 | 1.99 | GB | 2021-09-02 | -0.2796% | -0.5099% | 20211011 | 2021-10-11 | 1.985 | 2.001 | CVD | 2021-10-11 | 0.09581% | -0.4626%
20210903 | 2021-09-03 | 1.998 | 1.986 | GB | 2021-09-03 | -0.2297% | -0.7099% | 20211013 | 2021-10-13 | 1.979 | 1.998 | CVD | 2021-10-13 | -0.2067% | -0.6118%
20210906 | 2021-09-06 | 1.99 | 1.983 | CVD, Matrixx 1 | 2021-09-06 | -0.6292% | -0.8599% | 20211014 | 2021-10-14 | 1.979 | 2.001 | CVD | 2021-10-14 | -0.2067% | -0.4626%
20210908 | 2021-09-08 | 1.985 | 1.982 | ML | 2021-09-08 | -0.8789% | -0.9099% | 20211018 | 2021-10-18 | 1.989 | 2.007 | ML, Matrixx_1 | 2021-10-18 | 0.2975% | -0.1642%
20210909 | 2021-09-09 | 1.986 | 1.975 | GB | 2021-09-09 | -0.8289% | -1.26% | 20211020 | 2021-10-20 | 1.992 | 2.017 | ML | 2021-10-20 | 0.4488% | 0.3333%
20210910 | 2021-09-10 | 1.988 | 1.98 | GB | 2021-09-10 | -0.7291% | -1.01% | 20211021 | 2021-10-21 | 1.981 | 1.992 | GB | 2021-10-21 | -0.1059% | -0.9103%
20210914 | 2021-09-14 | 1.988 | 1.982 | ML | 2021-09-14 | -0.7291% | -0.9099% | 20211025 | 2021-10-25 | 1.991 | 2.009 | GB | 2021-10-25 | 0.3984% | -0.06467%
20210916 | 2021-09-16 | 1.993 | 1.986 | CVD | 2021-09-16 | -0.4794% | -0.7099% | 20211027 | 2021-10-27 | 1.987 | 2.007 | GB | 2021-10-27 | 0.1967% | -0.1642%
20210917 | 2021-09-17 | 1.996 | 1.985 | GB | 2021-09-17 | -0.3296% | -0.7599% | 20211028 | 2021-10-28 | 1.993 | 2.016 | GB | 2021-10-28 | 0.4992% | 0.2835%
20210921 | 2021-09-21 | 1.99 | 1.981 | CVD | 2021-09-21 | -0.6292% | -0.9599% | 20211029 | 2021-10-29 | 1.984 | 2.007 | KA | 2021-10-29 | 0.04538% | -0.1642%
20210922 | 2021-09-22 | 1.992 | 1.983 | GB | 2021-09-22 | -0.5293% | -0.8599% | 20211101 | 2021-11-01 | 1.987 | 2.008 | CVD | 2021-11-01 | 0.1967% | -0.1144%
20210924 | 2021-09-24 | 1.992 | 1.979 | ML | 2021-09-24 | -0.5293% | -1.06% | 20211103 | 2021-11-03 | 1.981 | 2.005 | GB | 2021-11-03 | -0.1059% | -0.2636%
20210928 | 2021-09-28 | 1.987 | 1.981 | GB | 2021-09-28 | -0.779% | -0.9599% | 20211104 | 2021-11-04 | 1.99 | 2.001 | ML | 2021-11-04 | 0.3479% | -0.4626%
20210929 | 2021-09-29 | 2.016 | 2.007 | GB | 2021-09-29 | 0.6691% | 0.34% | 20211108 | 2021-11-08 | 1.973 | 1.988 | GB/ wrong calib. Factor for Matrixx 1 | 2021-11-08 | -0.5093% | -1.109%
20210930 | 2021-09-30 | 2.013 | 2.005 | GB | 2021-09-30 | 0.5193% | 0.24% | 20211109 | 2021-11-09 | 1.986 | 2.005 | ML | 2021-11-09 | 0.1462% | -0.2636%
20211001 | 2021-10-01 | 2.017 | 2.011 | GB | 2021-10-01 | 0.7191% | 0.5399% | 20211112 | 2021-11-12 | 1.983 | 2.005 | ML | 2021-11-12 | -0.005043% | -0.2636%
20211004 | 2021-10-04 | 2.018 | 2.007 | GB | 2021-10-04 | 0.769% | 0.34% | 20211117 | 2021-11-17 | 1.98 | 2.003 | GB | 2021-11-17 | -0.1563% | -0.3631%
20211005 | 2021-10-05 | 2.012 | 1.998 | GB | 2021-10-05 | 0.4694% | -0.11% | 20211119 | 2021-11-19 | 1.991 | 2.013 | ML | 2021-11-19 | 0.3984% | 0.1343%
20211006 | 2021-10-06 | 2.013 | 2 | GB | 2021-10-06 | 0.5193% | -0.009999% | 20211122 | 2021-11-22 | 1.995 | 2.009 | ML | 2021-11-22 | 0.6001% | -0.06467%
20211007 | 2021-10-07 | 2.012 | 2.005 | GB | 2021-10-07 | 0.4694% | 0.24% | 20211124 | 2021-11-24 | 1.994 | 2.014 | ML | 2021-11-24 | 0.5496% | 0.1841%
20211008 | 2021-10-08 | 2.016 | 2.004 | GB | 2021-10-08 | 0.6691% | 0.19% | 20211125 | 2021-11-25 | 1.987 | 1.997 | CVD | 2021-11-25 | 0.1967% | -0.6616%
20211012 | 2021-10-12 | 2.008 | 1.998 | CVD | 2021-10-12 | 0.2696% | -0.11% | 20211126 | 2021-11-26 | 1.991 | 2.008 | GB | 2021-11-26 | 0.3984% | -0.1144%
20211013 | 2021-10-13 | 2.012 | 2.002 | CVD | 2021-10-13 | 0.4694% | 0.08999% | 20211129 | 2021-11-29 | 1.993 | 2.006 | GB | 2021-11-29 | 0.4992% | -0.2139%
20211014 | 2021-10-14 | 2.013 | 2.002 | CVD | 2021-10-14 | 0.5193% | 0.08999% | 20211202 | 2021-12-02 | 1.998 | 2.013 | GB | 2021-12-02 | 0.7513% | 0.1343%
20211017 | 2021-10-17 | 2.018 | 2.011 | EF | 2021-10-17 | 0.769% | 0.5399% | 20211207 | 2021-12-07 | 1.995 | 2.01 | GB | 2021-12-07 | 0.6001% | -0.01492%
20211019 | 2021-10-19 | 2.016 | 2.009 | GB | 2021-10-19 | 0.6691% | 0.44% | 20211209 | 2021-12-09 | 1.996 | 2.012 | ML | 2021-12-09 | 0.6505% | 0.08456%
20211020 | 2021-10-20 | 2.017 | 2.009 | ML | 2021-10-20 | 0.7191% | 0.44% | 20211210 | 2021-12-10 | 1.995 | 2.013 | ML | 2021-12-10 | 0.6001% | 0.1343%
20211021 | 2021-10-21 | 2.006 | 1.995 | GB | 2021-10-21 | 0.1698% | -0.26% | 20211213 | 2021-12-13 | 1.991 | 2.012 | ML | 2021-12-13 | 0.3984% | 0.08456%
20211022 | 2021-10-22 | 2.009 | 1.998 | GB | 2021-10-22 | 0.3196% | -0.11% | 20211216 | 2021-12-16 | 1.998 | 2.016 | GB | 2021-12-16 | 0.7513% | 0.2835%
20211027 | 2021-10-27 | 2.021 | 2.008 | GB | 2021-10-27 | 0.9188% | 0.39% | 20211220 | 2021-12-20 | 1.996 | 2.015 | CVD | 2021-12-20 | 0.6505% | 0.2338%
20211028 | 2021-10-28 | 2.01 | 2.004 | GB | 2021-10-28 | 0.3695% | 0.19% | 20211222 | 2021-12-22 | 1.993 | 2.009 | LL | 2021-12-22 | 0.4992% | -0.06467%
20211102 | 2021-11-02 | 2.009 | 1.999 | GB | 2021-11-02 | 0.3196% | -0.05999% | 20211223 | 2021-12-23 | 1.993 | 2.006 | LL | 2021-12-23 | 0.4992% | -0.2139%
20211103 | 2021-11-03 | 2.023 | 2.008 | GB | 2021-11-03 | 1.019% | 0.39% | 20220104 | 2022-01-04 | 1.997 | 1.996 | GB | 2022-01-04 | 0.7009% | -0.7113%
20211104 | 2021-11-04 | 1.989 | 1.978 | ML/ wrong calibration factor was applied | 2021-11-04 | -0.6791% | -1.11% | 20220104 | 2022-01-04 | 1.997 | 1.996 | ML | 2022-01-04 | 0.7009% | -0.7113%
20211110 | 2021-11-10 | 2.011 | 1.999 | GB | 2021-11-10 | 0.4195% | -0.05999% | 20220105 | 2022-01-05 | 1.979 | 1.996 | LL | 2022-01-05 | -0.2067% | -0.7113%
20211111 | 2021-11-11 | 2.018 | 2.008 | ML | 2021-11-11 | 0.769% | 0.39% | 20220106 | 2022-01-06 | 1.983 | 1.998 | GB | 2022-01-06 | -0.005043% | -0.6118%
20211112 | 2021-11-12 | 2.01 | 2.001 | ML | 2021-11-12 | 0.3695% | 0.04% | 20220107 | 2022-01-07 | 1.983 | 2.001 | GB | 2022-01-07 | -0.005043% | -0.4626%
20211116 | 2021-11-16 | 1.992 | 1.981 | GB | 2021-11-16 | -0.5293% | -0.9599% | 20220110 | 2022-01-10 | 1.984 | 2.005 | ML | 2022-01-10 | 0.04538% | -0.2636%
20211117 | 2021-11-17 | 2.001 | 1.99 | GB | 2021-11-17 | -0.0799% | -0.5099% | 20220112 | 2022-01-12 | 1.996 | 2.016 | GB | 2022-01-12 | 0.6505% | 0.2835%
20211118 | 2021-11-18 | 1.997 | 1.984 | ML | 2021-11-18 | -0.2796% | -0.8099% | 20220113 | 2022-01-13 | 1.988 | 2.006 | GB | 2022-01-13 | 0.2471% | -0.2139%
20211119 | 2021-11-19 | 2.009 | 1.994 | ML | 2021-11-19 | 0.3196% | -0.31% | 20220114 | 2022-01-14 | 1.995 | 2.007 | GB | 2022-01-14 | 0.6001% | -0.1642%
20211123 | 2021-11-23 | 2.002 | 1.992 | ML | 2021-11-23 | -0.02996% | -0.41% | 20220116 | 2022-01-16 | 1.976 | 2.011 | ML | 2022-01-16 | -0.358% | 0.03482%
20211124 | 2021-11-24 | 2.003 | 1.991 | ML | 2021-11-24 | 0.01997% | -0.46% | 20220117 | 2022-01-17 | 1.989 | 2.011 | ML | 2022-01-17 | 0.2975% | 0.03482%
20211125 | 2021-11-25 | 2.01 | 1.998 | CVD | 2021-11-25 | 0.3695% | -0.11% | 20220119 | 2022-01-19 | 1.998 | 2.022 | LL | 2022-01-19 | 0.7513% | 0.582%
20211130 | 2021-11-30 | 2.011 | 1.997 | GB | 2021-11-30 | 0.4195% | -0.16% | 20220120 | 2022-01-20 | 1.986 | 2.004 | ML | 2022-01-20 | 0.1462% | -0.3134%
20211201 | 2021-12-01 | 2.016 | 2.005 | GB | 2021-12-01 | 0.6691% | 0.24% | 20220121 | 2022-01-21 | 1.994 | 2.009 | ML | 2022-01-21 | 0.5496% | -0.06467%
20211202 | 2021-12-02 | 2.002 | 1.988 | GB | 2021-12-02 | -0.02996% | -0.6099% | 20220124 | 2022-01-24 | 1.993 | 2.01 | ML | 2022-01-24 | 0.4992% | -0.01492%
20211203 | 2021-12-03 | 2.005 | 1.993 | KA | 2021-12-03 | 0.1198% | -0.36% | 20220126 | 2022-01-26 | 1.986 | 2.005 | ML | 2022-01-26 | 0.1462% | -0.2636%
20211208 | 2021-12-08 | 2.001 | 1.991 | GB | 2021-12-08 | -0.0799% | -0.46% | 20220127 | 2022-01-27 | 1.996 | 2.017 | GB | 2022-01-27 | 0.6505% | 0.3333%
20211209 | 2021-12-09 | 2.005 | 1.993 | ML | 2021-12-09 | 0.1198% | -0.36% | 20220202 | 2022-02-02 | 1.995 | 2.014 | ML | 2022-02-02 | 0.6001% | 0.1841%
20211210 | 2021-12-10 | 2.005 | 1.996 | ML | 2021-12-10 | 0.1198% | -0.21% | 20220203 | 2022-02-03 | 1.99 | 2.009 | ML | 2022-02-03 | 0.3479% | -0.06467%
20211213 | 2021-12-13 | 2.005 | 1.992 | ML | 2021-12-13 | 0.1198% | -0.41% | 20220204 | 2022-02-04 | 1.997 | 2.015 | ML | 2022-02-04 | 0.7009% | 0.2338%
20211215 | 2021-12-15 | 2 | 1.993 | ML | 2021-12-15 | -0.1298% | -0.36% | 20220207 | 2022-02-07 | 1.988 | 2.011 | ML | 2022-02-07 | 0.2471% | 0.03482%
20211217 | 2021-12-17 | 1.999 | 1.988 | ML | 2021-12-17 | -0.1798% | -0.6099% | 20220209 | 2022-02-09 | 1.992 | 2.01 | GB | 2022-02-09 | 0.4488% | -0.01492%
20211221 | 2021-12-21 | 2.004 | 1.993 | CVD | 2021-12-21 | 0.06991% | -0.36% | 20220214 | 2022-02-14 | 1.987 | 2.002 | ML | 2022-02-14 | 0.1967% | -0.4129%
20211223 | 2021-12-23 | 2.011 | 2.001 | LL | 2021-12-23 | 0.4195% | 0.04% | 20220217 | 2022-02-17 | 1.988 | 2.007 | GB | 2022-02-17 | 0.2471% | -0.1642%
20220104 | 2022-01-04 | 2.028 | 2.007 | GB | 2022-01-04 | 1.268% | 0.34% | 20220218 | 2022-02-18 | 1.986 | 2.001 | ML | 2022-02-18 | 0.1462% | -0.4626%
20220105 | 2022-01-05 | 1.998 | 1.984 | GB, kört med 0,8 grader lägre temperatur i dosimetry manager. | 2022-01-05 | -0.2297% | -0.8099% | 20220221 | 2022-02-21 | 1.986 | 2.004 | ML | 2022-02-21 | 0.1462% | -0.3134%
20220106 | 2022-01-06 | 2.002 | 1.987 | GB | 2022-01-06 | -0.02996% | -0.6599% | 20220223 | 2022-02-23 | 1.991 | 2.008 | ML | 2022-02-23 | 0.3984% | -0.1144%
20220107 | 2022-01-07 | 2.006 | 1.992 | LL | 2022-01-07 | 0.1698% | -0.41% | 20220224 | 2022-02-24 | 1.991 | 2.008 | GB | 2022-02-24 | 0.3984% | -0.1144%
20220111 | 2022-01-11 | 2.014 | 2.002 | ML | 2022-01-11 | 0.5693% | 0.08999% | 20220228 | 2022-02-28 | 1.989 | 2.002 | ML | 2022-02-28 | 0.2975% | -0.4129%
20220112 | 2022-01-12 | 2.013 | 1.998 | GB | 2022-01-12 | 0.5193% | -0.11% | 20220302 | 2022-03-02 | 1.993 | 2.01 | ML | 2022-03-02 | 0.4992% | -0.01492%
20220113 | 2022-01-13 | 2.006 | 1.993 | ML | 2022-01-13 | 0.1698% | -0.36% | 20220303 | 2022-03-03 | 1.981 | 1.997 | GB | 2022-03-03 | -0.1059% | -0.6616%
20220114 | 2022-01-14 | 2.01 | 1.998 | GB | 2022-01-14 | 0.3695% | -0.11% | 20220304 | 2022-03-04 | 1.981 | 1.999 | GB | 2022-03-04 | -0.1059% | -0.5621%
20220118 | 2022-01-18 | 2.01 | 1.995 | ML | 2022-01-18 | 0.3695% | -0.26% | 20220307 | 2022-03-07 | 1.989 | 2.002 | CVD | 2022-03-07 | 0.2975% | -0.4129%
20220119 | 2022-01-19 | 2.008 | 1.998 | GB | 2022-01-19 | 0.2696% | -0.11% | 20220310 | 2022-03-10 | 1.986 | 2 | CVD | 2022-03-10 | 0.1462% | -0.5124%
20220120 | 2022-01-20 | 2.001 | 1.985 | ML | 2022-01-20 | -0.0799% | -0.7599% | 20220315 | 2022-03-15 | 1.994 | 2.008 | GB | 2022-03-15 | 0.5496% | -0.1144%
20220121 | 2022-01-21 | 2.006 | 1.991 | ML | 2022-01-21 | 0.1698% | -0.46% | 20220318 | 2022-03-18 | 1.991 | 2.011 | GB | 2022-03-18 | 0.3984% | 0.03482%
20220125 | 2022-01-25 | 1.998 | 1.986 | CVD | 2022-01-25 | -0.2297% | -0.7099% | 20220321 | 2022-03-21 | 1.992 | 2.009 | ML | 2022-03-21 | 0.4488% | -0.06467%
20220126 | 2022-01-26 | 2.003 | 1.986 | ML | 2022-01-26 | 0.01997% | -0.7099% | 20220322 | 2022-03-22 | 1.989 | 1.999 | CVD | 2022-03-22 | 0.2975% | -0.5621%
20220128 | 2022-01-28 | 2.005 | 1.987 | ML | 2022-01-28 | 0.1198% | -0.6599% | 20220324 | 2022-03-24 | 1.987 | 2.005 | GB | 2022-03-24 | 0.1967% | -0.2636%
20220201 | 2022-02-01 | 1.998 | 1.983 | CVD | 2022-02-01 | -0.2297% | -0.8599% | 20220328 | 2022-03-28 | 1.977 | 1.995 | GB | 2022-03-28 | -0.3076% | -0.7611%
20220202 | 2022-02-02 | 2.007 | 1.99 | ML | 2022-02-02 | 0.2197% | -0.5099% | 20220330 | 2022-03-30 | 1.998 | 2.017 | ML | 2022-03-30 | 0.7513% | 0.3333%
20220203 | 2022-02-03 | 2.005 | 1.991 | ML | 2022-02-03 | 0.1198% | -0.46% | 20220331 | 2022-03-31 | 1.991 | 2.009 | ML | 2022-03-31 | 0.3984% | -0.06467%
20220204 | 2022-02-04 | 2.008 | 1.993 | GB | 2022-02-04 | 0.2696% | -0.36% | 20220401 | 2022-04-01 | 1.984 | 2.008 | ML | 2022-04-01 | 0.04538% | -0.1144%
20220208 | 2022-02-08 | 2.007 | 1.993 | LL | 2022-02-08 | 0.2197% | -0.36% | 20220404 | 2022-04-04 | 1.986 | 2.003 | ML | 2022-04-04 | 0.1462% | -0.3631%
20220209 | 2022-02-09 | 2.02 | 2.005 | ML | 2022-02-09 | 0.8689% | 0.24% | 20220406 | 2022-04-06 | 1.981 | 1.999 | ML | 2022-04-06 | -0.1059% | -0.5621%
20220215 | 2022-02-15 | 1.995 | 1.984 | ML | 2022-02-15 | -0.3795% | -0.8099% | 20220407 | 2022-04-07 | 1.989 | 2.007 | GB | 2022-04-07 | 0.2975% | -0.1642%
20220216 | 2022-02-16 | 1.996 | 1.983 | GB | 2022-02-16 | -0.3296% | -0.8599% | 20220408 | 2022-04-08 | 1.976 | 1.995 | GB | 2022-04-08 | -0.358% | -0.7611%
20220217 | 2022-02-17 | 2.003 | 1.989 | ML | 2022-02-17 | 0.01997% | -0.5599% | 20220411 | 2022-04-11 | 1.994 | 2.019 | LL | 2022-04-11 | 0.5496% | 0.4328%
20220218 | 2022-02-18 | 2.002 | 1.988 | GB | 2022-02-18 | -0.02996% | -0.6099% | 20220414 | 2022-04-14 | 1.992 | 2.009 | CVD | 2022-04-14 | 0.4488% | -0.06467%
20220222 | 2022-02-22 | 2.007 | 1.993 | CVD | 2022-02-22 | 0.2197% | -0.36% | 20220416 | 2022-04-16 | 0 | 1.993 | CVD | 2022-04-16 |  | -0.8606%
20220223 | 2022-02-23 | 2 | 1.992 | ML | 2022-02-23 | -0.1298% | -0.41% | 20220419 | 2022-04-19 | 1.978 | 0 | AD | 2022-04-19 | -0.2572%
20220225 | 2022-02-25 | 2.007 | 1.992 | KA | 2022-02-25 | 0.2197% | -0.41% | 20220421 | 2022-04-21 | 1.996 | 2.007 | GB | 2022-04-21 | 0.6505% | -0.1642%
20220301 | 2022-03-01 | 2.007 | 1.995 | ML | 2022-03-01 | 0.2197% | -0.26% | 20220422 | 2022-04-22 | 1.988 | 2.005 | GB | 2022-04-22 | 0.2471% | -0.2636%
20220302 | 2022-03-02 | 2.008 | 1.993 | ML | 2022-03-02 | 0.2696% | -0.36% | 20220425 | 2022-04-25 | 1.988 | 2.009 | CVD | 2022-04-25 | 0.2471% | -0.06467%
20220303 | 2022-03-03 | 2.004 | 1.994 | ML | 2022-03-03 | 0.06991% | -0.31% | 20220427 | 2022-04-27 | 1.986 | 2.002 | GB | 2022-04-27 | 0.1462% | -0.4129%
20220304 | 2022-03-04 | 2.013 | 2 | GB | 2022-03-04 | 0.5193% | -0.009999% | 20220504 | 2022-05-04 | 2.004 | 2.022 | LL | 2022-05-04 | 1.054% | 0.582%
20220308 | 2022-03-08 | 2.011 | 1.999 | CVD | 2022-03-08 | 0.4195% | -0.05999% | 20220505 | 2022-05-05 | 1.986 | 2.006 | ML | 2022-05-05 | 0.1462% | -0.2139%
20220309 | 2022-03-09 | 2.013 | 2 | ML | 2022-03-09 | 0.5193% | -0.009999% | 20220509 | 2022-05-09 | 1.985 | 2.005 | LL | 2022-05-09 | 0.09581% | -0.2636%
20220310 | 2022-03-10 | 2.007 | 1.995 | CVD | 2022-03-10 | 0.2197% | -0.26% | 20220511 | 2022-05-11 | 1.983 | 2.002 | ML | 2022-05-11 | -0.005043% | -0.4129%
20220314 | 2022-03-14 | 2.016 | 2.001 | GB | 2022-03-14 | 0.6691% | 0.04% | 20220512 | 2022-05-12 | 1.988 | 2.002 | AD | 2022-05-12 | 0.2471% | -0.4129%
20220316 | 2022-03-16 | 2.008 | 1.997 | GB | 2022-03-16 | 0.2696% | -0.16% | 20220513 | 2022-05-13 | 1.981 | 2.003 | LL | 2022-05-13 | -0.1059% | -0.3631%
20220317 | 2022-03-17 | 2.014 | 2.001 | ML | 2022-03-17 | 0.5693% | 0.04% | 20220516 | 2022-05-16 | 1.988 | 2.008 | CVD | 2022-05-16 | 0.2471% | -0.1144%
20220318 | 2022-03-18 | 2.016 | 2.004 | GB | 2022-03-18 | 0.6691% | 0.19% | 20220518 | 2022-05-18 | 1.981 | 1.996 | ML | 2022-05-18 | -0.1059% | -0.7113%
20220323 | 2022-03-23 | 2.017 | 2.004 | ML | 2022-03-23 | 0.7191% | 0.19% | 20220519 | 2022-05-19 | 1.993 | 2.015 | GB | 2022-05-19 | 0.4992% | 0.2338%
20220324 | 2022-03-24 | 2.008 | 1.994 | CVD | 2022-03-24 | 0.2696% | -0.31% | 20220520 | 2022-05-20 | 1.986 | 2.004 | CVD | 2022-05-20 | 0.1462% | -0.3134%
20220325 | 2022-03-25 | 2.008 | 1.994 | ML | 2022-03-25 | 0.2696% | -0.31% | 20220523 | 2022-05-23 | 1.981 | 1.99 | ML | 2022-05-23 | -0.1059% | -1.01%
20220329 | 2022-03-29 | 2.005 | 1.995 | GB | 2022-03-29 | 0.1198% | -0.26% | 20220525 | 2022-05-25 | 1.994 | 2.014 | ML | 2022-05-25 | 0.5496% | 0.1841%
20220330 | 2022-03-30 | 2.014 | 2.001 | GB | 2022-03-30 | 0.5693% | 0.04% | 20220527 | 2022-05-27 | 1.995 | 2.018 | ML | 2022-05-27 | 0.6001% | 0.383%
20220401 | 2022-04-01 | 2.008 | 1.995 | ML | 2022-04-01 | 0.2696% | -0.26% | 20220530 | 2022-05-30 | 2.003 | 2.028 | GB | 2022-05-30 | 1.003% | 0.8805%
20220405 | 2022-04-05 | 2.013 | 1.998 | ML | 2022-04-05 | 0.5193% | -0.11% | 20220603 | 2022-06-03 | 1.989 | 2.009 | GB | 2022-06-03 | 0.2975% | -0.06467%
20220406 | 2022-04-06 | 2 | 1.987 | ML | 2022-04-06 | -0.1298% | -0.6599% | 20220607 | 2022-06-07 | 1.993 | 2.008 | YM | 2022-06-07 | 0.4992% | -0.1144%
20220407 | 2022-04-07 | 2.009 | 1.996 | ML | 2022-04-07 | 0.3196% | -0.21% | 20220609 | 2022-06-09 | 1.995 | 2.015 | GB | 2022-06-09 | 0.6001% | 0.2338%
20220408 | 2022-04-08 | 2.002 | 1.989 | GB | 2022-04-08 | -0.02996% | -0.5599% | 20220610 | 2022-06-10 | 1.992 | 2.015 | GB | 2022-06-10 | 0.4488% | 0.2338%
20220412 | 2022-04-12 | 2.005 | 1.993 | ML | 2022-04-12 | 0.1198% | -0.36% | 20220613 | 2022-06-13 | 2.007 | 2.033 | GB | 2022-06-13 | 1.205% | 1.129%
20220413 | 2022-04-13 | 2.008 | 1.992 | ML | 2022-04-13 | 0.2696% | -0.41% | 20220614 | 2022-06-14 | 1.991 | 2.008 | GB | 2022-06-14 | 0.3984% | -0.1144%
20220414 | 2022-04-14 | 2.008 | 1.994 | LL | 2022-04-14 | 0.2696% | -0.31% | 20220616 | 2022-06-16 | 1.996 | 2.016 | ML | 2022-06-16 | 0.6505% | 0.2835%
20220420 | 2022-04-20 | 2.007 | 1.995 | AD | 2022-04-20 | 0.2197% | -0.26% | 20220620 | 2022-06-20 | 1.986 | 2.005 | ML | 2022-06-20 | 0.1462% | -0.2636%
20220421 | 2022-04-21 | 2.017 | 2.004 | AD | 2022-04-21 | 0.7191% | 0.19% | 20220623 | 2022-06-23 | 1.985 | 2.006 | AD | 2022-06-23 | 0.09581% | -0.2139%
20220422 | 2022-04-22 | 2.006 | 1.993 | GB | 2022-04-22 | 0.1698% | -0.36% | 20220630 | 2022-06-30 | 1.999 | 2.02 | ML | 2022-06-30 | 0.8018% | 0.4825%
20220426 | 2022-04-26 | 2.003 | 1.99 | EF | 2022-04-26 | 0.01997% | -0.5099% | 20220701 | 2022-07-01 | 2.02 | 2.039 | ML | 2022-07-01 | 1.861% | 1.428%
20220503 | 2022-05-03 | 2.017 | 2.005 | GB | 2022-05-03 | 0.7191% | 0.24% | 20220706 | 2022-07-06 | 2.005 | 2.024 | ML | 2022-07-06 | 1.104% | 0.6815%
20220504 | 2022-05-04 | 2.01 | 1.999 | GB | 2022-05-04 | 0.3695% | -0.05999% | 20220707 | 2022-07-07 | 2.001 | 2.02 | ML | 2022-07-07 | 0.9026% | 0.4825%
20220506 | 2022-05-06 | 2.008 | 1.994 | LL | 2022-05-06 | 0.2696% | -0.31% | 20220708 | 2022-07-08 | 2.009 | 2.028 | CVD | 2022-07-08 | 1.306% | 0.8805%
20220510 | 2022-05-10 | 2.008 | 1.996 | LL | 2022-05-10 | 0.2696% | -0.21% | 20220712 | 2022-07-12 | 1.976 | 1.989 | ML | 2022-07-12 | -0.358% | -1.06%
20220511 | 2022-05-11 | 1.994 | 1.98 | LL | 2022-05-11 | -0.4294% | -1.01% | 20220714 | 2022-07-14 | 1.972 | 1.976 | ML | 2022-07-14 | -0.5597% | -1.706%
20220512 | 2022-05-12 | 1.992 | 1.98 | ML | 2022-05-12 | -0.5293% | -1.01% | 20220715 | 2022-07-15 | 1.976 | 1.992 | ef | 2022-07-15 | -0.358% | -0.9103%
20220513 | 2022-05-13 | 2.006 | 1.993 | AD | 2022-05-13 | 0.1698% | -0.36% | 20220718 | 2022-07-18 | 1.983 | 1.989 | YM | 2022-07-18 | -0.005043% | -1.06%
20220517 | 2022-05-17 | 1.991 | 1.98 | ML | 2022-05-17 | -0.5792% | -1.01% | 20220720 | 2022-07-20 | 1.975 | 1.986 | ML | 2022-07-20 | -0.4085% | -1.209%
20220518 | 2022-05-18 | 1.997 | 1.986 | CVD | 2022-05-18 | -0.2796% | -0.7099% | 20220721 | 2022-07-21 | 1.971 | 1.979 | ML | 2022-07-21 | -0.6102% | -1.557%
20220519 | 2022-05-19 | 1.997 | 1.985 | ML | 2022-05-19 | -0.2796% | -0.7599% | 20220722 | 2022-07-22 | 1.979 | 1.972 | ML | 2022-07-22 | -0.2067% | -1.905%
20220520 | 2022-05-20 | 1.994 | 1.983 | GB | 2022-05-20 | -0.4294% | -0.8599% | 20220725 | 2022-07-25 | 1.981 | 1.983 | ML | 2022-07-25 | -0.1059% | -1.358%
20220524 | 2022-05-24 | 2 | 1.987 | EF YM | 2022-05-24 | -0.1298% | -0.6599% | 20220727 | 2022-07-27 | 1.972 | 1.977 | AC | 2022-07-27 | -0.5597% | -1.656%
20220525 | 2022-05-25 | 2.013 | 2.001 | ML | 2022-05-25 | 0.5193% | 0.04% | 20220801 | 2022-08-01 | 1.979 | 1.983 | AC | 2022-08-01 | -0.2067% | -1.358%
20220601 | 2022-06-01 | 2.013 | 2 | ML | 2022-06-01 | 0.5193% | -0.009999% | 20220803 | 2022-08-03 | 1.984 | 1.988 | cvd | 2022-08-03 | 0.04538% | -1.109%
20220602 | 2022-06-02 | 2.011 | 1.997 | GB | 2022-06-02 | 0.4195% | -0.16% | 20220805 | 2022-08-05 | 1.977 | 1.977 | GB | 2022-08-05 | -0.3076% | -1.656%
20220608 | 2022-06-08 | 2.008 | 1.995 | YM | 2022-06-08 | 0.2696% | -0.26% | 20220808 | 2022-08-08 | 1.979 | 1.984 | AC, Matrixx2 | 2022-08-08 | -0.2067% | -1.308%
20220609 | 2022-06-09 | 2.005 | 1.995 | YM | 2022-06-09 | 0.1198% | -0.26% | 20220809 | 2022-08-09 | 1.97 | 1.981 | AC | 2022-08-09 | -0.6606% | -1.457%
20220610 | 2022-06-10 | 2.004 | 1.992 | GB | 2022-06-10 | 0.06991% | -0.41% | 20220810 | 2022-08-10 | 1.999 | 1.999 | AC | 2022-08-10 | 0.8018% | -0.5621%
20220615 | 2022-06-15 | 2.003 | 1.99 | GB | 2022-06-15 | 0.01997% | -0.5099% | 20220815 | 2022-08-15 | 1.98 | 1.985 | GB | 2022-08-15 | -0.1563% | -1.259%
20220616 | 2022-06-16 | 2.005 | 1.993 | GB | 2022-06-16 | 0.1198% | -0.36% | 20220818 | 2022-08-18 | 1.976 | 1.982 | AC | 2022-08-18 | -0.358% | -1.408%
20220617 | 2022-06-17 | 2.01 | 1.997 | GB | 2022-06-17 | 0.3695% | -0.16% | 20220819 | 2022-08-19 | 1.973 | 1.975 | AC | 2022-08-19 | -0.5093% | -1.756%
20220621 | 2022-06-21 | 2.008 | 1.997 | ML | 2022-06-21 | 0.2696% | -0.16% | 20220822 | 2022-08-22 | 1.961 | 1.973 | GB | 2022-08-22 | -1.114% | -1.855%
20220622 | 2022-06-22 | 2.005 | 1.991 | YM | 2022-06-22 | 0.1198% | -0.46% | 20220824 | 2022-08-24 | 1.97 | 1.98 | YM | 2022-08-24 | -0.6606% | -1.507%
20220623 | 2022-06-23 | 2.007 | 1.994 | ML | 2022-06-23 | 0.2197% | -0.31% | 20220825 | 2022-08-25 | 1.975 | 1.973 | AC | 2022-08-25 | -0.4085% | -1.855%
20220627 | 2022-06-27 | 2.012 | 2.002 | ML | 2022-06-27 | 0.4694% | 0.08999% | 20220826 | 2022-08-26 | 1.968 | 1.982 | YM | 2022-08-26 | -0.7614% | -1.408%
20220629 | 2022-06-29 | 2.018 | 2.004 | ML | 2022-06-29 | 0.769% | 0.19% | 20220829 | 2022-08-29 | 1.961 | 1.981 | YM | 2022-08-29 | -1.114% | -1.457%
20220630 | 2022-06-30 | 2.013 | 2.002 | AD | 2022-06-30 | 0.5193% | 0.08999% | 20220831 | 2022-08-31 | 1.976 | 1.982 | ML, Matrixx 1 | 2022-08-31 | -0.358% | -1.408%
20220701 | 2022-07-01 | 2.02 | 2.008 | ML | 2022-07-01 | 0.8689% | 0.39% | 20220901 | 2022-09-01 | 1.988 | 1.994 | GB | 2022-09-01 | 0.2471% | -0.8108%
20220705 | 2022-07-05 | 2.005 | 1.996 | ML | 2022-07-05 | 0.1198% | -0.21% | 20220902 | 2022-09-02 | 1.985 | 1.989 | YM | 2022-09-02 | 0.09581% | -1.06%
20220706 | 2022-07-06 | 2.005 | 1.998 | ML | 2022-07-06 | 0.1198% | -0.11% | 20220905 | 2022-09-05 | 1.991 | 1.989 | CVD KW | 2022-09-05 | 0.3984% | -1.06%
20220708 | 2022-07-08 | 2.006 | 1.999 | ML | 2022-07-08 | 0.1698% | -0.05999% | 20220907 | 2022-09-07 | 1.968 | 1.986 | ML | 2022-09-07 | -0.7614% | -1.209%
20220711 | 2022-07-11 | 1.995 | 1.98 | ML | 2022-07-11 | -0.3795% | -1.01% | 20220908 | 2022-09-08 | 1.988 | 1.988 | CVD | 2022-09-08 | 0.2471% | -1.109%
20220713 | 2022-07-13 | 2.003 | 1.987 | YM | 2022-07-13 | 0.01997% | -0.6599% | 20220909 | 2022-09-09 | 1.968 | 1.986 | YM | 2022-09-09 | -0.7614% | -1.209%
20220714 | 2022-07-14 | 1.994 | 1.982 | ML | 2022-07-14 | -0.4294% | -0.9099% | 20220912 | 2022-09-12 | 1.965 | 1.979 | ML | 2022-09-12 | -0.9127% | -1.557%
20220719 | 2022-07-19 | 1.989 | 1.974 | YM | 2022-07-19 | -0.6791% | -1.31% | 20220914 | 2022-09-14 | 1.981 | 1.983 | GB | 2022-09-14 | -0.1059% | -1.358%
20220720 | 2022-07-20 | 1.98 | 1.969 | YM | 2022-07-20 | -1.129% | -1.56% | 20220915 | 2022-09-15 | 1.986 | 1.992 | GB | 2022-09-15 | 0.1462% | -0.9103%
20220721 | 2022-07-21 | 1.988 | 1.973 | ML | 2022-07-21 | -0.7291% | -1.36% | 20220916 | 2022-09-16 | 1.989 | 1.984 | YM | 2022-09-16 | 0.2975% | -1.308%
20220722 | 2022-07-22 | 1.991 | 1.976 | ML | 2022-07-22 | -0.5792% | -1.21% | 20220919 | 2022-09-19 | 1.987 | 1.981 | ML | 2022-09-19 | 0.1967% | -1.457%
20220726 | 2022-07-26 | 1.991 | 1.977 | ML | 2022-07-26 | -0.5792% | -1.16% | 20220921 | 2022-09-21 | 1.981 | 1.986 | GB | 2022-09-21 | -0.1059% | -1.209%
20220728 | 2022-07-28 | 1.988 | 1.974 | ML | 2022-07-28 | -0.7291% | -1.31% | 20220922 | 2022-09-22 | 1.984 | 1.992 | GB | 2022-09-22 | 0.04538% | -0.9103%
20220729 | 2022-07-29 | 1.988 | 1.971 | AC | 2022-07-29 | -0.7291% | -1.46% | 20220923 | 2022-09-23 | 1.989 | 1.979 | YM | 2022-09-23 | 0.2975% | -1.557%
20220802 | 2022-08-02 | 1.994 | 1.976 | KA | 2022-08-02 | -0.4294% | -1.21% | 20220926 | 2022-09-26 | 1.978 | 1.979 | YM | 2022-09-26 | -0.2572% | -1.557%
20220803 | 2022-08-03 | 1.998 | 1.978 | AC | 2022-08-03 | -0.2297% | -1.11% | 20220928 | 2022-09-28 | 1.967 | 1.983 | CVD, MatriXX_2 | 2022-09-28 | -0.8119% | -1.358%
20220804 | 2022-08-04 | 1.998 | 1.983 | YM | 2022-08-04 | -0.2297% | -0.8599% | 20221009 | 2022-10-09 | 1.997 | 2.022 | AC, MatriXX2, afternoon measurements | 2022-10-09 | 0.7009% | 0.582%
20220809 | 2022-08-09 | 2.004 | 1.982 | GB | 2022-08-09 | 0.06991% | -0.9099% | 20221011 | 2022-10-11 | 2 | 2.023 | AC | 2022-10-11 | 0.8522% | 0.6317%
20220810 | 2022-08-10 | 1.998 | 1.984 | GB | 2022-08-10 | -0.2297% | -0.8099% | 20221012 | 2022-10-12 | 1.98 | 2.008 | ML,MatriXX_1 | 2022-10-12 | -0.1563% | -0.1144%
20220812 | 2022-08-12 | 1.992 | 1.995 | KA | 2022-08-12 | -0.5293% | -0.26% | 20221013 | 2022-10-13 | 1.986 | 2.008 | YM | 2022-10-13 | 0.1462% | -0.1144%
20220816 | 2022-08-16 | 2.009 | 1.984 | YM | 2022-08-16 | 0.3196% | -0.8099% | 20221014 | 2022-10-14 | 1.975 | 1.998 | CVD Matrixx1 | 2022-10-14 | -0.4085% | -0.6118%
20220817 | 2022-08-17 | 1.994 | 1.978 | YM | 2022-08-17 | -0.4294% | -1.11% | 20221021 | 2022-10-21 | 1.984 | 2.005 | YM | 2022-10-21 | 0.04538% | -0.2636%
20220818 | 2022-08-18 | 1.993 | 1.978 | GB | 2022-08-18 | -0.4794% | -1.11% | 20221024 | 2022-10-24 | 2.004 | 2.016 | ML, MatriXX_2 | 2022-10-24 | 1.054% | 0.2835%
20220819 | 2022-08-19 | 1.999 | 1.985 | AC | 2022-08-19 | -0.1798% | -0.7599% | 20221024 | 2022-10-24 | 1.985 | 2.01 | YM | 2022-10-24 | 0.09581% | -0.01492%
20220823 | 2022-08-23 | 2.004 | 1.98 | YM | 2022-08-23 | 0.06991% | -1.01% | 20221027 | 2022-10-27 | 2.006 | 2.022 | ML | 2022-10-27 | 1.155% | 0.582%
20220824 | 2022-08-24 | 2.001 | 1.984 | GB | 2022-08-24 | -0.0799% | -0.8099% | 20221031 | 2022-10-31 | 1.979 | 2.001 | ML | 2022-10-31 | -0.2067% | -0.4626%
20220825 | 2022-08-25 | 2.009 | 1.991 | AC | 2022-08-25 | 0.3196% | -0.46% | 20221102 | 2022-11-02 | 1.982 | 1.995 | KA | 2022-11-02 | -0.05547% | -0.7611%
20220826 | 2022-08-26 | 2.009 | 2.004 | AC | 2022-08-26 | 0.3196% | 0.19% | 20221103 | 2022-11-03 | 1.991 | 2.015 | AC | 2022-11-03 | 0.3984% | 0.2338%
20220830 | 2022-08-30 | 1.998 | 1.98 | CVD | 2022-08-30 | -0.2297% | -1.01% | 20221107 | 2022-11-07 | 1.985 | 2.001 | ML | 2022-11-07 | 0.09581% | -0.4626%
20220901 | 2022-09-01 | 1.999 | 1.979 | ML | 2022-09-01 | -0.1798% | -1.06% | 20221109 | 2022-11-09 | 1.989 | 2 | ML ? | 2022-11-09 | 0.2975% | -0.5124%
20220902 | 2022-09-02 | 1.999 | 1.98 | GB | 2022-09-02 | -0.1798% | -1.01% | 20221110 | 2022-11-10 | 1.972 | 1.99 | ML | 2022-11-10 | -0.5597% | -1.01%
20220906 | 2022-09-06 | 1.998 | 1.979 | ML | 2022-09-06 | -0.2297% | -1.06% | 20221111 | 2022-11-11 | 1.983 | 1.998 | CVD, Matrix 1 | 2022-11-11 | -0.005043% | -0.6118%
20220907 | 2022-09-07 | 1.999 | 1.982 | CVD KW | 2022-09-07 | -0.1798% | -0.9099% | 20221118 | 2022-11-18 | 1.979 | 1.997 | YM | 2022-11-18 | -0.2067% | -0.6616%
20220908 | 2022-09-08 | 1.997 | 1.981 | ML | 2022-09-08 | -0.2796% | -0.9599% | 20221118 | 2022-11-18 | 1.985 | 2.007 | ML, REFERENCE IMAGES | 2022-11-18 | 0.09581% | -0.1642%
20220909 | 2022-09-09 | 1.988 | 1.974 | CVD | 2022-09-09 | -0.7291% | -1.31% | 20221118 | 2022-11-18 | 1.982 | 2.008 | ML | 2022-11-18 | -0.05547% | -0.1144%
20220913 | 2022-09-13 | 1.995 | 1.98 | ML | 2022-09-13 | -0.3795% | -1.01% | 20221118 | 2022-11-18 | 1.985 | 2.006 | ML | 2022-11-18 | 0.09581% | -0.2139%
20220914 | 2022-09-14 | 1.996 | 1.982 | ML | 2022-09-14 | -0.3296% | -0.9099% | 20221118 | 2022-11-18 | 1.987 | 2.009 | ML, REFERENCE IMAGES | 2022-11-18 | 0.1967% | -0.06467%
20220915 | 2022-09-15 | 1.992 | 0 | GB, R31 not measured because of downtime in the morning | 2022-09-15 | -0.5293% | -100% | 20221118 | 2022-11-18 | 1.987 | 2.01 | ML | 2022-11-18 | 0.1967% | -0.01492%
20220916 | 2022-09-16 | 1.989 | 1.975 | GB | 2022-09-16 | -0.6791% | -1.26% | 20221118 | 2022-11-18 | 1.987 | 2.01 | ML | 2022-11-18 | 0.1967% | -0.01492%
20220920 | 2022-09-20 | 1.992 | 1.975 | YM | 2022-09-20 | -0.5293% | -1.26% | 20221118 | 2022-11-18 | 1.987 | 2.007 | ML, REFERENCE IMAGES | 2022-11-18 | 0.1967% | -0.1642%
20220921 | 2022-09-21 | 1.996 | 1.982 | ML | 2022-09-21 | -0.3296% | -0.9099% | 20221118 | 2022-11-18 | 1.983 | 2.006 | ML | 2022-11-18 | -0.005043% | -0.2139%
20220922 | 2022-09-22 | 1.993 | 1.984 | GB | 2022-09-22 | -0.4794% | -0.8099% | 20221118 | 2022-11-18 | 1.988 | 2.008 | ML | 2022-11-18 | 0.2471% | -0.1144%
20220923 | 2022-09-23 | 1.99 | 1.977 | GB | 2022-09-23 | -0.6292% | -1.16% | 20221121 | 2022-11-21 | 1.991 | 2.004 | ML | 2022-11-21 | 0.3984% | -0.3134%
20220927 | 2022-09-27 | 1.998 | 1.986 | GB | 2022-09-27 | -0.2297% | -0.7099% | 20221123 | 2022-11-23 | 1.983 | 1.997 | YM | 2022-11-23 | -0.005043% | -0.6616%
20220928 | 2022-09-28 | 2.004 | 1.991 | YM | 2022-09-28 | 0.06991% | -0.46% | 20221124 | 2022-11-24 | 1.963 | 1.984 | CVD | 2022-11-24 | -1.014% | -1.308%
20220930 | 2022-09-30 | 2.005 | 1.992 | ML | 2022-09-30 | 0.1198% | -0.41% | 20221125 | 2022-11-25 | 1.994 | 1.996 | GB | 2022-11-25 | 0.5496% | -0.7113%
20221009 | 2022-10-09 | 2.004 | 1.993 | CVD | 2022-10-09 | 0.06991% | -0.36% | 20221128 | 2022-11-28 | 1.992 | 2.005 | CVD, Matlab R2018b, new Citrix portal | 2022-11-28 | 0.4488% | -0.2636%
20221011 | 2022-10-11 | 2.001 | 1.997 | AC, MatriXX2, aftrernoon measurements | 2022-10-11 | -0.0799% | -0.16% | 20221130 | 2022-11-30 | 1.987 | 1.994 | CVD, Matlab R2018b, new Citrix portal | 2022-11-30 | 0.1967% | -0.8108%
20221012 | 2022-10-12 | 2.004 | 1.995 | AC | 2022-10-12 | 0.06991% | -0.26% | 20221202 | 2022-12-02 | 1.989 | 1.998 | CVD | 2022-12-02 | 0.2975% | -0.6118%
20221013 | 2022-10-13 | 2.009 | 1.997 | ML | 2022-10-13 | 0.3196% | -0.16% | 20221205 | 2022-12-05 | 1.954 | 1.997 | ML | 2022-12-05 | -1.467% | -0.6616%
20221014 | 2022-10-14 | 2.005 | 1.995 | YM | 2022-10-14 | 0.1198% | -0.26% | 20221207 | 2022-12-07 | 2.005 | 2.016 | ML | 2022-12-07 | 1.104% | 0.2835%
20221019 | 2022-10-19 | 2 | 1.989 | ML | 2022-10-19 | -0.1298% | -0.5599% | 20221208 | 2022-12-08 | 1.981 | 1.994 | ML | 2022-12-08 | -0.1059% | -0.8108%
20221020 | 2022-10-20 | 2.005 | 1.993 | ML | 2022-10-20 | 0.1198% | -0.36% | 20221209 | 2022-12-09 | 1.987 | 2.001 | ML | 2022-12-09 | 0.1967% | -0.4626%
20221021 | 2022-10-21 | 2 | 1.988 | CVD Matrixx2 | 2022-10-21 | -0.1298% | -0.6099% | 20221213 | 2022-12-13 | 1.981 | 1.997 | AC | 2022-12-13 | -0.1059% | -0.6616%
20221025 | 2022-10-25 | 2.002 | 1.991 | on | 2022-10-25 | -0.02996% | -0.46% | 20221213 | 2022-12-13 | 1.981 | 1.997 | ML | 2022-12-13 | -0.1059% | -0.6616%
20221101 | 2022-11-01 | 2.019 | 2.013 | YM | 2022-11-01 | 0.8189% | 0.6399% | 20221215 | 2022-12-15 | 1.978 | 1.994 | AC, MatriXX_2 | 2022-12-15 | -0.2572% | -0.8108%
20221102 | 2022-11-02 | 2.012 | 2.001 | ML | 2022-11-02 | 0.4694% | 0.04% | 20221219 | 2022-12-19 | 1.994 | 2.016 | AC, MatriXX_2 | 2022-12-19 | 0.5496% | 0.2835%
20221104 | 2022-11-04 | 2.016 | 2.004 | MatriXX_1 | 2022-11-04 | 0.6691% | 0.19% | 20221221 | 2022-12-21 | 1.995 | 2.014 | AC, MatriXX_2 | 2022-12-21 | 0.6001% | 0.1841%
20221108 | 2022-11-08 | 1.998 | 1.987 | YM | 2022-11-08 | -0.2297% | -0.6599% | 20221222 | 2022-12-22 | 1.995 | 2.015 | AC, MatriXX_2 | 2022-12-22 | 0.6001% | 0.2338%
20221109 | 2022-11-09 | 2.003 | 1.994 | ML | 2022-11-09 | 0.01997% | -0.31% | 20221223 | 2022-12-23 | 1.99 | 2.01 | AC, MatriXX_2 | 2022-12-23 | 0.3479% | -0.01492%
20221110 | 2022-11-10 | 1.998 | 1.991 | ML | 2022-11-10 | -0.2297% | -0.46% | 20221227 | 2022-12-27 | 1.987 | 2.011 | AC, MatriXX_2 | 2022-12-27 | 0.1967% | 0.03482%
20221111 | 2022-11-11 | 1.996 | 1.985 | ML | 2022-11-11 | -0.3296% | -0.7599% | 20221229 | 2022-12-29 | 2 | 2.019 | ML, MatriXX_2 | 2022-12-29 | 0.8522% | 0.4328%
20221116 | 2022-11-16 | 2.006 | 1.997 | YM | 2022-11-16 | 0.1698% | -0.16% | 20221230 | 2022-12-30 | 1.986 | 2.007 | ML, MatriXX_2 | 2022-12-30 | 0.1462% | -0.1642%
20221117 | 2022-11-17 | 1.999 | 1.998 | YM | 2022-11-17 | -0.1798% | -0.11% | 20230102 | 2023-01-02 | 1.98 | 1.994 | ML, MatriXX_2 | 2023-01-02 | -0.1563% | -0.8108%
20221118 | 2022-11-18 | 1.988 | 1.98 | CVD, students informed me about deviating gamma | 2022-11-18 | -0.7291% | -1.01% | 20230110 | 2023-01-10 | 1.974 | 1.996 | ML, MatriXX_2 | 2023-01-10 | -0.4589% | -0.7113%
20221123 | 2022-11-23 | 1.997 | 1.989 | ML | 2022-11-23 | -0.2796% | -0.5599% | 20230113 | 2023-01-13 | 1.976 | 1.996 | ML, MatriXX 2 | 2023-01-13 | -0.358% | -0.7113%
20221125 | 2022-11-25 | 2.004 | 1.997 | CVD | 2022-11-25 | 0.06991% | -0.16% | 20230116 | 2023-01-16 | 1.976 | 1.998 | ML, MatriXX 2 | 2023-01-16 | -0.358% | -0.6118%
20221129 | 2022-11-29 | 2.011 | 2.003 | CVD | 2022-11-29 | 0.4195% | 0.14% | 20230118 | 2023-01-18 | 1.976 | 1.998 | ML, MatriXX 2 | 2023-01-18 | -0.358% | -0.6118%
20221201 | 2022-12-01 | 2.008 | 1.997 | CVD | 2022-12-01 | 0.2696% | -0.16% | 20230124 | 2023-01-24 | 1.981 | 2.002 | ML, MatriXX 2 | 2023-01-24 | -0.1059% | -0.4129%
20221202 | 2022-12-02 | 2.005 | 1.997 | CVD, Matlab R2021b in new Citrix portal | 2022-12-02 | 0.1198% | -0.16% | 20230125 | 2023-01-25 | 1.982 | 2.001 | ML, MatriXX 2 | 2023-01-25 | -0.05547% | -0.4626%
20221206 | 2022-12-06 | 2.003 | 1.991 | ML | 2022-12-06 | 0.01997% | -0.46% | 20230130 | 2023-01-30 | 1.977 | 2 | ML, MatriXX 2 | 2023-01-30 | -0.3076% | -0.5124%
20221207 | 2022-12-07 | 2.021 | 2.012 | ML | 2022-12-07 | 0.9188% | 0.5899% | 20230209 | 2023-02-09 | 1.97 | 1.993 | ML, MatriXX1 | 2023-02-09 | -0.6606% | -0.8606%
20221208 | 2022-12-08 | 2.003 | 1.992 | ML | 2022-12-08 | 0.01997% | -0.41% | 20230210 | 2023-02-10 | 1.972 | 1.991 | ML, MatriXX1 | 2023-02-10 | -0.5597% | -0.9601%
20221209 | 2022-12-09 | 2.011 | 2.007 | ML, From DailyQA, Matrixx 2 | 2022-12-09 | 0.4195% | 0.34% | 20230213 | 2023-02-13 | 1.961 | 1.978 | ML, MatriXX1 | 2023-02-13 | -1.114% | -1.607%
20221220 | 2022-12-20 | 2.011 | 2.002 | AC, Matrixx3 | 2022-12-20 | 0.4195% | 0.08999% | 20230214 | 2023-02-14 | 1.962 | 1.974 | ML, MatriXX1 | 2023-02-14 | -1.064% | -1.806%
20221222 | 2022-12-22 | 2.016 | 2.002 | AC, Matrixx3 | 2022-12-22 | 0.6691% | 0.08999% | 20230216 | 2023-02-16 | 1.971 | 1.988 | ML, MatriXX1 | 2023-02-16 | -0.6102% | -1.109%
20221223 | 2022-12-23 | 2.01 | 1.993 | AC, Matrixx3 | 2022-12-23 | 0.3695% | -0.36% | 20230216 | 2023-02-16 | 1.971 | 1.988 | JR | 2023-02-16 | -0.6102% | -1.109%
20221229 | 2022-12-29 | 2.022 | 2.008 | ML, MatriXX_3 | 2022-12-29 | 0.9687% | 0.39% | 20230220 | 2023-02-20 | 1.968 | 1.984 | ML, MatriXX1 | 2023-02-20 | -0.7614% | -1.308%
20221230 | 2022-12-30 | 2.013 | 2.002 | ML, MatriXX_3 | 2022-12-30 | 0.5193% | 0.08999% | 20230220 | 2023-02-20 | 1.968 | 1.984 | AJ | 2023-02-20 | -0.7614% | -1.308%
20230103 | 2023-01-03 | 2.013 | 2.001 | ML, MatriXX_3 | 2023-01-03 | 0.5193% | 0.04% | 20230221 | 2023-02-21 | 1.956 | 1.977 | AE, Training session, beam not tunned | 2023-02-21 | -1.367% | -1.656%
20230111 | 2023-01-11 | 2.019 | 2.009 | ML, MatriXX_3 | 2023-01-11 | 0.8189% | 0.44% | 20230221 | 2023-02-21 | 1.955 | 1.973 | AR, Training session, beam not tunned | 2023-02-21 | -1.417% | -1.855%
20230113 | 2023-01-13 | 2.018 | 2.006 | ML, MatriXX_3 | 2023-01-13 | 0.769% | 0.29% | 20230224 | 2023-02-24 | 1.957 | 1.975 | OJ | 2023-02-24 | -1.316% | -1.756%
20230117 | 2023-01-17 | 2.004 | 1.994 | ML, MatriXX_3 | 2023-01-17 | 0.06991% | -0.31% | 20230227 | 2023-02-27 | 1.952 | 1.963 | GK,AJ,AR,OJ,EH | 2023-02-27 | -1.568% | -2.353%
20230119 | 2023-01-19 | 2.006 | 1.993 | ML, MatriXX_3 | 2023-01-19 | 0.1698% | -0.36% | 20230228 | 2023-02-28 | 1.965 | 1.984 | GF | 2023-02-28 | -0.9127% | -1.308%
20230127 | 2023-01-27 | 2.018 | 2.003 | ML, MatriXX_3 | 2023-01-27 | 0.769% | 0.14% | 20230301 | 2023-03-01 | 1.966 | 1.97 | SM, AE | 2023-03-01 | -0.8623% | -2.005%
20230131 | 2023-01-31 | 2.005 | 1.994 | YM | 2023-01-31 | 0.1198% | -0.31% | 20230302 | 2023-03-02 | 1.966 | 1.982 | JR, Remeber to record the GI vaules/ML | 2023-03-02 | -0.8623% | -1.408%
20230202 | 2023-02-02 | 2.007 | 1.994 | ML, MatriXX_3 | 2023-02-02 | 0.2197% | -0.31% | 20230303 | 2023-03-03 | 1.964 | 1.97 | SM, AR | 2023-03-03 | -0.9631% | -2.005%
20230203 | 2023-02-03 | 2.007 | 1.995 | ML, MatriXX_3 | 2023-02-03 | 0.2197% | -0.26% | 20230306 | 2023-03-06 | 1.989 | 1.981 | ML | 2023-03-06 | 0.2975% | -1.457%
20230207 | 2023-02-07 | 2.011 | 1.999 | YM | 2023-02-07 | 0.4195% | -0.05999% | 20230306 | 2023-03-06 | 0 | 1.98 | ML, only large cube was repeated | 2023-03-06 | -100% | -1.507%
20230208 | 2023-02-08 | 2.009 | 1.994 | ML, MatriXX_3 | 2023-02-08 | 0.3196% | -0.31% | 20230307 | 2023-03-07 | 1.964 | 1.976 | ML | 2023-03-07 | -0.9631% | -1.706%
20230209 | 2023-02-09 | 2.006 | 1.993 | ML, MatriXX_3 | 2023-02-09 | 0.1698% | -0.36% | 20230307 | 2023-03-07 | 0 | 1.976 | ML, only large cube was repeated | 2023-03-07 | -100% | -1.706%
20230210 | 2023-02-10 | 2.006 | 1.989 | ML, MatriXX_3 | 2023-02-10 | 0.1698% | -0.5599% | 20230308 | 2023-03-08 | 1.963 | 1.97 | ML | 2023-03-08 | -1.014% | -2.005%
20230214 | 2023-02-14 | 2.011 | 2.005 | ML, MatriXX_2 | 2023-02-14 | 0.4195% | 0.24% | 20230308 | 2023-03-08 | 1.962 | 1.978 | ML, Large cube: region approx. 3x3cm (right, down corner of the field) where dose deviates from reference in av. -2.5% | 2023-03-08 | -1.064% | -1.607%
20230215 | 2023-02-15 | 2.008 | 1.997 | ML, MatriXX_1 | 2023-02-15 | 0.2696% | -0.16% | 20230309 | 2023-03-09 | 1.971 | 1.989 | EH | 2023-03-09 | -0.6102% | -1.06%
20230216 | 2023-02-16 | 2.011 | 2 | ML, MatriXX_3 | 2023-02-16 | 0.4195% | -0.009999% | 20230309 | 2023-03-09 | 1.979 | 1.999 | ML | 2023-03-09 | -0.2067% | -0.5621%
20230217 | 2023-02-17 | 2.005 | 2.001 | ML, MatriXX_2 | 2023-02-17 | 0.1198% | 0.04% | 20230310 | 2023-03-10 | 1.97 | 1.983 | AR, GK | 2023-03-10 | -0.6606% | -1.358%
20230221 | 2023-02-21 | 2.001 | 1.994 | ML, MatriXX_2 | 2023-02-21 | -0.0799% | -0.31% | 20230313 | 2023-03-13 | 1.959 | 1.986 | AJ | 2023-03-13 | -1.215% | -1.209%
0 | #VALUE! | 0 | 0 | Intermission _ Matrixx from DailyQA | #VALUE! | -100% | -100% | 20230314 | 2023-03-14 | 1.969 | 1.991 | SM, EH | 2023-03-14 | -0.711% | -0.9601%
20230222 | 2023-02-22 | 1.998 | 1.99 | AR | 2023-02-22 | -0.2297% | -0.5099% | 20230315 | 2023-03-15 | 1.967 | 1.987 | GF | 2023-03-15 | -0.8119% | -1.159%
20230223 | 2023-02-23 | 2 | 1.993 | OJ | 2023-02-23 | -0.1298% | -0.36% | 20230316 | 2023-03-16 | 1.97 | 1.992 | EH | 2023-03-16 | -0.6606% | -0.9103%
20230224 | 2023-02-24 | 1.996 | 1.985 | EH | 2023-02-24 | -0.3296% | -0.7599% | 20230320 | 2023-03-20 | 1.973 | 1.989 | AC | 2023-03-20 | -0.5093% | -1.06%
20230227 | 2023-02-27 | 1.989 | 1.981 | AC | 2023-02-27 | -0.6791% | -0.9599% | 20230320 | 2023-03-20 | 1.968 | 1.987 | AC | 2023-03-20 | -0.7614% | -1.159%
20230228 | 2023-02-28 | 1.999 | 1.994 | AR, AJ | 2023-02-28 | -0.1798% | -0.31% | 20230320 | 2023-03-20 | 1.971 | 1.988 | AC | 2023-03-20 | -0.6102% | -1.109%
20230301 | 2023-03-01 | 2.01 | 2.003 | AR, JR | 2023-03-01 | 0.3695% | 0.14% | 20230321 | 2023-03-21 | 1.974 | 1.989 | GF,AE | 2023-03-21 | -0.4589% | -1.06%
20230302 | 2023-03-02 | 2.003 | 1.994 | SM, SF | 2023-03-02 | 0.01997% | -0.31% | 20230322 | 2023-03-22 | 1.969 | 1.989 | SF | 2023-03-22 | -0.711% | -1.06%
20230303 | 2023-03-03 | 2.006 | 1.996 | OJ | 2023-03-03 | 0.1698% | -0.21% | 20230323 | 2023-03-23 | 1.972 | 1.986 | SM, OJ | 2023-03-23 | -0.5597% | -1.209%
20230306 | 2023-03-06 | 2.003 | 1.994 | AD, GB, run 2, after ion source replacement | 2023-03-06 | 0.01997% | -0.31% | 20230324 | 2023-03-24 | 1.965 | 1.986 | SM, EH | 2023-03-24 | -0.9127% | -1.209%
20230307 | 2023-03-07 | 2.006 | 1.997 | OJ | 2023-03-07 | 0.1698% | -0.16% | 20230327 | 2023-03-27 | 1.963 | 1.982 | AJ, GK | 2023-03-27 | -1.014% | -1.408%
20230308 | 2023-03-08 | 2.003 | 1.993 | JR, After beam alignment | 2023-03-08 | 0.01997% | -0.36% | 20230328 | 2023-03-28 | 1.963 | 1.976 | AE, GF | 2023-03-28 | -1.014% | -1.706%
20230309 | 2023-03-09 | 2.002 | 1.996 | JR | 2023-03-09 | -0.02996% | -0.21% | 20230329 | 2023-03-29 | 1.969 | 1.984 | GK,ON,EH | 2023-03-29 | -0.711% | -1.308%
20230310 | 2023-03-10 | 2.006 | 1.997 | SM, AR | 2023-03-10 | 0.1698% | -0.16% | 20230330 | 2023-03-30 | 1.964 | 1.986 | JR,ON | 2023-03-30 | -0.9631% | -1.209%
20230313 | 2023-03-13 | 2 | 1.991 | SM, SF | 2023-03-13 | -0.1298% | -0.46% | 20230331 | 2023-03-31 | 1.963 | 1.985 | AR, SM | 2023-03-31 | -1.014% | -1.259%
20230314 | 2023-03-14 | 1.999 | 1.991 | AE, AR | 2023-03-14 | -0.1798% | -0.46% | 20230403 | 2023-04-03 | 1.975 | 1.996 | SM, AR | 2023-04-03 | -0.4085% | -0.7113%
20230315 | 2023-03-15 | 1.995 | 1.989 | SM, AE | 2023-03-15 | -0.3795% | -0.5599% | 20230404 | 2023-04-04 | 1.97 | 1.988 | ON | 2023-04-04 | -0.6606% | -1.109%
20230316 | 2023-03-16 | 1.992 | 1.983 | JR | 2023-03-16 | -0.5293% | -0.8599% | 20230405 | 2023-04-05 | 1.973 | 1.993 | AE, OJ | 2023-04-05 | -0.5093% | -0.8606%
20230320 | 2023-03-20 | 1.999 | 1.989 | AC | 2023-03-20 | -0.1798% | -0.5599% | 20230406 | 2023-04-06 | 1.965 | 1.988 | EH | 2023-04-06 | -0.9127% | -1.109%
20230321 | 2023-03-21 | 1.997 | 1.988 | ML, ML corrected the analysis | 2023-03-21 | -0.2796% | -0.6099% | 20230407 | 2023-04-07 | 1.97 | 1.991 | SM, JR | 2023-04-07 | -0.6606% | -0.9601%
20230322 | 2023-03-22 | 2 | 1.989 | GK,SM | 2023-03-22 | -0.1298% | -0.5599% | 20230411 | 2023-04-11 | 1.97 | 1.988 | GF, EH, ON | 2023-04-11 | -0.6606% | -1.109%
20230323 | 2023-03-23 | 2.003 | 1.994 | EH | 2023-03-23 | 0.01997% | -0.31% | 20230412 | 2023-04-12 | 1.972 | 1.991 | JR,GF | 2023-04-12 | -0.5597% | -0.9601%
20230324 | 2023-03-24 | 2.005 | 1.997 | AE | 2023-03-24 | 0.1198% | -0.16% | 20230413 | 2023-04-13 | 1.968 | 1.99 | SM, AE | 2023-04-13 | -0.7614% | -1.01%
20230327 | 2023-03-27 | 1.991 | 1.983 | SM, AR | 2023-03-27 | -0.5792% | -0.8599% | 20230414 | 2023-04-14 | 1.963 | 1.986 | OJ, SM | 2023-04-14 | -1.014% | -1.209%
20230328 | 2023-03-28 | 2.003 | 1.991 | GK,ON,SF | 2023-03-28 | 0.01997% | -0.46% | 20230417 | 2023-04-17 | 1.971 | 1.99 | AJ, OJ | 2023-04-17 | -0.6102% | -1.01%
20230329 | 2023-03-29 | 2.001 | 1.989 | JR | 2023-03-29 | -0.0799% | -0.5599% | 20230418 | 2023-04-18 | 1.968 | 1.991 | GK,GF | 2023-04-18 | -0.7614% | -0.9601%
20230330 | 2023-03-30 | 2.001 | 1.992 | SM, OJ | 2023-03-30 | -0.0799% | -0.41% | 20230419 | 2023-04-19 | 1.966 | 1.989 | OJ, AR | 2023-04-19 | -0.8623% | -1.06%
20230331 | 2023-03-31 | 2.002 | 1.991 | ON |  |  |  | 20230420 | 2023-04-20 | 1.971 | 1.993 | ON, EH | 2023-04-20 | -0.6102% | -0.8606%
20230403 | 2023-04-03 | 2.004 | 1.994 | OJ |  |  |  | 20230421 | 2023-04-21 | 1.967 | 1.988 | SM, ON | 2023-04-21 | -0.8119% | -1.109%
20230404 | 2023-04-04 | 2 | 1.992 | AE, GF |  |  |  | 20230424 | 2023-04-24 | 1.966 | 1.987 | GK,AE | 2023-04-24 | -0.8623% | -1.159%
20230405 | 2023-04-05 | 2.004 | 1.995 | ON, SM, JR |  |  |  | 20230425 | 2023-04-25 | 1.964 | 1.988 | AE, GF | 2023-04-25 | -0.9631% | -1.109%
20230406 | 2023-04-06 | 2.005 | 1.996 | OJ |  |  |  | 20230426 | 2023-04-26 | 1.992 | 2.001 | AR, GK | 2023-04-26 | 0.4488% | -0.4626%
20230407 | 2023-04-07 | 2.008 | 2 | JR,SM |  |  |  | 20230427 | 2023-04-27 | 1.98 | 1.998 | GK,OJ | 2023-04-27 | -0.1563% | -0.6118%
20230411 | 2023-04-11 | 2.004 | 1.992 | OJ, GK |  |  |  | 20230428 | 2023-04-28 | 1.969 | 1.992 | GK,JR | 2023-04-28 | -0.711% | -0.9103%
20230412 | 2023-04-12 | 1.996 | 1.986 | GK,EH |  |  |  | 20230502 | 2023-05-02 | 1.971 | 1.993 | AE, AR, EF | 2023-05-02 | -0.6102% | -0.8606%
20230413 | 2023-04-13 | 1.999 | 1.989 | AR |  |  |  | 20230503 | 2023-05-03 | 1.97 | 1.99 | ON | 2023-05-03 | -0.6606% | -1.01%
20230414 | 2023-04-14 | 1.994 | 1.985 | AR, JR |  |  |  | 20230504 | 2023-05-04 | 1.972 | 1.995 | JR | 2023-05-04 | -0.5597% | -0.7611%
20230417 | 2023-04-17 | 2.003 | 1.992 | GK,SM |  |  |  | 20230504 | 2023-05-04 | 1.975 | 1.995 | ON, EF | 2023-05-04 | -0.4085% | -0.7611%
20230418 | 2023-04-18 | 2.004 | 1.994 | AJ |  |  |  | 20230505 | 2023-05-05 | 1.971 | 1.992 | SM, GK | 2023-05-05 | -0.6102% | -0.9103%
20230419 | 2023-04-19 | 2.003 | 1.993 | JR |  |  |  | 20230508 | 2023-05-08 | 1.97 | 1.988 | OJ, SF | 2023-05-08 | -0.6606% | -1.109%
20230420 | 2023-04-20 | 2 | 1.992 | JR,AR |  |  |  | 20230509 | 2023-05-09 | 1.967 | 1.983 | GF | 2023-05-09 | -0.8119% | -1.358%
20230421 | 2023-04-21 | 2.006 | 1.993 | GK,AR |  |  |  | 20230510 | 2023-05-10 | 1.968 | 1.987 | ON | 2023-05-10 | -0.7614% | -1.159%
20230424 | 2023-04-24 | 1.996 | 1.989 | AR, AJ |  |  |  | 20230511 | 2023-05-11 | 1.973 | 1.991 | AR | 2023-05-11 | -0.5093% | -0.9601%
20230425 | 2023-04-25 | 2.001 | 1.989 | OJ, AR
20230426 | 2023-04-26 | 2.012 | 2.002 | ON, SM
20230427 | 2023-04-27 | 2.001 | 1.992
Formulas:
Date/Plan ID = [Reference]

## Updates
No | Date | Responsible | Change Description
1 | 2020-01-13 | EA, LS | penumbra is calculated for 20%-80% instead of 10%-90%
2 | 2020-01-13 | LS | Matrix is calibrated in D [Gy]
3 | 2020-02-13 | LS | new formatting with colours for limits for D, P and FWHM. Limits were set according to AAPM 224 report. For dose ref is TPS, for P and FWHM limit is 1 and 2 mm the same for x and y, left and right. Ref. value for P and FWHM is an average for Jan-Feb 2020
4 | 2020-02-20 | EA, LS | Updated definition of symmetry and flatness to be in accordance with AAPM 224. Also changed so that the flat part is defined as FWHM - 3 penumbra. Prevously it was FWHM-2 penumbra
5 | 2020-05-08 | LS | All results were recalculated with a new MatlabScript. Reference values were updated
6 | 2022-07-26 | ML | Reference values were updated