## GTR1_Setup check
ID.nr:
Titel: | [30 empty] | Dokumenttyp:
Protocol for Daily QA: Setup preparation fast check | [30 empty] | Strålskyddsinstruktion
Godkänt av: | [29 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [29 empty] | Skapat den:
Reviderat av: | [29 empty] | Reviderat den:
GTR1
Setup check: GTR1 MatriXX PT_2
Set-up: |  | 2.3 cm SW + MatrixPT position adjusted using X-rays (chambers at iso)
Detector |  | IBA MatrixPT_2
Phantom |  | SW
MyQA Project: |  | GTR1_DailyQA_Matrixx
Patient in AdaptDelivery |  | G1_Daily_MatriXX Skandion
Plan ID: |  | Single layer_E150MeV_F10x10
Detector Calib. Factor: |  | 85.956 | GTR1_Reference Calib. 2022-11-04
Reference dose [Gy] |  | 1.716 | (Calibration dose)
Dose from TPS[Gy] |  | 1.718
Good: within 0.5%
1.6988 | 1.7332
1.7008 | 1.7352
Warning: within 2%
1.7342 | 1.6978
1.7362 | 1.6998
Warning: over 2%
1.7513 | 1.6807
1.7534 | 1.6826
|  | Conditions in room |  |  | Detector condition | [6 empty] | IBA laser off CV (mm) | [6 empty] | Table position saved in Physic QA
Date | Responsible | p [hPa] | Troom [0C] | Troom [K] | Calibration Factor | TMatrixx [0C] | pMatriXX [hPa] | MyQA TMatrixx [0C] | MyQA pMatriXX [hPa] | MyQA Offset TMatrixx [0C] | MyQA Offset pMatriXX [hPa] | Δx (cm) | Δy (cm) | Δz (cm) | Rot (o) | Pitcht (o) | Roll (o) | CV (cm) | X (cm) | Y (cm) | Z (cm) | Rot (o) | Pitcht (o) | Roll (o) | CV (mm) | Z (cm) after applying Offset -3 cm | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | GI (1%, 1 mm) | GI (2%, 2 mm) | Notes
2022-11-04 | ML.AC | 1008.2 | 23.2 | 296.35 | 85.956 | 22.8 | 1008.2 | 25.6 | 1009 | -2.8 | -0.8 | [15 empty] | 1.714 | -0.1166% | -0.2328% |  |  | MatriXX rotated 270deg_run 5
2022-11-11 | ML,AC | 1004.6 | 23.3 | 296.45 | 85.956 | 23.2 | 1004.6 | 25.2 | 1005 | -2 | -0.4 | [15 empty] | 1.714 | -0.1166% | -0.2328% |  |  | REFERENCE IMAGE_run 2
2022-11-11 | ML,AC | 1004.6 | 23.3 | 296.45 | 85.956 | 23.2 | 1004.6 | 25.2 | 1005 | -2 | -0.4 | [15 empty] | 1.714 | -0.1166% | -0.2328% | 100 |  | run 3
2022-11-11 | ML,AC | 1004.6 | 23.3 | 296.45 | 85.956 | 23.2 | 1004.6 | 25.2 | 1005 | -2 | -0.4 | [15 empty] | 1.713 | -0.1748% | -0.291% | 100 |  | run 4
2022-12-09 | ML | 1003.1 | 23.9 | 297.05 | 85.956 | 22.2 | 1003.1 | 24.6 | 1004 | -2.4 | -0.9 | -0.03 | 0.03 | 0.03 | 0 | 0 | 0 | 0.051962 | 0 | 0 | 0.3 | 0 | 0 | 0 | 0.3 | -2.7 | 1.72 | 0.2331% | 0.1164% | 100 |  | run 01
2023-02-14 | AJ | 1025 | 23.3 | 296.45 | 85.596 | 22.5 | 1025 | 23.7 | 1025 | -1.2 | 0 | 0.01 | 0 | 0.03 | 0 | 0 | 0 | 0.031623 | 0 | 0.01 | 0.23 | 0 | 0 | 0 | 0.23022 | -2.77 | 1.721 | 0.2914% | 0.1746% | 94.9 |  | run#1
2023-02-17 | SR, EH, OJ | 1001 | 23.4 | 296.55 | 85.956 | 22.7 | 1001 | 23.5 | 1001 | -0.8 | 0 | -0.01 | 0.03 | 0 | 0 | 0 | 0 | 0.031623 | -0.01 | 0.03 | 0.27 | 0 | 0 | 0 | 0.27185 | -2.73 | 1.718 | 0.1166% | 0% | 100 |  | run 01
2023-02-21 | SM, GF, OJ, AR | 992.6 | 23.2 | 296.35 | 85.956 | 22.2 | 992.6 | 23.5 | 993 | -1.3 | -0.4 | 0.01 | -0.01 | 0.03 | 0 | 0 | 0 | 0.033166 | 0 | 0 | 0.2 | 0 | 0 | 0 | 0.2 | -2.8 | 1.7152 | -0.04662% | -0.163% | 100 |  | run 01
2023-02-22 | AR | 1016 | 22.7 | 295.85 | 85.956 | 22.2 | 1016 | 23.4 | 1016 | -1.2 | 0 | 0 | -0.01 | -0.02 | 0 | 0 | 0 | 0.022361 | 0 | 0 | 0.27 | 0 | 0 | 0 | 0.27 | -2.73 | 1.712 | -0.2331% | -0.3492% | 100
2023-02-23 | OJ, JR, AR | 1008.5 | 23.1 | 296.25 | 85.956 | 22.2 | 1008.5 | 23.3 | 1008 | -1.1 | 0.5 | 0.02 | 0.02 | -0.05 | 0 | 0 | 0 | 0.057446 | 0.02 | 0.02 | 0.21 | 0 | 0 | 0 | 0.2119 | -2.79 | 1.7098 | -0.3613% | -0.4773% | 100
2023-02-24 | EH | 993.9 | 22.9 | 296.05 | 85.956 | 22.2 | 993.9 | 23.4 | 994 | -1.2 | -0.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.27 | 0 | 0 | 0 | 0.27 | -2.73 | 1.7098 | -0.3596% | -0.4756% | 100 |  | run 01
2023-02-27 | GK,EH,AJ | 1020.5 | 23.1 | 296.25 | 85.956 | 23.1 | 1020.5 | 24.7 | 1021 | -1.6 | -0.5 | -0.04 | 0.01 | 0.02 | 0 | 0 | 0 | 0.045826 | 0 | 0 | 0.26 | 0 | 0 | 0 | 0.26 | -2.74 | 1.7072 | -0.5128% | -0.6286% | 89.6 | 98.3 | run 01
2023-02-28 | AR,AJ | 1021.8 | 23.1 | 296.25 | 85.956 | 23.2 | 1021.8 | 23.4 | 1022 | -0.2 | -0.2 | -0.07 | -0.01 | 0 | 0.1 | 0 | 0 | 0.070711 | -0.07 | -0.01 | 0.23 | 0.1 | 0 | 0 | 0.24062 | -2.77 | 1.724 | 0.4662% | 0.3492% | 100 |  | run 01
2023-02-28 | CVD, DJ | 1020.8 | 23.4 | 296.55 | 85.956 | 23.8 | 1020.8 | 23.8 | 1020.8 | 0 | 0 | -0.13 | 0.13 | -0.16 | 0 | 0 | 0 | 0.24372 | [8 empty] | 1.7229 | 0.4021% | 0.2852% | 98.3 | 100 | run 02 /Skandion physicists
2023-03-01 | AR,JR | 1022.1 | 23.1 | 296.25 | 85.956 | 23.4 | 1022.1 | 23 | 1022 | 0.4 | 0.1 | -0.01 | -0.01 | -0.1 | 0 | 0 | 0 | 0.101 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -2.73 | 1.7206 | 0.2652% | 0.1484% | 100 |  | run01 Why the table position has 0 everywhere? Please, pay an attention to correctly recording this position./ML
2023-03-02 | SM, SF | 1017.5 | 22.8 | 295.95 | 85.956 | 22.9 | 1017.5 | 23.8 | 1017 | -0.9 | 0.5 | 0 | 0 | 0 | -0.2 | 0 | 0 | 0 | -0.09 | 0.07 | 0.27 | 359.8 | -0.1 | 0 | 0.29309 | -2.73 | 1.717 | 0.05828% | -0.05821% | 100
2023-03-03 | OJ, SF | 1009.1 | 23.1 | 296.25 | 85.956 | 23.1 | 1009.1 | 23.5 | 1009 | -0.4 | 0.1 | -0.07 | 0.03 | -0.08 | 0 | 0 | 0 | 0.11045 | -0.07 | 0.03 | 0.27 | 0 | 0 | 0 | 0.28054 | -2.73 | 1.717 | 0.05828% | -0.05821% | 100
2023-03-06 | AJ, OJ | 993.2 | 23.2 | 296.35 | 85.956 | 22.9 | 993.2 | 24.3 | 993 | -1.4 | 0.2 | -0.02 | 0.04 | 0.02 | -0.1 | 0 | 0 | 0.04899 | -0.02 | 0.04 | 0.28 | 359.9 | 0 | 0 | 0.28355 | -2.72 | 1.7226 | 0.3846% | 0.2678% | 100 |  | run01 Matrixx got an error, after restarting Matrix it worked again.
2023-03-06 | AD, GB | 994.2 | 23.2 | 296.35 | 85.956 | 23.9 | 994.2 | 25.3 | 994 | -1.4 | 0.2 | -0.02 | 0.04 |  | -0.1 | 0 | 0 | 0.044721 | -0.02 | 0.04 | 0.28 | 359.9 | 0 | 0 |  | -2.72 | 1.7141 | -0.1107% | -0.227% | 100 |  | run02 Measurement after replacement of ion source (3 h later)
2023-03-07 | OJ, GF | 990.7 | 23 | 296.15 | 85.956 | 23.3 | 990.7 | 23.6 | 991 | -0.3 | -0.3 | 0.01 | 0 | 0.03 | -0.1 | 0 | 0 | 0.031623 | 0.01 | 0 | 0.27 | 359.9 | 0 | 0 |  | -2.73 | 1.718 | 0.1166% | 0% | 100
2023-03-08 | SM YM | 979.7 | 23.1 | 296.25 | 85.956 | 23.2 | 979.7 | 23.6 | 981 | -0.4 | -1.3 | 0.01 | 0 | 0.02 | 0 | 0 | 0 | 0.022361 | 0 | 0 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.714 | -0.1166% | -0.2328% | 99.6 | 100
2023-03-08 | JR YM | 979.7 | 23.1 | 296.25 | 85.956 | 23.2 | 979.7 | 23.6 | 981 | -0.4 | -1.3 | [14 empty] | -2.76 | 1.7144 | -0.09266% | -0.209% | 100 | 100 | After the beam alignment
2023-03-09 | JR YM | 992.2 | 22.8 | 295.95 | 85.956 | 23.2 | 992.2 | 23.3 | 993 | -0.1 | -0.8 | 0.02 | -0.02 | 0 | -0.1 | 0 | 0 | 0.028284 | 0.02 | -0.02 | 0.28 | -0.01 | 0 | 0 |  | -2.72 | 1.7155 | -0.02797% | -0.1444% | 100 | 100
2023-03-10 | SM, AR | 994.3 | 22.7 | 295.85 | 85.956 | 23 | 994.3 | 23.1 | 995 | -0.1 | -0.7 | 0.03 | -0.01 | 0 | -0.1 | 0 | 0 | 0.031623 | 0.03 | -0.01 | 0.29 | 359.9 | 0 | 0 |  | -2.71 | 1.7183 | 0.134% | 0.01746% | 100 | 100
2023-03-13 | SM, SF | 994.4 | 22.8 | 295.95 | 85.956 | 22.8 | 994.4 | 23.2 | 995 | -0.4 | -0.6 | 0.01 | 0 | 0 | -0.1 | 0 | 0 | 0.01 | 0.01 | 0 | 0.28 | 359.9 | 0 | 0 |  | -2.72 | 1.7122 | -0.2244% | -0.3405% | 96.7 | 100
2023-03-14 | AE, AR | 968.1 | 22.9 | 296.05 | 85.956 | 22.5 | 968.1 | 23.3 | 969 | -0.8 | -0.9 | -0.03 | 0 | 0 | 0 | 0 | 0 | 0.03 | -0.03 | 0.01 | 0.28 | 0 | 0 | 0 |  | -2.72 | 1.7131 | -0.1684% | -0.2846% | 100 | 100
2023-03-15 | SM, AE | 989.5 | 23.2 | 296.35 | 85.956 | 22.9 | 989.5 | 23.2 | 990 | -0.3 | -0.5 | -0.02 | -0.01 | 0.02 | 0 | 0 | 0 | 0.03 | -0.02 | -0.01 | 0.27 | 0 | 0 | 0 |  | -2.73 | 1.711 | -0.2914% | -0.4075% | 95.8 | 100
2023-03-16 | JR YM | 1010.1 | 22.5 | 295.65 | 85.956 | 23.2 | 1010.1 | 22.9 | 1011 | 0.3 | -0.9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.01 | 0.31 | 0 | 0 | 0 |  | -2.69 | 1.709 | -0.4079% | -0.5239% | 100 | 100 | in DM T=296.6 K not 295.6 K //ML
2023-03-20 | AC,ML | 1002.6 | 23.7 | 296.85 | 85.956 | 23.4 | 1002.6 | 24.5 | 1003 | -1.1 | -0.4 | 0.02 | -0.05 | 0.03 | 0 | 0 | 0 | 0.061644 | 0.02 | -0.05 | 0.29 | 0 | 0 | 0 |  | -2.71 | 1.7129 | -0.1818% | -0.298% | 100
2023-03-21 | GK,AJ | 1007.3 | 22.9 | 296.05 | 85.956 | 22.8 | 1007.3 | 23.4 | 1008 | -0.6 | -0.7 | -0.01 | -0.01 | 0 | 0 | 0 | 0 | 0.014142 | 0 | 0.01 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.7098 | -0.3631% | -0.479% | 100
2023-03-22 | GK,SM | 996.3 | 22.7 | 295.85 | 85.956 | 22.1 | 997.6 | 23.1 | 997 | -1 | 0.6 | 0.01 | 0.01 | 0.03 | 0 | 0 | 0 | 0.033166 | 0.01 | 0.01 | 0.3 | 0 | 0 | 0 |  | -2.7 | 1.7125 | -0.2045% | -0.3207% | 98.7 | 100
2023-03-23 | EH,GK | 998.2 | 22.8 | 295.95 | 85.956 | 23.1 | 988.2 | 23.2 | 989 | -0.1 | -0.8 | -0.01 | 0 | 0.02 | -0.1 | 0 | 0 | 0.022361 | 0 | 0 | 0.27 | 0 | 0 | 0 |  | -2.73 | 1.7182 | 0.13% | 0.01339% | 100
2023-03-24 | AE, AR | 988.7 | 22.9 | 296.05 | 85.956 | 23.2 | 988.7 | 23 | 989 | 0.2 | -0.3 | 0 | 0 | 0.03 | 0 | 0 | 0 | 0.03 | 0 | 0 | 0.26 | 0 | 0 | 0 |  | -2.74 | 1.7177 | 0.1002% | -0.0163% | 97.9 | 100
2023-03-27 | SM, AR | 997.6 | 23.2 | 296.35 | 85.956 | 22.2 | 997.6 | 23 | 998 | -0.8 | -0.4 | 0 | -0.01 | -0.02 | 0 | 0 | 0 | 0.022361 | 0 | 0.01 | 0.3 | 0 | 0 | 0 |  | -2.69 | 1.7076 | -0.4895% | -0.6054% | 97.9 | 100
2023-03-28 | GK,ON,SF | 1004.9 | 23 | 296.15 | 85.956 | 23 | 1004.9 | 23.8 | 1006 | -0.8 | -1.1 | -0.12 | 0 | -0.1 | -0.1 | 0 | 0 | 0.1562 | -0.12 | 0 | 0.23 | 359.9 | 0 | 0 |  | -2.77 | 1.7154 | -0.03671% | -0.1531% | 100 | 100
2023-03-29 | JR,AE | 1012.2 | 23.3 | 296.45 | 85.956 | 23.2 | 1012.2 | 23.2 | 1013 | 0 | -0.8 | -0.03 | 0 | -0.05 | 0 | 0 | 0 | 0.05831 | -0.03 | 0 | 0.27 | 0 | 0 | 0 |  | -2.73 | 1.7126 | -0.197% | -0.3132% | 100 | 100
2023-03-30 | SM, OJ | 1009.1 | 23 | 296.15 | 85.956 | 22.5 | 1009.1 | 22.9 | 1009 | -0.4 | 0.1 | -0.06 | 0.01 | 0 | -0.1 | 0 | 0 | 0.060828 | -0.06 | 0 | 0.2 | 359.9 | 0 | 0 |  | -2.8 | 1.7152 | -0.04779% | -0.1641% | 100 | 100
2023-03-31 | ON, EH, OJ | 1001.5 | 22.9 | 296.05 | 85.956 | 22.9 | 1001.5 | 23.3 | 1002 | -0.4 | -0.5 | -0.03 | -0.02 | 0.02 | 0 | 0 | 0 | 0.041231 | -0.02 | -0.01 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.714 | -0.1166% | -0.2328% | 99.2 | 100
2023-04-03 | OJ, ON, AJ | 1032.2 | 22.9 | 296.05 | 85.956 | 22.9 | 1032.2 | 23.3 | 1033 | -0.4 | -0.8 | -0.04 | 0 | 0.02 | 0 | 0 | 0 | 0.044721 | -0.04 | 0 | 0.21 | 0 | 0 | 0 |  | -2.79 | 1.715 | -0.05828% | -0.1746% | 94.5 | 100
2023-04-04 | AE, GF | 1024.7 | 22.8 | 295.95 | 85.956 | 22.7 | 1024.7 | 23.2 | 1025 | -0.5 | -0.3 | -0.05 | -0.01 | 0 | -0.1 | 0 | 0 | 0.05099 | -0.05 | -0.01 | 0.22 | 359.9 | 0 | 0 |  | -2.78 | 1.7151 | -0.05303% | -0.1694% | 100 | 100
2023-04-05 | ON, SM, JR | 1021.3 | 22.8 | 295.95 | 85.956 | 22.7 | 1021.3 | 23.3 | 1022 | -0.6 | -0.7 | 0.01 | 0.01 | 0 | 0 | 0 | 0 | 0.014142 | 0 | 0.01 | 0.17 | 0 | 0 | 0 |  | -2.83 | 1.7155 | -0.02914% | -0.1455% | 100 | 100
2023-04-06 | OJ, ON, JR | 1022.2 | 22.9 | 296.05 | 85.956 | 23.2 | 1022.2 | 22.7 | 1023 | 0.5 | -0.8 | 0.01 | 0.01 | 0.02 | 0 | 0 | 0 | 0.024495 | 0.04 | -0.01 | 0.19 | 0 | 0 | 0 |  | -2.82 | 1.719 | 0.1748% | 0.05821% | 100
2023-04-07 | JR,SM | 1023.1 | 22.7 | 295.85 | 85.956 | 22.7 | 1023.1 | 23 | 1023 | -0.3 | 0.1 | 0 | 0 | -0.02 | 0 | 0 | 0 | 0.02 | 0 | 0 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.7192 | 0.1882% | 0.07159% | 100 | 100 | ok/CVD
2023-04-11 | OJ, GK | 1010.4 | 22.9 | 296.05 | 85.956 | 22.9 | 1010.4 | 23.8 | 1011 | -0.9 | -0.6 | -0.02 | 0.06 | -0.03 | 0 | 0 | 0 | 0.07 | -0.03 | 0.1 | 0.2 | 0 | 0 | 0 |  | -2.8 | 1.717 | 0.05828% | -0.05821% | 100
2023-04-12 | GK,EH | 1005.6 | 22.7 | 295.85 | 85.956 | 22.9 | 1005.6 | 23.6 | 1006 | -0.7 | -0.4 | 0.05 | 0.03 | -0.02 | 0 | 0 | 0 | 0.061644 | 0.05 | 0.03 | 0.21 | 0 | 0 | 0 |  | -2.79 | 1.7137 | -0.134% | -0.2503% | 90.1 | 100 | ok/YM
2023-04-13 | AR | 1005.8 | 22.6 | 295.75 | 85.956 | 22.8 | 1005.8 | 23.7 | 1007 | -0.9 | -1.2 | -0.12 | 0.04 | 0 | -0.1 | 0 | 0 | 0.12649 | -0.12 | 0.04 | 0.19 | 359.9 | 0 | 0 |  | -2.81 | 1.7116 | -0.2587% | -0.3749% | 100 | 100 | OBS IBA laser CV out of tolerans./YM
2023-04-14 | AR | 1014.2 | 22.9 | 296.05 | 85.956 | 22.4 | 1014.2 | 23.1 | 1015 | -0.7 | -0.8 | -0.03 | -0.02 | 0 | 0 | 0 | 0 | 0.036056 | 0.03 | 0.01 | 0.19 | 0 | 0 | 0 |  | -2.81 | 1.7078 | -0.4779% | -0.5937% | 100 |  | output calibration older than 3 months
2023-04-17 | GK, SM | 1025.8 | 22.6 | 295.75 | 85.956 | 22.9 | 1025.8 | 22.9 | 1025 | 0 | 0.8 | 0.03 | 0.04 | 0 | 0 | 0 | 0 | 0.05 | 0.03 | 0.04 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.714 | -0.1166% | -0.2328% | 100 | 100
2023-04-18 | AJ | 1032.7 | 22.6 | 295.75 | 85.956 | 22.8 | 1032.7 | 23.6 | 1033 | -0.8 | -0.3 | -0.01 | -0.01 | 0.03 | 0 | 0 | 0 | 0.033166 | -0.01 | 0 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.718 | 0.1166% | 0% | 100
2023-04-19 | JR | 1033.3 | 22.8 | 295.95 | 85.956 | 23.1 | 1033.3 | 23.3 | 1034 | -0.2 | -0.7 | 0 | 0 | -0.01 | 0 | 0 | 0 | 0.01 | 0 | 0.01 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.7153 | -0.04254% | -0.1589% | 100 | 100
2023-04-20 | JR,AR | 1029.4 | 22.8 | 295.95 | 85.956 | 23.1 | 1029.4 | 22.8 | 1029 | 0.3 | 0.4 | -0.03 | 0 | 0 | 0 | 0 | 0 | 0.03 | -0.03 | 0 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.7142 | -0.1026% | -0.2189% | 100
2023-04-21 | GK,AR | 1022.2 | 22 | 295.15 | 85.956 | 22.9 | 1022.2 | 23.2 | 1022 | -0.3 | 0.2 | -0.03 | -0.01 | 0.03 | 0.1 | 0 | 0 | 0.043589 | -0.03 | -0.01 | 0.21 | 359.9 | 0 | 0 |  | -2.79 | 1.715 | -0.05828% | -0.1746% | 92.9 | 100 | in DM T=295.8 K
2023-04-24 | AR, AJ | 999.9 | 22.7 | 295.85 | 85.956 | 22.9 | 999.9 | 23.1 | 1000 | -0.2 | -0.1 | -0.03 | 0 | 0 | -0.1 | 0 | 0 | 0.03 | -0.03 | 0 | 0.21 | 359.9 | 0 | 0 |  | -2.79 | 1.71 | -0.3497% | -0.4657% | 100
2023-04-25 | OJ, AR | 995.8 | 22.6 | 295.75 | 85.956 | 22.9 | 995.8 | 24 | 996 | -1.1 | -0.2 | -0.06 | -0.01 | 0 | 0 | 0 | 0 | 0.060828 | -0.15 | 0.08 | 0.2 | 0 | -0.1 | 0 |  | -2.8 | 1.716 | 0% | -0.1164% | 100
2023-04-26 | ON, SM | 995.5 | 22.8 | 295.95 | 85.956 | 22.9 | 995.5 | 23.1 | 995 | -0.2 | 0.5 | 0.01 | 0.01 | 0 | 0 | 0 | 0 | 0.014142 | 0 | 0 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.721 | 0.2914% | 0.1746% | 99.6 | 100
2023-04-27 | SM, JR | 1000.8 | 22.7 | 295.85 | 85.956 | 23.1 | 1000.8 | 23.1 | 1001 | 0 | -0.2 | 0 | 0.01 | 0.02 | 0 | 0 | 0 | 0.022361 | 0 | 0.01 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.7131 | -0.1707% | -0.287% | 100
2023-04-28 | OJ, SM | 1007.8 | 22.7 | 295.85 | 85.956 | 22.8 | 1007.8 | 23.1 | 1008 | -0.3 | -0.2 | 0.01 | 0 | 0 | -0.1 | 0 | 0 | 0.01 | 0 | 0 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.718 | 0.1166% | 0% | 100
2023-05-02 | AJ, EN | 1002.4 | 22.8 | 295.95 | 85.956 | 23 | 1002.4 | 24.2 | 1003 | -1.2 | -0.6 | 0.02 | 0.01 | 0.07 | -0.2 | 0 | 0 | 0.073485 | 0.02 | 0.01 | 0.23 | 359.8 | 0 | 0 |  | -2.77 | 1.7123 | -0.2179% | -0.3341% | 100
2023-05-03 | JR | 1015.2 | 23.2 | 296.35 | 85.956 | 22.9 | 1015.2 | 23.3 | 1015 | -0.4 | 0.2 | 0 | -0.01 | -0.05 | 0 | 0 | 0 | 0.05099 | 0 | -0.01 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.7098 | -0.3613% | -0.4773% | 100 | 100
2023-05-04 | GK,AR,EF | 1023.4 | 22.8 | 295.95 | 85.956 | 22.4 | 1023.4 | 22.9 | 1024 | -0.5 | -0.6 | 0.04 | 0.07 | 0 | -0.1 | 0 | 0 | 0.080623 | -0.02 | 0.1 | 0.24 | 359.9 | 0 | 0 |  | -2.76 | 1.7161 | 0.002914% | -0.1135% | 100 | 100
2023-05-05 | OJ, EN, EK | 1028 | 22.8 | 295.95 | 85.956 | 23 | 1028 | 23.4 | 1028 | -0.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -0.14 | 0.08 | 0.21 | 0 | -0.1 | 0 |  | -2.79 | 1.721 | 0.2914% | 0.1746% | 100 |  | ok/AC
2023-05-08 | SM, AJ | 1025.4 | 22.9 | 296.05 | 85.956 | 22.8 | 1025.4 | 22.8 | 1025 | 0 | 0.4 | -0.02 | 0.02 | 0.02 | -0.1 | 0 | 0 | 0.034641 | -0.02 | 0.02 | 0.21 | 359.9 | 0 | 0 |  | -2.79 | 1.716 | 0% | -0.1164% | 100 | 100 | ok/AC
2023-05-09 | EN,GF,OJ | 1020 | 24.2 | 297.35 | 85.956 | 23.4 | 1020 | 23.7 | 1020 | -0.3 | 0 | 0.02 | 0.01 | -0.02 | 0.1 | 0 | 0 | 0.03 | 0.02 | 0 | 0.23 | 0.1 | 0 | 0 |  | -2.78 | 1.712 | -0.2331% | -0.3492% | 100 | 100
2023-05-10 | AR, EN, JR | 1015.8 | 23.1 | 296.25 | 85.956 | 23.4 | 1015.8 | 23.9 | 1016 | -0.5 | -0.2 | -0.03 | -0.04 | 0 | 0 | 0 | 0 | 0.05 | -0.03 | -0.04 | 0.21 | 0 | 0 | 0 |  | -2.79 | 1.716 | 0% | -0.1164% | 94.1 | 100
2023-05-11 | JR | 1017.7 | 23.1 | 296.25 | 85.956 | 23.1 | 1017.7 | 23.4 | 1018 | -0.3 | -0.3 | 0.04 | 0.03 | 0 | 0 | 0 | 0 | 0.05 | 0.04 | 0.02 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.7177 | 0.09732% | -0.01921% | 89.4 | 100 | ok/CVD
2023-05-12 | EN,GK,ON | 1021.4 | 22.8 | 295.95 | 85.956 | 22.7 | 1021.4 | 23.3 | 1022 | -0.6 | -0.6 | 0.01 | -0.03 | 0.07 | 0 | 0 | 0 | 0.076811 | 0.01 | -0.02 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.716 | 0% | -0.1164% | 98.3 | 100 | ok/EAL
2023-05-15 | AR, EM, | 1013.5 | 22.6 | 295.75 | 85.956 | 23 | 1013.5 | 23.2 | 1013 | -0.2 | 0.5 | -0.02 | 0.01 | -0.02 | 0.2 | 0 | 0 | 0.03 | -0.02 | 0.01 | 0.22 | 0.2 | 0 | 0 |  | -2.77 | 1.709 | -0.4079% | -0.5239% | 100 | 100 | OK/GB
2023-05-16 | GF, AE, EF | 1003.3 | 23.1 | 296.25 | 85.956 | 23 | 1003.3 | 23.3 | 1003 | -0.3 | 0.3 | -0.02 | 0.01 | -0.03 | 0 | 0 | 0 | 0.037417 | -0.02 | 0.01 | 0.18 | 0 | 0 | 0 |  | -2.82 | 1.715 | -0.05828% | -0.1746% | 100 | 100 | OK/ON YM
2023-05-17 | SM, GK | 997.7 | 23.2 | 296.35 | 85.956 | 22.8 | 997.7 | 22.8 | 998 | 0 | -0.3 | -0.01 | 0 | 0.03 | 0 | 0 | 0 | 0.031623 | 0.03 | 0.02 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.718 | 0.1166% | 0% | 100 | 100 | ok/YM ON
2023-05-19 | EF, AJ, OJ | 1025.2 | 22.8 | 295.95 | 85.956 | 23 | 1025.2 | 23.3 | 1025 | -0.3 | 0.2 | 0 | 0.04 | 0 | 0 | 0 | 0 | 0.04 | 0 | 0.04 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.72 | 0.2331% | 0.1164% | 100 |  | OK/AD
2023-05-22 | AJ, OJ | 1020.8 | 22.8 | 295.95 | 85.956 | 22.5 | 1020.8 | 23.2 | 1020 | -0.7 | 0.8 | -0.01 | 0.04 | 0 | -0.1 | 0 | 0 | 0.041231 | -0.01 | 0.04 | 0.23 | 359.9 | 0 | 0 |  | -2.77 | 1.712 | -0.2331% | -0.3492% | 99.6 |  | ok/GB
2023-05-23 | AE, EF | 1010.6 | 22.6 | 295.75 | 85.956 | 22.7 | 1010.6 | 23.5 | 1010 | -0.8 | 0.6 | 0 | -0.06 | 0.03 | 0 | 0 | 0 | 0.067082 | -0.09 | 0.04 | 0.31 | 0 | -0.1 | 0 |  | -2.69 | 1.714 | -0.1166% | -0.2328% | 100 | 100 | ok/CVD
2023-05-24 | JR | 1012.6 | 22.6 | 295.75 | 85.956 | 22.8 | 1012.6 | 23.2 | 1012 | -0.4 | 0.6 | 0.01 | 0 | -0.02 | 0 | 0 | 0 | 0.022361 | 0.01 | 0 | 0.24 | 0 | 0 | 0 |  | -2.64 | 1.712 | -0.2331% | -0.3492% | 88 | 100 | AD
2023-05-25 | GK,JR | 1009.7 | 22.5 | 295.65 | 85.956 | 22 | 1009.7 | 22.8 | 1009 | -0.8 | 0.7 | 0.03 | 0.01 | 0 | 0 | 0 | 0 | 0.031623 | 0.03 | 0 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.717 | 0.05828% | -0.05821% | 100 | 100
2023-05-26 | GK,SF | 1010.7 | 22.6 | 295.75 | 85.956 | 22.7 | 1010.7 | 23.1 | 1011 | -0.4 | -0.3 | 0.01 | 0.04 | 0 | 0 | 0 | 0 | 0.041231 | 0.01 | 0.05 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.715 | -0.05828% | -0.1746% | 89.8 | 100 | ok/AC
2023-05-29 | GK,SF | 1016.6 | 22.6 | 295.75 | 85.956 | 22.2 | 1016.6 | 22.8 | 1016 | -0.6 | 0.6 | -0.01 | 0.04 | 0.05 | 0 | 0 | 0 | 0.064807 | -0.01 | 0.05 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.712 | -0.2331% | -0.3492% | 99.6 | 100
2023-05-30 | SM, JR | 1018.6 | 22.5 | 295.65 | 85.956 | 22.7 | 1018.6 | 22.8 | 1018 | -0.1 | 0.6 | 0.01 | 0 | -0.07 | 0 | 0 | 0 | 0.070711 | 0.01 | 0 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.71 | -0.3497% | -0.4657% | 100 | 100 | ok/YM
2023-05-31 | GF, SF | 1011.7 | 22.6 | 295.75 | 85.956 | 22.8 | 1011.7 | 23.3 | 1011 | -0.5 | 0.7 | 0.01 | 0.01 | 0 | 0 | 0 | 0 | 0.014142 | 0 | 0 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.716 | 0% | -0.1164% | 100 | 100 | ok/YM
2023-06-01 | ON, AJ | 1012.2 | 22.6 | 295.75 | 85.956 | 22.6 | 1012.2 | 23 | 1012 | -0.4 | 0.2 | -0.06 | 0.05 | -0.03 | 0.1 | 0 | 0 | 0.083666 | -0.06 | 0.05 | 0.24 | 0.1 | 0 | 0 |  | -2.76 | 1.713 | -0.1748% | -0.291% | 100 |  | AD
2023-06-02 | EF, EN | 1015.9 | 22.7 | 295.85 | 85.956 | 22.7 | 1015.9 | 24.1 | 1016 | -1.4 | -0.1 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.08 | 0.08 | 0.22 | 0 | -0.1 | 0 |  | -2.78 | 1.702 | -0.8159% | -0.9313% | 87 | 100 | AC
2023-06-05 | AR, EN | 1016.7 | 22.9 | 296.05 | 85.956 | 23 | 1016.7 | 23.3 | 1016 | -0.3 | 0.7 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0.02 | -0.01 | -0.03 | 0.21 | 0 | 0 | 0 |  | -2.79 | 1.715 | -0.05828% | -0.1746% | 100
2023-06-07 | AJ, EF | 1018.3 | 22.6 | 295.75 | 85.956 | 22.6 | 1018.3 | 23.5 | 1018 | -0.9 | 0.3 | -0.02 | 0.03 | 0.05 | 0 | 0 | 0 | 0.061644 | -0.02 | -0.03 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.713 | -0.1748% | -0.291% | 96.6 |  | ok/YM
2023-06-08 | EF, AR | 1013.8 | 22.7 | 295.85 | 85.956 | 22.6 | 1013.8 | 23.4 | 1014 | -0.8 | -0.2 | -0.03 | 0 | -0.1 | 0.1 | 0 | 0 | 0.1044 | -0.03 | 0 | 0.25 | 0.1 | 0 | 0 |  | -2.75 | 1.715 | -0.05828% | -0.1746% | 100 |  | GB
2023-06-09 | ON, EN | 1021.6 | 22.7 | 295.85 | 85.956 | 22.6 | 1021.6 | 23.5 | 1022 | -0.9 | -0.4 | 0.01 | 0.01 | 0.03 | -0.1 | 0 | 0 | 0.033166 | 0.01 | 0.01 | 0.21 | 359.9 | 0 | 0 |  | -2.89 | 1.707 | -0.5245% | -0.6403% | 94 | 100
2023-06-12 | EF, EN | 1023.5 | 22.6 | 295.75 | 85.956 | 22.6 | 1023.5 | 24.4 | 1023 | -1.8 | 0.5 | 0 | 0 | -0.02 | 0 | 0 | 0 | 0.02 | 0 | 0 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.708 | -0.4662% | -0.5821% | 100 |  | AC
2023-06-13 | EF, EN | 1018 | 22.9 | 296.05 | 85.956 | 22.7 | 1018 | 23.7 | 1018 | -1 | 0 | -0.01 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0.01 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.713 | -0.1748% | -0.291% | 100 |  | GB
2023-06-14 | GK,ON | 1017.7 | 22.6 | 295.75 | 85.956 | 22.7 | 1017.7 | 23 | 1018 | -0.3 | -0.3 | 0 | -0.01 | 0.03 | 0 | 0 | 0 | 0.031623 | 0 | -0.03 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.713 | -0.1748% | -0.291% | 100 | 100 | ok / YM
2023-06-15 | ON | 1017 | 22.6 | 295.75 | 85.956 | 22.7 | 1017 | 23.3 | 1017 | -0.6 | 0 | 0.01 | 0.01 | 0 | -0.1 | 0 | 0 | 0.014142 | 0.01 | 0.01 | 0.22 | 359.9 | 0 | 0 |  | -2.78 | 1.714 | -0.1166% | -0.2328% | 100 |  | ok / YM
2023-06-16 | OJ | 1013.6 | 22.7 | 295.85 | 85.956 | 22.7 | 1013.6 | 24.5 | 1014 | -1.8 | -0.4 | 0 | 0.05 | 0.02 | 0.2 | 0 | 0 | 0.053852 | -0.02 | 0.06 | 0.2 | 0.2 | 0 | 0 |  | -2.8 | 1.706 | -0.5828% | -0.6985% | 100 |  | AC
2023-06-19 | GK | 1012.4 | 22.6 | 295.75 | 85.956 | 22.9 | 1012.4 | 23 | 1013 | -0.1 | -0.6 | -0.03 | 0 | 0.1 | -0.1 | 0 | 0 | 0.1044 | -0.03 | 0 | 0.21 | 0 | 0 | 0 |  | 2.79 | 1.736 | 1.166% | 1.048% | 36.8 | 100 | After LLRF calibration /ML
2023-06-20 | EN, ON | 1012.7 | 22.7 | 295.85 | 85.956 | 22.6 | 1012.7 | 23 | 1014 | -0.4 | -1.3 | 0.04 | 0 | 0.03 | 0 | 0 | 0 | 0.05 | -0.05 | 0.05 | 0.28 | 0 | -0.1 | 0 |  | -2.72 | 1.711 | -0.2914% | -0.4075% | 97.9 | 100
2023-06-21 | ON, EN | 1011.9 | 22.6 | 295.75 | 85.956 | 22.7 | 1011.9 | 23.1 | 1013 | -0.4 | -1.1 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0.01 | -0.06 | 0.06 | 0.29 | 0 | -0.1 | 0 |  | -2.71 | 1.71 | -0.3497% | -0.4657% | 89.4 | 100 | CVD// the value of Dcentral was wrong, ML corrected it
2023-06-22 | OJ | 1006.1 | 22.7 | 295.85 | 85.956 | 22.6 | 1006.1 | 23.8 | 1007 | -1.2 | -0.9 | -0.02 | 0.03 | 0 | 0 | 0 | 0 | 0.036056 | -0.02 | 0.03 | 0.3 | 0 | 0 | 0 |  | -2.7 | 1.716 | 0% | -0.1164% | 100
2023-06-26 | OJ, EN | 1011.7 | 22.8 | 295.95 | 85.956 | 22.7 | 1011.7 | 23.8 | 1013 | -1.1 | -1.3 | -0.01 | -0.02 | -0.03 | 0 | 0 | 0 | 0.037417 | -0.01 | -0.02 | 0.29 | 0 | 0 | 0 |  | -2.71 | 1.721 | 0.2914% | 0.1746% | 89.1 | 100
2023-06-27 | GK,AR | 1003.9 | 22.9 | 296.05 | 85.956 | 22.9 | 1003.9 | 23 | 1005 | -0.1 | -1.1 | -0.01 | 0.01 | -0.03 | -0.1 | 0 | 0 | 0.033166 | -0.01 | 0.01 | 0.25 | 359.9 | 0 | 0 |  | -2.75 | 1.717 | 0.05828% | -0.05821% | 89.3 | 100 | ok/AC
2023-06-28 | EN, ON | 1007.5 | 22.7 | 295.85 | 85.956 | 22.7 | 1007.5 | 22.7 | 1008 | 0 | -0.5 | -0.01 | 0 | 0.03 | 0 | 0 | 0 | 0.031623 | -0.01 | 0 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.718 | 0.1166% | 0% | 89.1 | 100 | AD
2023-06-29 | OJ, AR | 1004.9 | 22.7 | 295.85 | 85.956 | 22.2 | 1004.9 | 22.7 | 1006 | -0.5 | -1.1 | 0.03 | 0 | -0.02 | 0 | 0 | 0 | 0.036056 | 0.03 | 0.01 | 0.2 | 0 | 0 | 0 |  | -2.8 | 1.712 | -0.2331% | -0.3492% | 100 |  | ok / ML YM
2023-06-30 | AR, OJ | 1001.1 | 22.7 | 295.85 | 85.956 | 22.7 | 1001.1 | 23.1 | 1002 | -0.4 | -0.9 | -0.01 | 0.01 | 0 | 0 | 0 | 0 | 0.014142 | -0.01 | 0.01 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.715 | -0.05828% | -0.1746% | 90.2 | 100 | ok / ML YM
2023-07-03 | OJ, EN | 984.8 | 22.7 | 295.85 | 85.956 | 22.6 | 984.8 | 23.3 | 986 | -0.7 | -1.2 | -0.01 | 0.01 | 0.02 | 0 | 0 | 0 | 0.024495 | -0.01 | 0.01 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.712 | -0.2331% | -0.3492% | 100 |  | ok/AC
2023-07-04 | AR | 991.6 | 23.1 | 296.25 | 85.956 | 23.1 | 991.6 | 23.1 | 993 | 0 | -1.4 | -0.01 | 0.01 | 0 | 0 | 0 | 0 | 0.014142 | -0.01 | 0.01 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.711 | -0.2914% | -0.4075% | 97 | 100 | ok/AC
2023-07-05 | ON, EN | 1000.6 | 23.1 | 296.25 | 85.956 | 22.8 | 1000.6 | 23.1 | 1002 | -0.3 | -1.4 | 0.03 | -0.03 | 0 | 0 | 0 | 0 | 0.042426 | 0 | 0.05 | 0.22 | 0.1 | -0.2 | 0.1 |  | -2.78 | 1.709 | -0.4079% | -0.5239% | 100 |  | AD
2023-07-06 | ON | 1005.6 | 23 | 296.15 | 85.956 | 22.7 | 1005.6 | 23.3 | 1006 | -0.6 | -0.4 | 0 | 0 | -0.07 | 0 | 0 | 0 | 0.07 | 0 | 0 | 0.2 | 0 | 0 | 0 |  | -2.8 | 1.712 | -0.2331% | -0.3492% | 100 |  | AD
2023-07-07 | AR | 1007.9 | 22.5 | 295.65 | 85.956 | 22.9 | 1007.9 | 23.1 | 1008 | -0.2 | -0.1 | 0.01 | 0.01 | 0.03 | 0 | 0 | 0 | 0.033166 | 0.01 | 0.01 | 0.2 | 0 | 0 | 0 |  | -2.8 | 1.702 | -0.8159% | -0.9313% | 97.9 | 100 | ML
2023-07-10 | AR, EN | 1017.6 | 22.7 | 295.85 | 85.956 | 22.8 | 1017.6 | 23 | 1018 | -0.2 | -0.4 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0.02 | 0.02 | 0 | 0.19 | 0 | 0 | 0 |  | -2.81 | 1.7121 | -0.2273% | -0.3434% | 89.5 | 100 | AC
2023-07-11 | EN, AR | 1013 | 23 | 296.15 | 85.956 | 22.9 | 1013 | 23 | 1014 | -0.1 | -1 | 0.01 | -0.01 | 0 | 0 | 0 | 0 | 0.014142 | 0.01 | -0.01 | 0.25 | 0 | 0 | 0 |  | -2.75 | 1.714 | -0.1166% | -0.2328% | 100 |  | ok/YM
2023-07-12 | OJ | 1008.1 | 22.9 | 296.05 | 85.956 | 22.8 | 1008.1 | 23.9 | 1008 | -1.1 | 0.1 | -0.02 | 0.02 | 0 | 0 | 0 | 0 | 0.028284 | -0.02 | 0.02 | 0.23 | 0 | 0 | 0 |  | -2.77 | 1.711 | -0.2914% | -0.4075% | 100 |  | ok/YM
2023-07-13 | OJ, AR | 1000.7 | 23.1 | 296.25 | 85.956 | 22.9 | 1000.7 | 23.2 | 1002 | -0.3 | -1.3 | 0.01 | 0.01 | -0.03 | 0 | 0 | 0 | 0.033166 | 0.01 | 0.01 | 0.22 | 0 | 0 | 0 |  | -2.78 | 1.714 | -0.1166% | -0.2328% | 99.1 |  | AD
2023-07-14 | OJ, EN | 1003.8 | 22.9 | 296.05 | 85.956 | 22.7 | 1003.8 | 23 | 1005 | -0.3 | -1.2 | -0.02 | 0.01 | -0.12 | 0 | 0 | 0 | 0.12207 | -0.02 | 0 | 0.21 | 0 | 0 | 0 |  | -2.79 | 1.72 | 0.2331% | 0.1164% | 100 |  | ok/YM
2023-07-17 | EN, AR | 1003.4 | 22.8 | 295.95 | 85.956 | 22.4 | 1003.4 | 23.1 | 1005 | -0.7 | -1.6 | -0.01 | 0.02 | 0 | 0.1 | 0 | 0 | 0.022361 | -0.01 | 0.02 | 0.23 | 0.1 | 0 | 0 |  | -2.77 | 1.711 | -0.2914% | -0.4075% | 100 |  | AD/YM
2023-07-18 | AR | 999.5 | 23.2 | 296.35 | 85.956 | 22.9 | 999.5 | 24.4 | 1001 | -1.5 | -1.5 | -0.01 | -0.01 | 0 | 0 | 0 | 0 | 0.014142 | 0 | 0 | 0.21 | 0 | 0 | 0 |  | -2.79 | 1.7176 | 0.09324% | -0.02328% | 100 |  | AC,AD
2023-07-19 | AR | 1000 | 22.9 | 296.05 | 85.956 | 22.8 | 1000 | 23.2 | 1001 | -0.4 | -1 | 0.01 | 0 | -0.02 | 0 | 0 | 0 | 0.022361 | 0.01 | 0 | 0.2 | 0 | 0 | 0 |  | -2.79 | 1.7176 | 0.09091% | -0.02561% | 94.9 | 100 | AC,YM
2023-07-20 | ON | 997 | 22.8 | 295.95 | 85.956 | 22.3 | 997 | 23.1 | 998 | -0.8 | -1 | 0.04 | 0 | 0 | 0 | 0 | 0 | 0.04 | 0.04 | 0 | 0.18 | 0 | 0 | 0 |  | -2.82 | 1.715 | -0.05828% | -0.1746% | 100 |  | YM
2023-07-21 | EN | 997.4 | 23.1 | 296.25 | 85.956 | 22.6 | 997.4 | 22.6 | 999 | 0 | -1.6 | -0.04 | -0.01 | 0.05 | 0 | 0 | 0 | 0.064807 | -0.04 | -0.01 | 0.19 | 0 | 0 | 0 |  | -2.81 | 1.717 | 0.05828% | -0.05821% | 94.8 |  | YM,NIC
2023-07-24 | ON, EN | 999.4 | 22.7 | 295.85 | 85.956 | 22.6 | 998.8 | 23.7 | 1000 | -1.1 | -1.2 | -0.05 | 0.02 | 0.03 | 0 | 0 | 0 | 0.061644 | -0.05 | 0.02 | 0.24 | 0 | 0 | 0 |  | -2.76 | 1.715 | -0.05828% | -0.1746% | 100 |  | KW, YM
2023-07-25 | AR, ON | 990.7 | 23.2 | 296.35 | 85.956 | 23 | 990.7 | 22.3 | 992 | 0.7 | -1.3 | -0.05 | 0 | 0.05 | 0 | 0 | 0 | 0.070711 | -0.05 | 0.01 | 0.2 | 0 | 0 | 0 |  | -2.8 | 1.721 | 0.2914% | 0.1746% | 100 |  | KW, YM
2023-07-26 | EN | 994 | 22.9 | 296.05 | 85.956 | 22.7 | 994 | 22.8 | 996 | -0.1 | -2 | -0.01 | 0.01 | 0.02 | 0 | 0 | 0 | 0.024495 | -0.01 | 0.01 | 0.19 | 0 | 0 | 0 |  | -2.81 | 1.716 | 0% | -0.1164% | 100 |  | ok/YM
2023-07-27 | EN, AR | 998.8 | 22.7 | 295.85 | 85.956 | 22.7 | 998.8 | 22.7 | 1001 | 0 | -2.2 | -0.03 | 0.02 | 0 | 0.2 | 0 | 0 | 0.036056 | -0.03 | 0.02 | 0.21 | 0.2 | 0 | 0 |  | -2.79 | 1.717 | 0.05828% | -0.05821% | 100 |  | ok/ YM KW
2023-07-28 | EN | 997.4 | 23 | 296.15 | 85.956 | 22.8 | 997.4 | 22.8 | 999 | 0 | -1.6 | -0.01 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0 | 0.18 | 0 | 0 | 0 |  | -2.82 | 1.7155 | -0.02797% | -0.1444% | 100 |  | ok /YM CMN
Note on F81: Christina Vallhagen Dahlgren: inmatning i dosimetry manager blev av misstag 296,1, men bra mätta doser så ok!

## GTR1_150MeV
ID.nr:
Titel: | [24 empty] | Dokumenttyp:
Protocol for Daily QA: Consistency check of single layer and cubic plans dosimetric parameters | [24 empty] | Strålskyddsinstruktion
Godkänt av: | [23 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [23 empty] | Skapat den:
Reviderat av: | [23 empty] | Reviderat den:
GTR1
Daily single layer parameters
MatriXX PT_2
Parameters and its criterias
Parameter | Value (Gy) | Good: within 1% |  | Warning: within 2% |  | Action level: over 2%
Ref. dose | 1.716 | 1.7332 | 1.6988 | 1.7342 | 1.6978 | 1.7513 | 1.6807
Dose from TPS | 1.718 | 1.7352 | 1.7008 | 1.7362 | 1.6998 | 1.7534 | 1.6826
Parameter | Value (mm) |  |  | +/- 1.001 mm |  | +/- 2.001 mm
Penumbra (average) | 8.1417 | 9.1417 | 7.1417 | 9.1427 | 7.1407 | 10.1427 | 6.1407
FWHM (average) | 104.8556 | 105.8556 | 103.8556 | 105.8566 | 103.8546 | 106.8566 | 102.8546
Parameter | Value (Gy) | Good: within +/-0.5% |  | Warning: within +/-1% |  | Action level: over 1%
Symmetry | x | 0.5 | -0.5 | 0.5 | -0.5 | 1 | 1
Parameter | Value (Gy) | Good: within +/-1% |  | Warning: within +/-2% |  | Action level: over 3%
Flatness | x | 1 | -1 | 2 | -2 | 3 | -3
| Output Calibration (150 MeV 10x10 cm2, 1MU per spot) | [19 empty] | Plot
Date | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Sign | Checked | Notes | Date | Difference from reference (a.u.) | Medelvärde +/- 10 dygn | Medelvärde +/- 20 dygn | Ref | 1% | -1% | 2% | -2%
20221104 | 1.714 | -0.1166% | -0.2328% | 0.16 | 0.22 | 0.01 | 0.06 | 0.42 | 0.05 | 0.12 | 1.714 | 8.7 | 8 | 8 | 7.6 | 104.7 | 105.1 | ML | x | GTR1_20221104_OF_05 | 2022-11-04 | 0.99883 | 1 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221111 | 1.714 | -0.1166% | -0.2328% | 0.31 | 0.37 | 0.18 | 0.15 | 0.73 | 0.15 | 0.2 | 1.714 | 8.1 | 8.8 | 8 | 7.7 | 104.6 | 105 | ML | x | GTR1_20221111_OF_02 | 2022-11-11 | 0.99883 | 1 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221111 | 1.714 | -0.1166% | -0.2328% | 0.39 | 0.43 | 0.23 | 0.12 | 0.8 | 0.18 | 0.18 | 1.714 | 8.1 | 8.8 | 7.9 | 7.9 | 104.6 | 105 | ML | x | GTR1_20221111_OF_03 | 2022-11-11 | 0.99883 | 1 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221111 | 1.713 | -0.1748% | -0.291% | 0.47 | 0.38 | 0.29 | 0.18 | 0.7 | 0.19 | 0.18 | 1.713 | 8.1 | 8.8 | 7.9 | 7.9 | 104.6 | 105 | ML | x | GTR1_20221111_OF_04 | 2022-11-11 | 0.99825 | 1 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221209 | 1.72 | 0.2331% | 0.1164% | 0.31 | 0.33 | 0.17 | 0.16 | 0.76 | 0.16 | 0.19 | 1.72 | 8.1 | 8.8 | 7.8 | 7.8 | 104.7 | 105 | ML | x | GTR1_20221209_OF_01 | 2022-12-09 | 1.0023 | 1 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230214 | 1.721 | 0.2914% | 0.1746% | 0.31 | 0.24 | 0.12 | 0.09 | 0.56 | 0.15 | 0.1 | 1.721 | 8.7 | 8.2 | 8 | 7.6 | 104.7 | 105.1 | AJ | ML | GTR1_20230214_OF_01 | 2023-02-14 | 1.0029 | 1 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230214 | 1.721 | 0.2914% | 0.1746% | 0.31 | 0.24 | 0.12 | 0.09 | 0.56 | 0.15 | 0.1 | 1.721 | 8.7 | 8.2 | 8 | 7.6 | 104.7 | 105.1 | AE | ML corr | GTR1_20230217_OF_01 | 2023-02-14 | 1.0029 | 1.0001 | 1.0001 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230217 | 1.718 | 0.1166% | 0% | 0.28 | 0.4 | 0.13 | 0.14 | 0.66 | 0.14 | 0.14 | 1.718 | 8.4 | 8.6 | 8.1 | 7.6 | 104.7 | 105 | Students | ML corr | GTR1_20230217_OF_01 | 2023-02-17 | 1.0012 | 1.0002 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230221 | 1.715 | -0.05828% | -0.1746% | 0.29 | 0.42 | 0.07 | 0.17 | 0.57 | 0.1 | 0.11 | 1.715 | 8.5 | 8.4 | 8.2 | 7.5 | 104.7 | 105.1 | SM, GF, AR, OJ | ML | GTR1_20230221_OF_01 | 2023-02-21 | 0.99942 | 1.0004 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230222 | 1.712 | -0.2331% | -0.3492% | 0.35 | 0.34 | 0.14 | 0.12 | 0.57 | 0.13 | 0.12 | 1.712 | 8.5 | 8.4 | 8.3 | 7.5 | 104.7 | 105 | AR | ML | GTR1_20230222_OF_01 | 2023-02-22 | 0.99767 | 1.0003 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230223 | 1.71 | -0.3497% | -0.4657% | 0.32 | 0.26 | 0.12 | 0.07 | 0.59 | 0.14 | 0.12 | 1.71 | 8.6 | 8.3 | 7.9 | 7.9 | 104.7 | 105 | OJ | ML | GTR1_20230223_OF_01 | 2023-02-23 | 0.9965 | 1.0003 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230224 | 1.71 | -0.3497% | -0.4657% | 0.26 | 0.37 | 0.06 | 0.12 | 0.63 | 0.09 | 0.13 | 1.71 | 8.2 | 8.7 | 8.2 | 7.5 | 104.7 | 105.1 | EH | AD | GTR1_20230224_OF_01 | 2023-02-24 | 0.9965 | 1.0003 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230227 | 1.707 | -0.5245% | -0.6403% | 0.35 | 0.31 | 0.16 | 0.13 | 0.61 | 0.14 | 0.1 | 1.707 | 7.6 | 9.1 | 7.8 | 8 | 104.6 | 105 | AC | AC | GTR1_20230227_OF_01 | 2023-02-27 | 0.99476 | 1.0004 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230228 | 1.724 | 0.4662% | 0.3492% | 0.35 | 0.35 | 0.14 | 0.15 | 0.75 | 0.18 | 0.16 | 1.724 | 8.6 | 8.5 | 7.9 | 8 | 104.7 | 105 | AR, AJ | YM | GTR1_20230228_OF_01 | 2023-02-28 | 1.0047 | 1.0005 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230228 | 1.723 | 0.4079% | 0.291% | 0.35 | 0.38 | 0.17 | 0.17 | 0.6 | 0.16 | 0.16 | 1.723 | 8.5 | 8.5 | 8.1 | 7.7 | 104.7 | 105 | ML | x | GTR1_20230228_OF_02 | 2023-02-28 | 1.0041 | 1.0005 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230301 | 1.721 | 0.2914% | 0.1746% | 0.29 | 0.35 | 0.09 | 0.14 | 0.54 | 0.11 | 0.12 | 1.721 | 8.5 | 8.5 | 8.1 | 7.7 | 104.7 | 105.1 | AR, JR | GB | GTR1_20230301_OF_01 | 2023-03-01 | 1.0029 | 0.99988 | 0.99954 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230302 | 1.717 | 0.05828% | -0.05821% | 0.29 | 0.35 | 0.07 | 0.15 | 0.5 | 0.13 | 0.13 | 1.717 | 8.6 | 8.3 | 8 | 7.7 | 104.7 | 105.1 | SM, SF | YM | GTR1_20230302_OF_01 | 2023-03-02 | 1.0006 | 1.0002 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230303 | 1.717 | 0.05828% | -0.05821% | 0.29 | 0.31 | 0.13 | 0.15 | 0.63 | 0.12 | 0.18 | 1.717 | 8.1 | 8.8 | 8.2 | 7.6 | 104.6 | 105 | OJ | AC | GTR1_20230303_OF_01 | 2023-03-03 | 1.0006 | 1.0002 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230306 | 1.723 | 0.4079% | 0.291% | 0.33 | 0.35 | 0.14 | 0.18 | 0.7 | 0.14 | 0.16 | 1.723 | 8 | 8.7 | 7.8 | 7.6 | 104.7 | 105.1 | ML | x | GTR1_20230306_OF_01 | 2023-03-06 | 1.0041 | 1.0002 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230306 | 1.723 | 0.4079% | 0.291% | 0.33 | 0.35 | 0.14 | 0.18 | 0.7 | 0.14 | 0.16 | 1.723 | 8 | 8.7 | 7.8 | 7.6 | 104.7 | 105.1 | x | ML | GTR1_20230306_OF_01 | 2023-03-06 | 1.0041 | 1.0004 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230306 | 1.714 | -0.1166% | -0.2328% | 0.35 | 0.3 | 0.16 | 0.1 | 0.56 | 0.16 | 0.12 | 1.714 | 8.7 | 8.3 | 8 | 7.8 | 104.7 | 105 | AD, GB | ML | GTR1_20230306_OF_02 // run 2, after ion source replacement | 2023-03-06 | 0.99883 | 1.0007 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230307 | 1.718 | 0.1166% | 0% | 0.32 | 0.33 | 0.11 | 0.13 | 0.57 | 0.16 | 0.12 | 1.718 | 8.7 | 8.3 | 8.2 | 7.5 | 104.6 | 105 | OJ | YM KW | GTR1_20230307_OF_01 | 2023-03-07 | 1.0012 | 1.001 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230308 | 1.714 | -0.1166% | -0.2328% | 0.28 | 0.28 | 0.11 | 0.13 | 0.61 | 0.16 | 0.13 | 1.714 | 8.6 | 8.3 | 8.3 | 7.2 | 104.7 | 105.1 | JR | YM | GTR1_20230308_OF_02 | 2023-03-08 | 0.99883 | 1.0014 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230309 | 1.716 | 0% | -0.1164% | 0.37 | 0.37 | 0.17 | 0.16 | 0.66 | 0.17 | 0.17 | 1.716 | 8.7 | 8.3 | 8 | 7.7 | 104.7 | 105 | JR | YM | GTR1_20230309_OF_01 | 2023-03-09 | 1 | 1.0012 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230310 | 1.718 | 0.1166% | 0% | 0.23 | 0.35 | 0.09 | 0.12 | 0.56 | 0.11 | 0.13 | 1.718 | 8.4 | 8.5 | 7.9 | 7.8 | 104.7 | 105.1 | SM, AR | AD | GTR1_20230310_OF_01 | 2023-03-10 | 1.0012 | 1.0009 | 1.0003 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230313 | 1.712 | -0.2331% | -0.3492% | 0.29 | 0.34 | 0.09 | 0.12 | 0.62 | 0.14 | 0.13 | 1.712 | 8.7 | 8.2 | 8 | 7.8 | 104.7 | 105.1 | SM, SF | AC | GTR1_20230313_OF_01 | 2023-03-13 | 0.99767 | 1.0007 | 1.0004 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230314 | 1.713 | -0.1748% | -0.291% | 0.31 | 0.4 | 0.13 | 0.18 | 0.67 | 0.14 | 0.17 | 1.713 | 8.1 | 8.8 | 8.2 | 7.5 | 104.7 | 105 | AE | YM | GTR1_20230314_OF_01 | 2023-03-14 | 0.99825 | 0.99928 | 0.99961 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230315 | 1.711 | -0.2914% | -0.4075% | 0.37 | 0.29 | 0.16 | 0.11 | 0.74 | 0.15 | 0.16 | 1.711 | 8.7 | 8.1 | 7.9 | 7.8 | 104.7 | 105 | SM, AE | CVD | GTR1_20230315_OF_01 | 2023-03-15 | 0.99709 | 0.99922 | 0.99951 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230316 | 1.709 | -0.4079% | -0.5239% | 0.19 | 0.31 | 0.05 | 0.13 | 0.58 | 0.09 | 0.13 | 1.709 | 8.4 | 8.5 | 7.8 | 7.9 | 104.7 | 105.1 | JR | ML | GTR1_20230316_OF_01 | 2023-03-16 | 0.99592 | 0.999 | 0.99947 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230320 | 1.713 | -0.1748% | -0.291% | 0.26 | 0.35 | 0.09 | 0.13 | 0.53 | 0.09 | 0.13 | 1.713 | 8.4 | 8.5 | 7.7 | 8.1 | 104.7 | 105 | AC | AC | GTR1_20230320_OF_01 | 2023-03-20 | 0.99825 | 0.99878 | 0.99943 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230321 | 1.71 | -0.3497% | -0.4657% | 0.25 | 0.4 | 0.09 | 0.15 | 0.55 | 0.08 | 0.12 | 1.71 | 8.4 | 8.6 | 8.4 | 7.4 | 104.7 | 105 | GK,AJ | GB | GTR1_20230321_OF_01 | 2023-03-21 | 0.9965 | 0.99883 | 0.99912 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230322 | 1.712 | -0.2331% | -0.3492% | 0.27 | 0.31 | 0.05 | 0.13 | 0.63 | 0.12 | 0.15 | 1.712 | 8.7 | 8.2 | 8 | 7.7 | 104.7 | 105 | GK,SM | CVD | GTR1_20230322_OF_01 | 2023-03-22 | 0.99767 | 0.99886 | 0.99907 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230323 | 1.718 | 0.1166% | 0% | 0.31 | 0.39 | 0.08 | 0.18 | 0.57 | 0.1 | 0.13 | 1.718 | 8.4 | 8.5 | 8.1 | 7.6 | 104.7 | 105.1 | EH | ML | GTR1_20230323_OF_01 | 2023-03-23 | 1.0012 | 0.999 | 0.99926 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230324 | 1.718 | 0.1166% | 0% | 0.27 | 0.28 | 0.07 | 0.14 | 0.63 | 0.15 | 0.16 | 1.718 | 8.6 | 8.2 | 8 | 7.7 | 104.7 | 105.1 | AE | KW | GTR1_20230324_OF_01 | 2023-03-24 | 1.0012 | 0.99903 | 0.99923 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230327 | 1.708 | -0.4662% | -0.5821% | 0.33 | 0.26 | 0.1 | 0.09 | 0.53 | 0.12 | 0.13 | 1.708 | 8.6 | 8.4 | 8.1 | 7.7 | 104.7 | 105.1 | SM, AR | AC | GTR1_20230327_OF_01 | 2023-03-27 | 0.99534 | 0.99892 | 0.99927 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230328 | 1.715 | -0.05828% | -0.1746% | 0.3 | 0.33 | 0.08 | 0.09 | 0.59 | 0.14 | 0.12 | 1.715 | 8.6 | 8.3 | 7.9 | 7.8 | 104.7 | 105.1 | GK,ON,SF | YM | GTR1_20230328_OF_01 | 2023-03-28 | 0.99942 | 0.99892 | 0.99922 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230329 | 1.713 | -0.1748% | -0.291% | 0.34 | 0.38 | 0.12 | 0.18 | 0.67 | 0.13 | 0.16 | 1.713 | 8.7 | 8.3 | 8 | 7.7 | 104.7 | 105.1 | JR | YM ON | GTR1_20230329_OF_01 | 2023-03-29 | 0.99825 | 0.99878 | 0.99918 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230330 | 1.715 | -0.05828% | -0.1746% | 0.31 | 0.35 | 0.15 | 0.16 | 0.76 | 0.13 | 0.19 | 1.715 | 8.1 | 8.8 | 8.2 | 7.6 | 104.7 | 105 | SM, OJ | ON | GTR1_20230330_OF_01 | 2023-03-30 | 0.99942 | 0.99886 | 0.99914 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230331 | 1.715 | -0.05828% | -0.1746% | 0.28 | 0.3 | 0.09 | 0.13 | 0.62 | 0.14 | 0.14 | 1.715 | 8.7 | 8.2 | 8 | 7.8 | 104.7 | 105 | ON | GB | GTR1_20230331_OF_01 | 2023-03-31 | 0.99942 | 0.99911 | 0.99913 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230403 | 1.715 | -0.05828% | -0.1746% | 0.29 | 0.34 | 0.13 | 0.17 | 0.62 | 0.15 | 0.18 | 1.715 | 8.8 | 8.2 | 7.8 | 7.9 | 104.7 | 105.1 | OJ | YM | GTR1_20230403_OF_01 | 2023-04-03 | 0.99942 | 0.99917 | 0.99911 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230404 | 1.715 | -0.05828% | -0.1746% | 0.3 | 0.28 | 0.13 | 0.11 | 0.62 | 0.16 | 0.13 | 1.715 | 8.7 | 8.3 | 7.8 | 8 | 104.7 | 105 | AE | GB | GTR1_20230404_OF_01 | 2023-04-04 | 0.99942 | 0.99896 | 0.99906 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230405 | 1.716 | 0% | -0.1164% | 0.37 | 0.28 | 0.2 | 0.14 | 0.68 | 0.2 | 0.15 | 1.716 | 8.6 | 8.3 | 7.8 | 7.8 | 104.7 | 105.1 | ON, SM, JR | GB | GTR1_20230405_OF_01 | 2023-04-05 | 1 | 0.99889 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230406 | 1.719 | 0.1748% | 0.05821% | 0.26 | 0.26 | 0.09 | 0.11 | 0.63 | 0.14 | 0.14 | 1.719 | 8.6 | 8.4 | 8 | 7.7 | 104.7 | 105.1 | OJ | CVD | GTR1_20230406_OF_01 | 2023-04-06 | 1.0017 | 0.99908 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230407 | 1.719 | 0.1748% | 0.05821% | 0.37 | 0.32 | 0.17 | 0.13 | 0.62 | 0.15 | 0.11 | 1.719 | 8.4 | 8.5 | 8 | 7.8 | 104.7 | 105.1 | JR,SM | CVD | GTR1_20230407_OF_01 | 2023-04-07 | 1.0017 | 0.99904 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230411 | 1.717 | 0.05828% | -0.05821% | 0.34 | 0.33 | 0.1 | 0.14 | 0.63 | 0.15 | 0.16 | 1.717 | 8.6 | 8.4 | 8.3 | 7.5 | 104.7 | 105 | OJ, GK | GB | GTR1_20230411_OF_01 | 2023-04-11 | 1.0006 | 0.99908 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230412 | 1.714 | -0.1166% | -0.2328% | 0.25 | 0.32 | 0.08 | 0.09 | 0.53 | 0.13 | 0.1 | 1.714 | 8.8 | 8.1 | 7.7 | 8.1 | 104.7 | 105.1 | GK,EH | YM | GTR1_20230412_OF_01 | 2023-04-12 | 0.99883 | 0.99908 | 0.99898 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230413 | 1.712 | -0.2331% | -0.3492% | 0.34 | 0.38 | 0.12 | 0.2 | 0.65 | 0.13 | 0.14 | 1.712 | 8.5 | 8.5 | 8.3 | 7.4 | 104.7 | 105.1 | AR | YM | GTR1_20230413_OF_01 | 2023-04-13 | 0.99767 | 0.99909 | 0.99896 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230414 | 1.708 | -0.4662% | -0.5821% | 0.29 | 0.32 | 0.11 | 0.13 | 0.66 | 0.16 | 0.15 | 1.708 | 8.6 | 8.4 | 8.3 | 7.4 | 104.7 | 105.1 | AR, JR | GB | GTR1_20230414_OF_01 | 2023-04-14 | 0.99534 | 0.99909 | 0.99897 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230417 | 1.714 | -0.1166% | -0.2328% | 0.32 | 0.33 | 0.11 | 0.12 | 0.65 | 0.15 | 0.17 | 1.714 | 8.6 | 8.4 | 7.8 | 7.9 | 104.7 | 105.1 | GK | GB | GTR1_20230417_OF_01 | 2023-04-17 | 0.99883 | 0.99908 | 0.99898 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230418 | 1.718 | 0.1166% | 0% | 0.42 | 0.33 | 0.18 | 0.2 | 0.67 | 0.15 | 0.14 | 1.718 | 8.5 | 8.5 | 8.2 | 7.6 | 104.7 | 105.1 | AJ | GB | GTR1_20230418_OF_01 | 2023-04-18 | 1.0012 | 0.99908 | 0.999 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230419 | 1.715 | -0.05828% | -0.1746% | 0.25 | 0.43 | 0.11 | 0.1 | 0.55 | 0.11 | 0.1 | 1.715 | 8.3 | 8.6 | 8 | 7.8 | 104.7 | 105.1 | JR | CVD | GTR1_20230419_OF_01 | 2023-04-19 | 0.99942 | 0.99907 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230420 | 1.714 | -0.1166% | -0.2328% | 0.25 | 0.41 | 0.1 | 0.22 | 0.69 | 0.08 | 0.2 | 1.714 | 8 | 8.8 | 8.3 | 7.5 | 104.7 | 105 | JR,AR | CVD | GTR1_20230420_OF_01 | 2023-04-20 | 0.99883 | 0.99906 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230421 | 1.715 | -0.05828% | -0.1746% | 0.36 | 0.36 | 0.15 | 0.14 | 0.62 | 0.17 | 0.15 | 1.715 | 8.8 | 8 | 8 | 7.8 | 104.7 | 105.1 | GK,AR | CVD | GTR1_20230421_OF_01 | 2023-04-21 | 0.99942 | 0.99903 | 0.99907 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230424 | 1.71 | -0.3497% | -0.4657% | 0.36 | 0.38 | 0.11 | 0.16 | 0.58 | 0.14 | 0.11 | 1.71 | 8.5 | 8.5 | 8.2 | 7.5 | 104.7 | 105.1 | AR, AJ | AC,YM | GTR1_20230424_OF_01 | 2023-04-24 | 0.9965 | 0.99899 | 0.99909 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230425 | 1.716 | 0% | -0.1164% | 0.36 | 0.29 | 0.14 | 0.13 | 0.58 | 0.15 | 0.15 | 1.716 | 8.6 | 8.3 | 8.2 | 7.4 | 104.7 | 105.1 | OJ, AR | CVD | GTR1_20230425_OF_01 | 2023-04-25 | 1 | 0.99897 | 0.99906 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230426 | 1.721 | 0.2914% | 0.1746% | 0.31 | 0.35 | 0.13 | 0.15 | 0.64 | 0.19 | 0.13 | 1.721 | 8.6 | 8.3 | 8.2 | 7.3 | 104.7 | 105.1 | ON, SM | YM | GTR1_20230426_OF_01 | 2023-04-26 | 1.0029 | 0.99897 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230427 | 1.713 | -0.1748% | -0.291% | 0.28 | 0.36 | 0.12 | 0.14 | 0.69 | 0.13 | 0.16 | 1.713 | 8.3 | 8.6 | 8.1 | 7.5 | 104.7 | 105.1 | SM, JR | CVD | GTR1_20230427_OF_01 | 2023-04-27 | 0.99825 | 0.99899 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230428 | 1.718 | 0.1166% | 0% | 0.22 | 0.36 | 0.07 | 0.13 | 0.67 | 0.11 | 0.15 | 1.718 | 8.3 | 8.6 | 8.2 | 7.6 | 104.7 | 105 | OJ, SM | AC | GTR1_20230428_OF_01 | 2023-04-28 | 1.0012 | 0.99904 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230502 | 1.712 | -0.2331% | -0.3492% | 0.32 | 0.41 | 0.14 | 0.14 | 0.71 | 0.15 | 0.16 | 1.712 | 8.3 | 8.7 | 8.2 | 7.5 | 104.6 | 105.1 | EN,AJ | CVD | GTR1_20230502_OF_01 | 2023-05-02 | 0.99767 | 0.99904 | 0.99909 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230503 | 1.71 | -0.3497% | -0.4657% | 0.33 | 0.35 | 0.14 | 0.14 | 0.71 | 0.18 | 0.17 | 1.71 | 8.6 | 8.3 | 8 | 7.7 | 104.7 | 105.1 | JR | CVD | GTR1_20230503_OF_01 | 2023-05-03 | 0.9965 | 0.99901 | 0.99909 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230504 | 1.716 | 0% | -0.1164% | 0.27 | 0.34 | 0.1 | 0.18 | 0.72 | 0.12 | 0.17 | 1.716 | 8.3 | 8.7 | 8.1 | 7.5 | 104.7 | 105.1 | GK,AR,EF | YM | GTR1_20230504_OF_01 | 2023-05-04 | 1 | 0.99901 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230505 | 1.721 | 0.2914% | 0.1746% | 0.27 | 0.28 | 0.09 | 0.1 | 0.49 | 0.1 | 0.1 | 1.721 | 8.5 | 8.4 | 7.9 | 7.8 | 104.7 | 105.1 | OJ | AC | GTR1_20230505_OF_01 | 2023-05-05 | 1.0029 | 0.99901 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230508 | 1.716 | 0% | -0.1164% | 0.31 | 0.27 | 0.09 | 0.11 | 0.55 | 0.11 | 0.11 | 1.716 | 8.5 | 8.5 | 8 | 7.7 | 104.7 | 105.1 | SM, AJ | AC | GTR1_20230508_OF_01 | 2023-05-08 | 1 | 0.999 | 0.99907 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230509 | 1.712 | -0.2331% | -0.3492% | 0.26 | 0.37 | 0.1 | 0.12 | 0.7 | 0.13 | 0.14 | 1.712 | 8.4 | 8.6 | 8.3 | 7.2 | 104.7 | 105.1 | EN | GB | GTR1_20230509_OF_01 | 2023-05-09 | 0.99767 | 0.99904 | 0.99906 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230510 | 1.716 | 0% | -0.1164% | 0.29 | 0.21 | 0.11 | 0.03 | 0.5 | 0.1 | 0.08 | 1.716 | 8.8 | 8.1 | 7.8 | 7.7 | 104.6 | 105.1 | AR | GB | GTR1_20230510_OF_01 | 2023-05-10 | 1 | 0.99903 | 0.99903 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230511 | 1.718 | 0.1166% | 0% | 0.39 | 0.37 | 0.14 | 0.13 | 0.69 | 0.15 | 0.15 | 1.718 | 8.9 | 8 | 8.1 | 7.5 | 104.6 | 105.1 | JR | CVD | GTR1_20230511_OF_01 | 2023-05-11 | 1.0012 | 0.99896 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230512 | 1.716 | 0% | -0.1164% | 0.3 | 0.35 | 0.11 | 0.14 | 0.62 | 0.15 | 0.14 | 1.716 | 8.7 | 8.2 | 8.2 | 7.4 | 104.7 | 105.1 | EN,GK,ON | EAL | GTR1_20230512_OF_01 | 2023-05-12 | 1 | 0.99898 | 0.99897 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230515 | 1.709 | -0.4079% | -0.5239% | 0.39 | 0.43 | 0.18 | 0.18 | 0.72 | 0.19 | 0.14 | 1.709 | 8.5 | 8.5 | 7.8 | 8.1 | 104.6 | 105 | AR | GB | GTR1_20230515_OF_01 | 2023-05-15 | 0.99592 | 0.99894 | 0.99897 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230516 | 1.715 | -0.05828% | -0.1746% | 0.31 | 0.32 | 0.12 | 0.13 | 0.66 | 0.14 | 0.16 | 1.715 | 8.4 | 8.5 | 7.9 | 7.8 | 104.7 | 105.1 | GF, AE, EF | ON, YM | GTR1_20230516_OF_01 | 2023-05-16 | 0.99942 | 0.99896 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230517 | 1.718 | 0.1166% | 0% | 0.29 | 0.27 | 0.1 | 0.09 | 0.59 | 0.13 | 0.14 | 1.718 | 8.4 | 8.6 | 8.3 | 7.4 | 104.7 | 105.1 | SM, GK | YM ON | GTR1_20230517_OF_01 | 2023-05-17 | 1.0012 | 0.999 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230519 | 1.72 | 0.2331% | 0.1164% | 0.2 | 0.31 | 0.06 | 0.15 | 0.56 | 0.06 | 0.15 | 1.72 | 8.3 | 8.6 | 7.9 | 7.6 | 104.7 | 105.1 | EF, AJ, OJ | AD | GTR1_20230519_OF_01 | 2023-05-19 | 1.0023 | 0.99899 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230522 | 1.712 | -0.2331% | -0.3492% | 0.32 | 0.32 | 0.15 | 0.12 | 0.63 | 0.15 | 0.13 | 1.712 | 8.7 | 8.3 | 8.1 | 7.8 | 104.7 | 105 | AJ, OJ | GB | GTR1_20230522_OF_01 | 2023-05-22 | 0.99767 | 0.99892 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230523 | 1.714 | -0.1166% | -0.2328% | 0.33 | 0.27 | 0.1 | 0.1 | 0.67 | 0.16 | 0.14 | 1.714 | 8.6 | 8.3 | 8 | 7.8 | 104.7 | 105 | AE | CVD | GTR1_20230523_OF_01 | 2023-05-23 | 0.99883 | 0.9989 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230524 | 1.712 | -0.2331% | -0.3492% | 0.38 | 0.35 | 0.16 | 0.12 | 0.57 | 0.14 | 0.12 | 1.712 | 9.1 | 7.5 | 8 | 7.9 | 104.6 | 105 | JR | AD | GTR1_20230524_OF_01 | 2023-05-24 | 0.99767 | 0.99892 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230525 | 1.717 | 0.05828% | -0.05821% | 0.36 | 0.41 | 0.13 | 0.13 | 0.67 | 0.14 | 0.12 | 1.717 | 8.4 | 8.5 | 8 | 7.9 | 104.7 | 105 | GK | GB | GTR1_20230525_OF_01 | 2023-05-25 | 1.0006 | 0.9989 | 0.999 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230526 | 1.715 | -0.05828% | -0.1746% | 0.39 | 0.37 | 0.16 | 0.12 | 0.64 | 0.16 | 0.13 | 1.715 | 8.8 | 8.1 | 8.1 | 7.7 | 104.7 | 105 | GK | AC | GTR1_20230526_OF_01 | 2023-05-26 | 0.99942 | 0.99886 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230529 | 1.712 | -0.2331% | -0.3492% | 0.24 | 0.36 | 0.06 | 0.12 | 0.56 | 0.1 | 0.13 | 1.712 | 8.6 | 8.2 | 8.2 | 7.6 | 104.7 | 105 | GK,SF | CVD | GTR1_20230529_OF_01 | 2023-05-29 | 0.99767 | 0.99888 | 0.99898 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230530 | 1.71 | -0.3497% | -0.4657% | 0.25 | 0.36 | 0.06 | 0.16 | 0.65 | 0.11 | 0.12 | 1.71 | 8.3 | 8.6 | 8.1 | 7.6 | 104.7 | 105 | SM, JR | YM | GTR1_20230530_OF_01 | 2023-05-30 | 0.9965 | 0.99883 | 0.99894 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230531 | 1.716 | 0% | -0.1164% | 0.28 | 0.35 | 0.13 | 0.15 | 0.65 | 0.15 | 0.18 | 1.716 | 8.6 | 8.3 | 8 | 7.7 | 104.7 | 105 | GF, SF | YM | GTR1_20230531_OF_01 | 2023-05-31 | 1 | 0.99876 | 0.99896 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230601 | 1.713 | -0.1748% | -0.291% | 0.28 | 0.37 | 0.13 | 0.15 | 0.66 | 0.15 | 0.16 | 1.713 | 8.3 | 8.7 | 8.1 | 7.7 | 104.7 | 105 | ON | AD | GTR1_20230601_OF_01 | 2023-06-01 | 0.99825 | 0.99878 | 0.999 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230602 | 1.702 | -0.8159% | -0.9313% | 0.32 | 0.35 | 0.1 | 0.14 | 0.65 | 0.13 | 0.15 | 1.702 | 8.7 | 8.2 | 8.3 | 7.3 | 104.7 | 105 | EF | AC | GTR1_20230602_OF_01 | 2023-06-02 | 0.99184 | 0.99878 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230605 | 1.715 | -0.05828% | -0.1746% | 0.4 | 0.36 | 0.18 | 0.17 | 0.64 | 0.16 | 0.13 | 1.715 | 8.4 | 8.5 | 8 | 7.8 | 104.7 | 105 | AR | AC/KW | GTR1_20230605_OF_01 | 2023-06-05 | 0.99942 | 0.99881 | 0.99892 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230607 | 1.713 | -0.1748% | -0.291% | 0.33 | 0.26 | 0.15 | 0.11 | 0.6 | 0.16 | 0.13 | 1.713 | 8.6 | 8.3 | 7.4 | 8.3 | 104.6 | 105 | AJ, EF | YM | GTR1_20230607_OF_01 | 2023-06-07 | 0.99825 | 0.99877 | 0.9989 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230608 | 1.715 | -0.05828% | -0.1746% | 0.31 | 0.39 | 0.11 | 0.16 | 0.57 | 0.09 | 0.12 | 1.715 | 8.4 | 8.5 | 8.2 | 7.5 | 104.7 | 105 | EF | GB | GTR1_20230608_OF_01 | 2023-06-08 | 0.99942 | 0.99875 | 0.99892 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230609 | 1.707 | -0.5245% | -0.6403% | 0.33 | 0.36 | 0.15 | 0.09 | 0.72 | 0.18 | 0.12 | 1.707 | 8.8 | 8.1 | 8 | 7.7 | 104.6 | 105 | ON | AC | GTR1_20230609_OF_01 | 2023-06-09 | 0.99476 | 0.99875 | 0.9989 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230612 | 1.708 | -0.4662% | -0.5821% | 0.33 | 0.35 | 0.13 | 0.13 | 0.76 | 0.16 | 0.17 | 1.708 | 8.2 | 8.7 | 7.9 | 7.8 | 104.6 | 105 | EF | AC | GTR1_20230612_OF_01 | 2023-06-12 | 0.99534 | 0.99875 | 0.99886 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230613 | 1.713 | -0.1748% | -0.291% | 0.37 | 0.38 | 0.12 | 0.19 | 0.61 | 0.15 | 0.15 | 1.713 | 8.6 | 8.3 | 8.1 | 7.7 | 104.7 | 105 | EF | GB | GTR1_20230613_OF_01 | 2023-06-13 | 0.99825 | 0.99878 | 0.99883 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230614 | 1.713 | -0.1748% | -0.291% | 0.39 | 0.35 | 0.16 | 0.14 | 0.66 | 0.18 | 0.16 | 1.713 | 8.6 | 8.3 | 7.9 | 7.9 | 104.7 | 105 | GK | YM | GTR1_20230614_OF_01 | 2023-06-14 | 0.99825 | 0.99883 | 0.99889 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230615 | 1.715 | -0.05828% | -0.1746% | 0.4 | 0.28 | 0.16 | 0.15 | 0.65 | 0.16 | 0.12 | 1.715 | 8.5 | 8.4 | 8.1 | 7.6 | 104.7 | 105 | ON | YM | GTR1_20230615_OF_01 | 2023-06-15 | 0.99942 | 0.99906 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230616 | 1.706 | -0.5828% | -0.6985% | 0.3 | 0.39 | 0.18 | 0.13 | 0.59 | 0.14 | 0.1 | 1.706 | 8.3 | 8.7 | 8.3 | 7.4 | 104.7 | 105 | EN | AC | GTR1_20230616_OF_01 | 2023-06-16 | 0.99417 | 0.99903 | 0.99907 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230619 | 1.736 | 1.166% | 1.048% | 0.33 | 0.46 | 0.1 | 0.22 | 0.84 | 0.19 | 0.22 | 1.736 | 8.9 | 7.8 | 7.8 | 8 | 104.6 | 105 | GK | ML | GTR1_20230619_OF_01 | 2023-06-19 | 1.0117 | 0.99899 | 0.99909 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230620 | 1.711 | -0.2914% | -0.4075% | 0.39 | 0.38 | 0.19 | 0.18 | 0.78 | 0.17 | 0.2 | 1.711 | 8.6 | 8.3 | 8 | 7.8 | 104.6 | 105 | EN | GB | GTR1_20230620_OF_01 | 2023-06-20 | 0.99709 | 0.99897 | 0.99906 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230621 | 1.71 | -0.3497% | -0.4657% | 0.32 | 0.39 | 0.13 | 0.18 | 0.67 | 0.13 | 0.19 | 1.71 | 8.9 | 7.9 | 8 | 7.7 | 104.6 | 105 | ON | CVD | GTR1_20230621_OF_01 | 2023-06-21 | 0.9965 | 0.99897 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230622 | 1.716 | 0% | -0.1164% | 0.28 | 0.38 | 0.12 | 0.16 | 0.61 | 0.1 | 0.16 | 1.716 | 8.1 | 8.7 | 8.1 | 7.5 | 104.7 | 105 | OJ | ML | GTR1_20230622_OF_01 | 2023-06-22 | 1 | 0.99899 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230626 | 1.721 | 0.2914% | 0.1746% | 0.28 | 0.38 | 0.08 | 0.14 | 0.49 | 0.07 | 0.15 | 1.721 | 8.8 | 7.9 | 7.9 | 7.7 | 104.6 | 105 | OJ | ML, YM | GTR1_20230626_OF_01 | 2023-06-26 | 1.0029 | 0.99904 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230627 | 1.717 | 0.05828% | -0.05821% | 0.28 | 0.4 | 0.07 | 0.14 | 0.6 | 0.15 | 0.12 | 1.717 | 8.9 | 7.9 | 7.9 | 7.9 | 104.6 | 105 | GK | AC | GTR1_20230627_OF_01 | 2023-06-27 | 1.0006 | 0.99904 | 0.99909 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230628 | 1.718 | 0.1166% | 0% | 0.38 | 0.33 | 0.16 | 0.08 | 0.68 | 0.18 | 0.13 | 1.718 | 8.8 | 8 | 7.9 | 7.7 | 104.6 | 105 | EN | AD | GTR1_20230628_OF_01 | 2023-06-28 | 1.0012 | 0.99901 | 0.99909 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230629 | 1.712 | -0.2331% | -0.3492% | 0.28 | 0.29 | 0.09 | 0.12 | 0.62 | 0.16 | 0.12 | 1.712 | 8.4 | 8.4 | 8.1 | 7.6 | 104.7 | 105 | OJ | ML | GTR1_20230629_OF_01 | 2023-06-29 | 0.99767 | 0.99901 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230630 | 1.715 | -0.05828% | -0.1746% | 0.29 | 0.44 | 0.14 | 0.18 | 0.7 | 0.16 | 0.2 | 1.715 | 8.7 | 8.1 | 7.9 | 7.9 | 104.6 | 105 | AR | ML | GTR1_20230630_OF_01 | 2023-06-30 | 0.99942 | 0.99901 | 0.99908 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230703 | 1.712 | -0.2331% | -0.3492% | 0.34 | 0.38 | 0.13 | 0.1 | 0.63 | 0.15 | 0.11 | 1.712 | 8.6 | 8.3 | 8.2 | 7.4 | 104.7 | 105 | OJ | AC YM | GTR1_20230703_OF_01 | 2023-07-03 | 0.99767 | 0.999 | 0.99907 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230704 | 1.711 | -0.2914% | -0.4075% | 0.32 | 0.33 | 0.13 | 0.12 | 0.58 | 0.16 | 0.13 | 1.711 | 8.7 | 8.2 | 8 | 7.7 | 104.6 | 105.1 | AR | AC | GTR1_20230704_OF_01 | 2023-07-04 | 0.99709 | 0.99904 | 0.99906 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230705 | 1.71 | -0.3497% | -0.4657% | 0.37 | 0.34 | 0.19 | 0.18 | 0.72 | 0.18 | 0.2 | 1.71 | 8.6 | 8.4 | 8.1 | 7.5 | 104.6 | 105 | ON | AD | GTR1_20230705_OF_01 | 2023-07-05 | 0.9965 | 0.99903 | 0.99903 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230706 | 1.712 | -0.2331% | -0.3492% | 0.33 | 0.33 | 0.12 | 0.14 | 0.7 | 0.15 | 0.16 | 1.712 | 8.7 | 8.3 | 8 | 7.7 | 104.7 | 105.1 | ON | AD | GTR1_20230706_OF_01 | 2023-07-06 | 0.99767 | 0.99896 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230707 | 1.702 | -0.8159% | -0.9313% | 0.34 | 0.39 | 0.16 | 0.15 | 0.69 | 0.17 | 0.15 | 1.702 | 8.5 | 8.4 | 8.3 | 7.5 | 104.7 | 105 | AR | ML | GTR1_20230707_OF_01 | 2023-07-07 | 0.99184 | 0.99898 | 0.99897 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230710 | 1.712 | -0.2331% | -0.3492% | 0.42 | 0.25 | 0.19 | 0.1 | 0.6 | 0.16 | 0.13 | 1.712 | 8.9 | 8 | 7.5 | 8.1 | 104.6 | 105 | AR | AC | GTR1_20230710_OF_01 | 2023-07-10 | 0.99767 | 0.99894 | 0.99897 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230711 | 1.714 | -0.1166% | -0.2328% | 0.36 | 0.41 | 0.16 | 0.15 | 0.66 | 0.16 | 0.15 | 1.714 | 8.7 | 8.2 | 7.9 | 7.9 | 104.6 | 105 | EN | YM | GTR1_20230711_OF_01 | 2023-07-11 | 0.99883 | 0.99896 | 0.99899 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230712 | 1.711 | -0.2914% | -0.4075% | 0.31 | 0.38 | 0.16 | 0.16 | 0.65 | 0.13 | 0.18 | 1.711 | 8.5 | 8.3 | 8.1 | 7.4 | 104.7 | 105.1 | OJ | YM | GTR1_20230712_OF_01 | 2023-07-12 | 0.99709 | 0.999 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230713 | 1.714 | -0.1166% | -0.2328% | 0.33 | 0.33 | 0.17 | 0.15 | 0.72 | 0.16 | 0.18 | 1.714 | 8.7 | 8.2 | 8 | 7.8 | 104.6 | 105 | OJ | AD | GTR1_20230713_OF_01 | 2023-07-13 | 0.99883 | 0.99899 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230714 | 1.72 | 0.2331% | 0.1164% | 0.35 | 0.36 | 0.22 | 0.18 | 0.69 | 0.2 | 0.16 | 1.72 | 8.4 | 8.5 | 7.9 | 7.6 | 104.6 | 105.1 | OJ | YM | GTR1_20230714_OF_01 | 2023-07-14 | 1.0023 | 0.99892 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230717 | 1.711 | -0.2914% | -0.4075% | 0.49 | 0.4 | 0.27 | 0.17 | 0.73 | 0.22 | 0.19 | 1.711 | 8.6 | 8.4 | 8.1 | 7.6 | 104.7 | 105 | EN | AD | GTR1_20230717_OF_01 | 2023-07-17 | 0.99709 | 0.9989 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230718 | 1.718 | 0.1166% | 0% | 0.38 | 0.4 | 0.22 | 0.17 | 0.76 | 0.2 | 0.25 | 1.718 | 8.6 | 8.5 | 8.3 | 7.4 | 104.7 | 105.1 | AR | AC,AD | GTR1_20230718_OF_01 | 2023-07-18 | 1.0012 | 0.99892 | 0.99901 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230719 | 1.718 | 0.1166% | 0% | 0.36 | 0.36 | 0.15 | 0.14 | 0.62 | 0.16 | 0.15 | 1.718 | 8.7 | 8.2 | 7.8 | 7.8 | 104.7 | 105.1 | AR | AC,YM | GTR1_20230719_OF_01 | 2023-07-19 | 1.0012 | 0.9989 | 0.999 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230720 | 1.715 | -0.05828% | -0.1746% | 0.26 | 0.42 | 0.1 | 0.15 | 0.69 | 0.12 | 0.12 | 1.715 | 8.4 | 8.6 | 7.8 | 7.9 | 104.7 | 105 | ON | YM | GTR1_20230720_OF_01 | 2023-07-20 | 0.99942 | 0.99886 | 0.99904 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230721 | 1.717 | 0.05828% | -0.05821% | 0.26 | 0.42 | 0.09 | 0.12 | 0.75 | 0.13 | 0.13 | 1.717 | 8.7 | 8.1 | 8 | 7.6 | 104.7 | 105.1 | EN | YM, NIC | GTR1_20230721_OF_01 | 2023-07-21 | 1.0006 | 0.99883 | 0.99903 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230724 | 1.715 | -0.05828% | -0.1746% | 0.26 | 0.36 | 0.11 | 0.16 | 0.63 | 0.16 | 0.16 | 1.715 | 8.5 | 8.5 | 7.8 | 7.8 | 104.7 | 105.1 | ON | YM KW | GTR1_20230724_OF_01 | 2023-07-24 | 0.99942 | 0.99889 | 0.99896 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230725 | 1.721 | 0.2914% | 0.1746% | 0.29 | 0.41 | 0.08 | 0.18 | 0.69 | 0.14 | 0.18 | 1.721 | 8.4 | 8.5 | 7.9 | 7.9 | 104.7 | 105 | AR | YM KW | GTR1_20230725_OF_01 | 2023-07-25 | 1.0029 | 0.99888 | 0.99898 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230726 | 1.716 | 0% | -0.1164% | 0.32 | 0.37 | 0.14 | 0.16 | 0.65 | 0.15 | 0.14 | 1.716 | 8.5 | 8.3 | 7.9 | 7.7 | 104.7 | 105.1 | EN | YM | GTR1_20230726_OF_01 | 2023-07-26 | 1 | 0.99883 | 0.99894 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230727 | 1.717 | 0.05828% | -0.05821% | 0.33 | 0.38 | 0.14 | 0.15 | 0.77 | 0.19 | 0.18 | 1.717 | 8.7 | 8.3 | 8 | 7.7 | 104.6 | 105 | EN | YM KW | GTR1_20230727_OF_01 | 2023-07-27 | 1.0006 | 0.99876 | 0.99896 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230728 | 1.716 | 0% | -0.1164% | 0.32 | 0.39 | 0.17 | 0.17 | 0.65 | 0.12 | 0.18 | 1.716 | 8.2 | 8.6 | 8 | 7.7 | 104.7 | 105 | EN | YM CMN | GTR1_20230728_OF_01 | 2023-07-28 | 1 | 0.99878 | 0.999 | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
Formulas:
Date = DATE(LEFT([Parameter],4), MID([Parameter],5,2), RIGHT([Parameter],2))

## GTR1_Cubes
ID.nr:
Titel: | [52 empty] | Dokumenttyp:
Protocol for Daily QA: Consistency check of single layer and cubic plans dosimetric parameters | [52 empty] | Strålskyddsinstruktion
Godkänt av: | [51 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [51 empty] | Skapat den:
Reviderat av: | [51 empty] | Reviderat den:
GTR1
Daily cubes parameters
MatriXX PT_2
Set-up: |  | MatriXX (MatrixPT_2 - position adjusted using X-rays (chambers at iso))
Detector: |  | IBA MatrixPT_2
Phantom: |  | SW
MyQA Project: |  | GTR1_DailyQA_Matrixx
Patient in AdaptDelivery: |  | G1_Daily_MatriXX Skandion
Plan ID: |  | Cube_R15_SOBP10; Cube_R31_SOBP10
Parameters and its criterias
Plan ID | Parameter | Value (Gy) | Good: within 2% |  | Warning: over 2% |  | Action level: over 3%
R15M10 | Ref. dose (Roos IC SW)* | 2.006 | 2.0461 | 1.9659 | 2.0471 | 1.9649 | 2.0672 | 1.9448 |  | -0.5946%
R31M10 | Ref. dose (Roos IC SW)* | 1.997 | 2.0369 | 1.9571 | 2.0379 | 1.9561 | 2.0579 | 1.9361 |  | -1.09%
R15M10 | Dose from TPS | 2.018 | 2.0584 | 1.9776 | 2.0594 | 1.9766 | 2.0795 | 1.9565
R31M10 | Dose from TPS | 2.019 | 2.0584 | 1.9776 | 2.0594 | 1.9766 | 2.0795 | 1.9565
Plan ID | Parameter | Value (mm) | Good |  | +/- 1.001 mm |  | +/- 2.001 mm
R15M10 | Penumbra (average) | 8.9306 | 9.9306 | 7.9306 | 9.9316 | 7.9296 | 10.9316 | 6.9296
R31M10 | Penumbra (average) | 11.4306 | 12.4306 | 10.4306 | 12.4316 | 10.4296 | 13.4316 | 9.4296
Plan ID | Parameter | Value (mm) | Good |  | +/- 1.001 mm |  | +/- 2.001 mm
R15M10 | FWHM (average) | 119.7722 | 120.7722 | 118.7722 | 120.7732 | 118.7712 | 121.7732 | 117.7712
R31M10 | FWHM (average) | 119.5333 | 120.5333 | 118.5333 | 120.5343 | 118.5323 | 121.5343 | 117.5323
Plan ID | Parameter |  | Good: within +/-0.5% |  | Warning: within +/-1% |  | Action level: over 1%
| Symmetry | x | 0.5 | -0.5 | 0.5 | -0.5 | 1 | 1
Plan ID |  |  | Good: within +/-1% |  | Warning: within +/-2% |  | Action level: over 3%
| Flatness | x | 1 | -1 | 2 | -2 | 3 | -3
* reference measurements, Roos IC #2519, 04.11.2022
| R15M10 |  |  | R31M10 |  |  |  | R15_SOBP10_2Gy, 10 cm WET, 9.2 SW | [13 empty] | R31_SOBP10_2Gy , 26 cm WET, 24.8 SW | [17 empty] | R15_SOBP10_2Gy |  | R31_SOBP10_2Gy |  | Plot
Date | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | R31M10 vs. R15M10 | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Sign | Checked | File name | Notes | GI (1%, 1 mm) | GI (2%, 2 mm) | GI (1%, 1 mm) | GI (2%, 2 mm) | Date | Difference from reference (a.u.) R15M10 | Medelvärde +/- 10 dygn R15M10 | Medelvärde +/- 20 dygn R15M10 | Difference from reference (a.u.) R31M10 | Medelvärde +/- 10 dygn R31M10 | Medelvärde +/- 20 dygn R31M10 | Ref | 2% | -2% | 3% | -3%
20221111 | 1.996 | -0.4985% | -1.09% | 1.987 | -0.5008% | -1.585% | 1.0045 | 0.58 | 0.52 | 0.37 | 0.24 | 0.95 | 0.29 | 0.24 | 1.996 | 8.7 | 9.4 | 8.8 | 8.6 | 119.6 | 120 | 0.45 | 0.16 | 0.23 | 0.01 | 0.63 | 0.19 | 0.02 | 1.987 | 11.4 | 11.5 | 11.4 | 11.1 | 119.4 | 119.7 | ML | x | GTR1_20221111_R15M10_01, GTR1_20221111_R31M10_01 | REFERENCE images |  |  |  |  | 2022-11-11 | 0.99501 | 0.99761 | 0.99823 | 0.99499 | 0.9985 | 0.99892 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20221209 | 2.005 | -0.04985% | -0.6442% | 1.999 | 0.1002% | -0.9906% | 1.003 | 0.53 | 0.47 | 0.28 | 0.25 | 0.94 | 0.25 | 0.3 | 2.005 | 9 | 9.2 | 8.7 | 8.7 | 119.6 | 119.9 | 0.31 | 0.2 | 0.13 | -0.03 | 0.61 | 0.16 | 0.02 | 1.999 | 11.3 | 11.4 | 11.3 | 11.1 | 119.4 | 119.7 | ML | x | GTR1_20221209_R15M10_01, GTR1_20221209_R31M10_01 |  | 100 |  | 100 |  | 2022-12-09 | 0.9995 | 0.99751 | 0.99832 | 1.001 | 0.9985 | 0.99897 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230214 | 2.011 | 0.2493% | -0.3469% | 2.005 | 0.4006% | -0.6934% | 1.003 | 0.35 | 0.58 | 0.23 | 0.22 | 1.02 | 0.22 | 0.27 | 2.011 | 9.5 | 8.7 | 8.6 | 8.9 | 119.6 | 120 | 0.26 | 0.16 | 0.07 | 0 | 0.54 | 0.11 | 0 | 2.005 | 11.5 | 11.4 | 11.6 | 11.3 | 119.4 | 119.7 | AJ | ML | GTR1_20230214_R15M10_01, GTR1_20230214_R31M10_01 |  | 88.5 | 100 | 84.7 | 100 | 2023-02-14 | 1.0025 | 0.99788 | 0.99826 | 1.004 | 0.99887 | 0.99887 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230214 | 2.011 | 0.2493% | -0.3469% | 2.005 | 0.4006% | -0.6934% | 1.003 | 0.35 | 0.58 | 0.23 | 0.22 | 1.02 | 0.22 | 0.27 | 2.011 | 9.5 | 8.7 | 8.6 | 8.9 | 119.6 | 120 | 0.26 | 0.16 | 0.07 | 0 | 0.54 | 0.11 | 0 | 2.005 | 11.5 | 11.4 | 11.6 | 11.3 | 119.4 | 119.7 | AE | ML | GTR1_20230217_R15M10_01, GTR1_20230217_R31M10_01, GTR1_20230214_R15M10_01, GTR1_20230214_R31M10_01 |  | 100 |  | 99.7 |  | 2023-02-14 | 1.0025 | 0.99793 | 0.99817 | 1.004 | 0.99884 | 0.99878 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230217 | 2.005 | -0.04985% | -0.6442% | 2.001 | 0.2003% | -0.8915% | 1.002 | 0.56 | 0.45 | 0.31 | 0.24 | 0.95 | 0.26 | 0.27 | 2.005 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 119.9 | 0.33 | 0.12 | 0.11 | 0 | 0.55 | 0.12 | 0.01 | 2.001 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | EH SR OJ | ML | GTR1_20230217_R15M10_01, GTR1_20230217_R31M10_01 |  | 100 |  | 99.7 |  | 2023-02-17 | 0.9995 | 0.99793 | 0.99779 | 1.002 | 0.99884 | 0.99841 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230221 | 2.001 | -0.2493% | -0.8424% | 1.994 | -0.1502% | -1.238% | 1.0035 | 0.52 | 0.36 | 0.25 | 0.16 | 0.88 | 0.23 | 0.18 | 2.001 | 9.1 | 9.2 | 9 | 8.5 | 119.6 | 120 | 0.37 | 0.13 | 0.11 | 0.01 | 0.62 | 0.11 | 0.03 | 1.994 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | SM, GF, AR, OJ | ML | GTR1_20230221_R15M10_01, GTR1_20230221_R31M10_01 | Remeber to record the GI vaules | 100 |  | 100 |  | 2023-02-21 | 0.99751 | 0.99808 | 0.99774 | 0.9985 | 0.99889 | 0.99831 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230222 | 1.998 | -0.3988% | -0.9911% | 1.99 | -0.3505% | -1.436% | 1.004 | 0.44 | 0.46 | 0.16 | 0.23 | 0.89 | 0.2 | 0.26 | 1.998 | 9.1 | 9.2 | 8.9 | 8.6 | 119.6 | 119.9 | 0.3 | 0.16 | 0.07 | 0 | 0.59 | 0.11 | 0.03 | 1.99 | 11.5 | 11.6 | 11.6 | 11.5 | 119.3 | 119.7 | AR | ML | GTR1_20230222_R15M10_01, GTR1_20230222_R31M10_01 |  | 100 |  | 100 |  | 2023-02-22 | 0.99601 | 0.99811 | 0.99765 | 0.99649 | 0.99886 | 0.9982 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230223 | 2 | -0.2991% | -0.892% | 1.993 | -0.2003% | -1.288% | 1.0035 | 0.46 | 0.41 | 0.18 | 0.22 | 0.89 | 0.2 | 0.24 | 2 | 9.1 | 9.2 | 8.7 | 8.8 | 119.6 | 119.9 | 0.29 | 0.14 | 0.05 | 0.02 | 0.55 | 0.11 | 0.01 | 1.993 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | OJ | ML | GTR1_20230223_R15M10_01, GTR1_20230223_R31M10_01 |  | 100 |  | 100 |  | 2023-02-23 | 0.99701 | 0.99822 | 0.99762 | 0.998 | 0.99894 | 0.99811 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230224 | 1.996 | -0.4985% | -1.09% | 1.985 | -0.6009% | -1.684% | 1.0055 | 0.47 | 0.38 | 0.26 | 0.18 | 0.83 | 0.23 | 0.22 | 1.996 | 9.3 | 9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.25 | 0.18 | 0.06 | 0.03 | 0.48 | 0.09 | 0.01 | 1.985 | 11.4 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | EH | AD | GTR1_20230224_R15M10_01, GTR1_20230224_R31M10_01 |  | 98.1 |  | 100 |  | 2023-02-24 | 0.99501 | 0.99824 | 0.99766 | 0.99399 | 0.99888 | 0.99813 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230227 | 1.989 | -0.8475% | -1.437% | 1.981 | -0.8012% | -1.882% | 1.004 | 0.48 | 0.41 | 0.3 | 0.18 | 0.86 | 0.25 | 0.21 | 1.989 | 8.2 | 9.7 | 8.7 | 8.8 | 119.5 | 120 | 0.37 | 0.17 | 0.13 | -0.05 | 0.54 | 0.16 | -0.04 | 1.981 | 11.9 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | AC | AC | GTR1_20230227_R15M10_01, GTR1_20230227_R31M10_01 |  | 99 |  | 95.6 |  | 2023-02-27 | 0.99153 | 0.99823 | 0.99772 | 0.99199 | 0.99892 | 0.99819 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230228 | 1.999 | -0.349% | -0.9415% | 1.994 | -0.1502% | -1.238% | 1.0025 | 0.41 | 0.43 | 0.18 | 0.22 | 0.91 | 0.2 | 0.25 | 1.999 | 9.2 | 9 | 8.8 | 8.5 | 119.6 | 120 | 0.36 | 0.21 | 0.14 | 0.03 | 0.61 | 0.13 | -0.02 | 1.994 | 11.4 | 11.4 | 11.4 | 11.1 | 119.4 | 119.8 | AR, AJ | YM | GTR1_20230228_R15M10_01, GTR1_20230228_R31M10_01 |  | 99.7 |  | 100 |  | 2023-02-28 | 0.99651 | 0.9985 | 0.99754 | 0.9985 | 0.99919 | 0.99801 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230301 | 2.01 | 0.1994% | -0.3964% | 2.003 | 0.3005% | -0.7925% | 1.0035 | 0.4 | 0.46 | 0.16 | 0.23 | 0.87 | 0.19 | 0.24 | 2.01 | 9.1 | 9.2 | 8.8 | 8.6 | 119.6 | 120 | 0.29 | 0.15 | 0.05 | 0.01 | 0.57 | 0.09 | 0.03 | 2.003 | 11.4 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | AR, JR | GB | GTR1_20230301_R15M10_01, GTR1_20230301_R31M10_01 |  | 97.1 |  | 96.9 |  | 2023-03-01 | 1.002 | 0.99837 | 0.99757 | 1.003 | 0.99897 | 0.99798 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230302 | 2.003 | -0.1496% | -0.7433% | 1.994 | -0.1502% | -1.238% | 1.0045 | 0.43 | 0.46 | 0.15 | 0.25 | 0.78 | 0.17 | 0.22 | 2.003 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.28 | 0.19 | 0.08 | -0.02 | 0.46 | 0.09 | -0.01 | 1.994 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | SM, SF | YM | GTR1_20230302_R15M10_01, GTR1_20230302_R31M10_01 |  | 93 | 100 | 99.7 | 100 | 2023-03-02 | 0.9985 | 0.99803 | 0.99757 | 0.9985 | 0.99858 | 0.99792 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230303 | 2.006 | 0% | -0.5946% | 1.996 | -0.05008% | -1.139% | 1.005 | 0.41 | 0.48 | 0.23 | 0.24 | 0.9 | 0.24 | 0.25 | 2.006 | 9 | 9.3 | 8.8 | 8.6 | 119.6 | 119.9 | 0.3 | 0.12 | 0.07 | -0.02 | 0.53 | 0.11 | 0 | 1.996 | 11.5 | 11.6 | 11.5 | 11.4 | 119.3 | 119.7 | OJ | AC | GTR1_20230303_R15M10_01, GTR1_20230303_R31M10_01 |  | 100 |  | 100 |  | 2023-03-03 | 1 | 0.99719 | 0.99762 | 0.9995 | 0.99755 | 0.99789 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230306 | 2.003 | -0.1496% | -0.7433% | 1.994 | -0.1502% | -1.238% | 1.0045 | 0.48 | 0.47 | 0.23 | 0.16 | 0.89 | 0.2 | 0.2 | 2.003 | 9.1 | 9.2 | 8.6 | 8.9 | 119.6 | 119.9 | 0.33 | 0.18 | 0.09 | -0.08 | 0.52 | 0.1 | -0.04 | 1.994 | 11.5 | 11.6 | 11.5 | 11.3 | 119.3 | 119.6 | AD, GB | ML | GTR1_20230306_R15M10_02, GTR1_20230306_R31M10_02 | run 2, after ion source replacement | ¨100 |  | 100 |  | 2023-03-06 | 0.9985 | 0.99709 | 0.99761 | 0.9985 | 0.99739 | 0.99788 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230307 | 2.006 | 0% | -0.5946% | 1.997 | 0% | -1.09% | 1.0045 | 0.44 | 0.41 | 0.18 | 0.18 | 0.85 | 0.21 | 0.23 | 2.006 | 9.2 | 9.1 | 8.8 | 8.6 | 119.6 | 120 | 0.3 | 0.13 | 0.09 | -0.05 | 0.46 | 0.09 | -0.02 | 1.997 | 11.4 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | OJ | YM KW | GTR1_20230307_R15M10_01, GTR1_20230307_R31M10_01 |  | 100 |  | 100 |  | 2023-03-07 | 1 | 0.99714 | 0.99765 | 1 | 0.99736 | 0.99791 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230308 | 2.003 | -0.1496% | -0.7433% | 1.993 | -0.2003% | -1.288% | 1.005 | 0.44 | 0.42 | 0.22 | 0.2 | 0.85 | 0.17 | 0.22 | 2.003 | 9.3 | 8.9 | 9 | 8.3 | 119.6 | 120 | 0.35 | 0.12 | 0.12 | -0.02 | 0.53 | 0.1 | -0.01 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.3 | 119.7 | JR | YM | GTR1_20230308_R15M10_02, GTR1_20230308_R31M10_02 | After beam alignment | 92.1 | 100 | 100 | 100 | 2023-03-08 | 0.9985 | 0.99722 | 0.9977 | 0.998 | 0.99739 | 0.99796 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230309 | 2.002 | -0.1994% | -0.7929% | 1.996 | -0.05008% | -1.139% | 1.003 | 0.45 | 0.46 | 0.21 | 0.23 | 0.81 | 0.2 | 0.22 | 2.002 | 9.2 | 9 | 8.8 | 8.6 | 119.6 | 120 | 0.27 | 0.15 | 0.08 | -0.01 | 0.48 | 0.09 | 0 | 1.996 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | JR | YM | GTR1_20230309_R15M10_01, GTR1_20230309_R31M10_01 |  | 99.7 | 100 | 100 | 100 | 2023-03-09 | 0.99801 | 0.99746 | 0.99778 | 0.9995 | 0.99771 | 0.99805 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230310 | 2.006 | 0% | -0.5946% | 1.997 | 0% | -1.09% | 1.0045 | 0.33 | 0.63 | 0.18 | 0.29 | 0.87 | 0.16 | 0.29 | 2.006 | 9 | 9.2 | 8.6 | 8.8 | 119.6 | 120 | 0.27 | 0.16 | 0.07 | 0 | 0.47 | 0.08 | -0.01 | 1.997 | 11.3 | 11.5 | 11.4 | 11.2 | 119.4 | 119.7 | SM, AR | AD | GTR1_20230310_R15M10_01, GTR1_20230310_R31M10_01 |  | 98.1 | 100 | 97.8 | 100 | 2023-03-10 | 1 | 0.99751 | 0.99789 | 1 | 0.99776 | 0.99812 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230313 | 2 | -0.2991% | -0.892% | 1.991 | -0.3005% | -1.387% | 1.0045 | 0.39 | 0.46 | 0.13 | 0.25 | 0.76 | 0.16 | 0.23 | 2 | 9.3 | 9 | 8.8 | 8.8 | 119.6 | 119.9 | 0.2 | 0.11 | 0.03 | -0.01 | 0.5 | 0.09 | 0.01 | 1.991 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | SM, SF | AC | GTR1_20230313_R15M10_01, GTR1_20230313_R31M10_01 |  | 97.1 | 100 | 100 |  | 2023-03-13 | 0.99701 | 0.99761 | 0.99777 | 0.997 | 0.99768 | 0.99794 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230314 | 1.999 | -0.349% | -0.9415% | 1.991 | -0.3005% | -1.387% | 1.004 | 0.55 | 0.43 | 0.27 | 0.2 | 0.9 | 0.23 | 0.2 | 1.999 | 8.9 | 9.3 | 8.8 | 8.6 | 119.6 | 120 | 0.34 | 0.13 | 0.12 | -0.02 | 0.59 | 0.14 | 0.01 | 1.991 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | AE, AR | YM | GTR1_20230314_R15M10_01, GTR1_20230314_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-03-14 | 0.99651 | 0.99738 | 0.99761 | 0.997 | 0.99731 | 0.99773 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230315 | 1.995 | -0.5484% | -1.14% | 1.989 | -0.4006% | -1.486% | 1.003 | 0.42 | 0.4 | 0.18 | 0.2 | 0.94 | 0.2 | 0.24 | 1.995 | 9.3 | 9 | 8.7 | 8.8 | 119.6 | 120 | 0.27 | 0.16 | 0.06 | 0.01 | 0.52 | 0.09 | 0.03 | 1.989 | 11.5 | 11.4 | 11.5 | 11.2 | 119.3 | 119.7 | SM, AE | CVD | GTR1_20230315_R15M10_01, GTR1_20230315_R31M10_01 |  | 96.8 | 100 | 100 |  | 2023-03-15 | 0.99452 | 0.99732 | 0.99739 | 0.99599 | 0.99726 | 0.99747 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230316 | 1.992 | -0.6979% | -1.288% | 1.983 | -0.7011% | -1.783% | 1.0045 | 0.45 | 0.56 | 0.27 | 0.26 | 0.94 | 0.24 | 0.28 | 1.992 | 9.1 | 9.1 | 8.6 | 8.8 | 119.6 | 120 | 0.41 | 0.17 | 0.18 | 0.02 | 0.58 | 0.16 | -0.01 | 1.983 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | JR | ML | GTR1_20230316_R15M10_01, GTR1_20230316_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-03-16 | 0.99302 | 0.99736 | 0.99742 | 0.99299 | 0.99725 | 0.99747 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230320 | 1.999 | -0.349% | -0.9415% | 1.989 | -0.4006% | -1.486% | 1.005 | 0.43 | 0.42 | 0.2 | 0.17 | 0.74 | 0.18 | 0.2 | 1.999 | 9.2 | 9.1 | 8.6 | 8.9 | 119.6 | 119.9 | 0.3 | 0.12 | 0.08 | -0.02 | 0.53 | 0.09 | 0 | 1.989 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.7 | AC | AC | GTR1_20230320_R15M10_01, GTR1_20230320_R31M10_01 |  | 100 |  | 100 |  | 2023-03-20 | 0.99651 | 0.99744 | 0.99741 | 0.99599 | 0.99731 | 0.99738 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230321 | 1.997 | -0.4487% | -1.041% | 1.988 | -0.4507% | -1.535% | 1.0045 | 0.47 | 0.43 | 0.15 | 0.25 | 0.97 | 0.16 | 0.23 | 1.997 | 9.2 | 9.1 | 9 | 8.4 | 119.6 | 120 | 0.31 | 0.15 | -0.02 | 0.01 | 0.53 | 0.04 | 0.02 | 1.988 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | ML | ML | GTR1_20230321_R15M10_01, GTR1_20230321_R31M10_01 | ML corrected the analysis | 100 |  | 100 |  | 2023-03-21 | 0.99551 | 0.99729 | 0.99743 | 0.99549 | 0.99721 | 0.99737 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230322 | 2 | -0.2991% | -0.892% | 1.989 | -0.4006% | -1.486% | 1.0055 | 0.26 | 0.63 | 0.14 | 0.28 | 1.01 | 0.15 | 0.3 | 2 | 9.3 | 8.9 | 8.5 | 9 | 119.6 | 119.9 | 0.26 | 0.16 | 0.03 | 0.01 | 0.51 | 0.06 | 0.02 | 1.989 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | GK,SM | CVD | GTR1_20230322_R15M10_01, GTR1_20230322_R31M10_01 |  | 92.3 | 100 | 100 | 100 | 2023-03-22 | 0.99701 | 0.99732 | 0.99747 | 0.99599 | 0.99723 | 0.99738 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230323 | 2.003 | -0.1496% | -0.7433% | 1.994 | -0.1502% | -1.238% | 1.0045 | 0.51 | 0.45 | 0.27 | 0.23 | 0.89 | 0.24 | 0.26 | 2.003 | 9 | 9.3 | 8.8 | 8.6 | 119.6 | 120 | 0.4 | 0.13 | 0.15 | -0.03 | 0.56 | 0.14 | 0 | 1.994 | 11.7 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | EH | ML | GTR1_20230323_R15M10_01, GTR1_20230323_R31M10_01 |  | 100 |  | 100 |  | 2023-03-23 | 0.9985 | 0.99729 | 0.99747 | 0.9985 | 0.99721 | 0.99738 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230324 | 2.005 | -0.04985% | -0.6442% | 1.997 | 0% | -1.09% | 1.004 | 0.48 | 0.51 | 0.16 | 0.28 | 0.92 | 0.18 | 0.26 | 2.005 | 9.2 | 9 | 8.8 | 8.7 | 119.6 | 120 | 0.27 | 0.17 | 0.04 | 0.01 | 0.54 | 0.07 | 0.02 | 1.997 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | AE | KW | GTR1_20230324_R15M10_01, GTR1_20230324_R31M10_01 |  | 97.8 | 100 | 100 | 100 | 2023-03-24 | 0.9995 | 0.99741 | 0.99743 | 1 | 0.99738 | 0.99736 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230327 | 1.991 | -0.7478% | -1.338% | 1.983 | -0.7011% | -1.783% | 1.004 | 0.38 | 0.47 | 0.14 | 0.23 | 0.74 | 0.14 | 0.21 | 1.991 | 9.2 | 9 | 8.9 | 8.5 | 119.6 | 120 | 0.27 | 0.15 | 0.04 | 0 | 0.5 | 0.08 | 0.02 | 1.983 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | SM, AR | AC | GTR1_20230327_R15M10_01, GTR1_20230327_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-03-27 | 0.99252 | 0.99746 | 0.99747 | 0.99299 | 0.99728 | 0.9974 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230328 | 2.003 | -0.1496% | -0.7433% | 1.991 | -0.3005% | -1.387% | 1.006 | 0.5 | 0.46 | 0.23 | 0.23 | 0.95 | 0.22 | 0.24 | 2.003 | 9.3 | 8.9 | 8.7 | 8.9 | 119.6 | 119.9 | 0.28 | 0.15 | 0.07 | -0.02 | 0.53 | 0.1 | 0.01 | 1.991 | 11.5 | 11.4 | 11.6 | 11.4 | 119.4 | 119.7 | GK,ON,SF | YM | GTR1_20230328_R15M10_01, GTR1_20230328_R31M10_01 |  | 91.1 | 100 | 100 | 100 | 2023-03-28 | 0.9985 | 0.99722 | 0.99748 | 0.997 | 0.99702 | 0.99738 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230329 | 2.001 | -0.2493% | -0.8424% | 1.989 | -0.4006% | -1.486% | 1.006 | 0.43 | 0.44 | 0.18 | 0.22 | 0.9 | 0.19 | 0.25 | 2.001 | 9.3 | 9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.24 | 0.19 | 0.02 | 0.03 | 0.53 | 0.07 | 0.05 | 1.989 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.7 | JR | YM ON | GTR1_20230329_R15M10_01, GTR1_20230329_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-03-29 | 0.99751 | 0.9972 | 0.99742 | 0.99599 | 0.99697 | 0.99732 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230330 | 2.001 | -0.2493% | -0.8424% | 1.992 | -0.2504% | -1.337% | 1.0045 | 0.42 | 0.5 | 0.19 | 0.28 | 0.93 | 0.22 | 0.3 | 2.001 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 120 | 0.3 | 0.13 | 0.07 | -0.02 | 0.63 | 0.12 | 0.05 | 1.992 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | SM, OJ | ON | GTR1_20230330_R15M10_01, GTR1_20230330_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-03-30 | 0.99751 | 0.99708 | 0.99742 | 0.9975 | 0.99683 | 0.99731 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230331 | 2.002 | -0.1994% | -0.7929% | 1.991 | -0.3005% | -1.387% | 1.0055 | 0.45 | 0.42 | 0.16 | 0.21 | 0.76 | 0.15 | 0.24 | 2.002 | 9.3 | 9 | 8.7 | 8.9 | 119.6 | 119.9 | 0.27 | 0.14 | 0.06 | -0.05 | 0.54 | 0.07 | -0.01 | 1.991 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ON | GB | GTR1_20230331_R15M10_01, GTR1_20230331_R31M10_01 |  | 100 |  | 100 |  | 2023-03-31 | 0.99801 | 0.99727 | 0.99737 | 0.997 | 0.9969 | 0.99727 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230403 | 2.004 | -0.0997% | -0.6938% | 1.994 | -0.1502% | -1.238% | 1.005 | 0.39 | 0.59 | 0.24 | 0.29 | 0.97 | 0.2 | 0.32 | 2.004 | 9.3 | 8.9 | 8.6 | 8.9 | 119.6 | 119.9 | 0.34 | 0.13 | 0.09 | 0.01 | 0.61 | 0.11 | 0.02 | 1.994 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | OJ | YM | GTR1_20230403_R15M10_01, GTR1_20230403_R31M10_01 |  | 90.7 | 100 | 100 | 100 | 2023-04-03 | 0.999 | 0.99755 | 0.99759 | 0.9985 | 0.99716 | 0.99744 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230404 | 2 | -0.2991% | -0.892% | 1.992 | -0.2504% | -1.337% | 1.004 | 0.55 | 0.43 | 0.23 | 0.2 | 0.96 | 0.2 | 0.24 | 2 | 9.2 | 9.1 | 8.7 | 8.8 | 119.6 | 120 | 0.25 | 0.12 | 0.04 | 0 | 0.48 | 0.07 | 0.02 | 1.992 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | AE, GF | GB | GTR1_20230404_R15M10_01, GTR1_20230404_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-04-04 | 0.99701 | 0.99765 | 0.99765 | 0.9975 | 0.99726 | 0.99754 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230405 | 2.004 | -0.0997% | -0.6938% | 1.995 | -0.1002% | -1.189% | 1.0045 | 0.47 | 0.46 | 0.25 | 0.23 | 0.85 | 0.22 | 0.23 | 2.004 | 9.4 | 8.8 | 8.6 | 8.7 | 119.6 | 120 | 0.36 | 0.12 | 0.14 | -0.03 | 0.51 | 0.13 | -0.02 | 1.995 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.8 | ON, SM, JR | GB | GTR1_20230405_R15M10_01, GTR1_20230405_R31M10_01 |  | 91.1 | 100 | 100 | 100 | 2023-04-05 | 0.999 | 0.99757 | 0.99764 | 0.999 | 0.99735 | 0.99753 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230406 | 2.005 | -0.04985% | -0.6442% | 1.996 | -0.05008% | -1.139% | 1.0045 | 0.45 | 0.51 | 0.2 | 0.25 | 0.93 | 0.22 | 0.27 | 2.005 | 8.7 | 9.4 | 8.8 | 8.7 | 119.6 | 120 | 0.29 | 0.12 | 0.08 | 0 | 0.55 | 0.12 | 0.03 | 1.996 | 11.9 | 11.7 | 11.6 | 11.7 | 119.4 | 119.8 | OJ | CVD | GTR1_20230406_R15M10_01, GTR1_20230406_R31M10_01 |  | 100 |  | 100 |  | 2023-04-06 | 0.9995 | 0.99757 | 0.99762 | 0.9995 | 0.99739 | 0.9975 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230407 | 2.008 | 0.0997% | -0.4955% | 2 | 0.1502% | -0.9411% | 1.004 | 0.45 | 0.44 | 0.21 | 0.21 | 0.85 | 0.21 | 0.27 | 2.008 | 9.1 | 9.2 | 8.7 | 8.9 | 119.6 | 119.9 | 0.33 | 0.13 | 0.11 | 0 | 0.6 | 0.13 | 0.03 | 2 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | JR,SM | CVD | GTR1_20230407_R15M10_01, GTR1_20230407_R31M10_01 |  | 99.4 | 100 | 99.7 | 100 | 2023-04-07 | 1.001 | 0.99746 | 0.99761 | 1.0015 | 0.99729 | 0.9975 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230411 | 2.004 | -0.0997% | -0.6938% | 1.992 | -0.2504% | -1.337% | 1.006 | 0.41 | 0.43 | 0.18 | 0.21 | 0.85 | 0.19 | 0.26 | 2.004 | 9.1 | 9.1 | 8.9 | 8.7 | 119.6 | 119.9 | 0.34 | 0.17 | 0.13 | -0.01 | 0.59 | 0.13 | -0.01 | 1.992 | 11.5 | 11.6 | 11.6 | 11.6 | 119.4 | 119.7 | OJ, GK | GB | GTR1_20230411_R15M10_01, GTR1_20230411_R31M10_01 |  | 100 |  | 100 |  | 2023-04-11 | 0.999 | 0.99736 | 0.99761 | 0.9975 | 0.99718 | 0.99748 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230412 | 1.996 | -0.4985% | -1.09% | 1.986 | -0.5508% | -1.634% | 1.005 | 0.46 | 0.46 | 0.19 | 0.22 | 0.91 | 0.21 | 0.23 | 1.996 | 9.3 | 9 | 8.7 | 8.8 | 119.6 | 120 | 0.28 | 0.15 | 0.06 | -0.02 | 0.58 | 0.1 | 0.02 | 1.986 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | GK,EH | YM | GTR1_20230412_R15M10_01, GTR1_20230412_R31M10_01 |  | 95.5 | 100 | 100 | 100 | 2023-04-12 | 0.99501 | 0.99745 | 0.99758 | 0.99449 | 0.99727 | 0.99745 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230413 | 1.999 | -0.349% | -0.9415% | 1.989 | -0.4006% | -1.486% | 1.005 | 0.39 | 0.5 | 0.16 | 0.26 | 0.91 | 0.19 | 0.21 | 1.999 | 9.1 | 9.1 | 8.9 | 8.4 | 119.6 | 120 | 0.31 | 0.14 | 0.11 | 0 | 0.55 | 0.12 | -0.01 | 1.989 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | AR | YM | GTR1_20230413_R15M10_01, GTR1_20230413_R31M10_01 |  | 100 |  | 100 |  | 2023-04-13 | 0.99651 | 0.99744 | 0.99759 | 0.99599 | 0.9973 | 0.99745 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230414 | 1.994 | -0.5982% | -1.189% | 1.985 | -0.6009% | -1.684% | 1.0045 | 0.48 | 0.46 | 0.21 | 0.22 | 0.85 | 0.22 | 0.23 | 1.994 | 9.1 | 9.1 | 8.9 | 8.6 | 119.6 | 120 | 0.37 | 0.16 | 0.14 | -0.04 | 0.51 | 0.14 | -0.01 | 1.985 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | AR, JR | GB | GTR1_20230414_R15M10_01, GTR1_20230414_R31M10_01 |  | 100 |  | 100 |  | 2023-04-14 | 0.99402 | 0.99745 | 0.9976 | 0.99399 | 0.99733 | 0.99746 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230417 | 2.003 | -0.1496% | -0.7433% | 1.992 | -0.2504% | -1.337% | 1.0055 | 0.29 | 0.65 | 0.2 | 0.28 | 0.93 | 0.17 | 0.29 | 2.003 | 9.3 | 9 | 8.5 | 8.9 | 119.6 | 119.9 | 0.29 | 0.18 | 0.06 | -0.04 | 0.52 | 0.09 | 0 | 1.992 | 11.5 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | GK,SM | GB | GTR1_20230417_R15M10_01, GTR1_20230417_R31M10_01 |  | 93.3 | 100 | 100 | 100 | 2023-04-17 | 0.9985 | 0.99739 | 0.99763 | 0.9975 | 0.99727 | 0.99747 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230418 | 2.004 | -0.0997% | -0.6938% | 1.994 | -0.1502% | -1.238% | 1.005 | 0.49 | 0.53 | 0.18 | 0.23 | 0.92 | 0.17 | 0.19 | 2.004 | 9.1 | 9.1 | 8.9 | 8.6 | 119.6 | 120 | 0.22 | 0.15 | 0.04 | 0.02 | 0.56 | 0.06 | 0.03 | 1.994 | 11.5 | 11.5 | 11.5 | 11.5 | 119.4 | 119.8 | AJ | GB | GTR1_20230418_R15M10_01, GTR1_20230418_R31M10_01 |  | 100 |  | 100 |  | 2023-04-18 | 0.999 | 0.99768 | 0.99768 | 0.9985 | 0.99752 | 0.99752 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230419 | 2.003 | -0.1496% | -0.7433% | 1.993 | -0.2003% | -1.288% | 1.005 | 0.44 | 0.5 | 0.23 | 0.28 | 0.85 | 0.2 | 0.25 | 2.003 | 9.4 | 8.8 | 8.8 | 8.7 | 119.6 | 119.9 | 0.37 | 0.14 | 0.16 | 0.03 | 0.56 | 0.12 | 0 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | JR | CVD | GTR1_20230419_R15M10_01, GTR1_20230419_R31M10_01 |  | 100 |  | 100 |  | 2023-04-19 | 0.9985 | 0.99774 | 0.9977 | 0.998 | 0.99762 | 0.99754 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230420 | 2 | -0.2991% | -0.892% | 1.992 | -0.2504% | -1.337% | 1.004 | 0.45 | 0.56 | 0.2 | 0.3 | 0.83 | 0.18 | 0.29 | 2 | 8.9 | 9.3 | 8.9 | 8.5 | 119.6 | 120 | 0.33 | 0.2 | 0.07 | 0.01 | 0.61 | 0.13 | 0.1 | 1.992 | 11.6 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | JR,AR | CVD | GTR1_20230420_R15M10_01, GTR1_20230420_R31M10_01 |  | 100 |  | 100 |  | 2023-04-20 | 0.99701 | 0.99775 | 0.99772 | 0.9975 | 0.99762 | 0.99756 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230421 | 2.006 | 0% | -0.5946% | 1.993 | -0.2003% | -1.288% | 1.0065 | 0.53 | 0.46 | 0.23 | 0.21 | 0.91 | 0.21 | 0.24 | 2.006 | 9.4 | 8.8 | 8.7 | 8.7 | 119.6 | 120 | 0.26 | 0.11 | 0.06 | 0 | 0.57 | 0.09 | 0.01 | 1.993 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.8 | GK,AR | CVD | GTR1_20230421_R15M10_01, GTR1_20230421_R31M10_01 |  | 91.1 | 100 | 99.7 | 100 | 2023-04-21 | 1 | 0.99773 | 0.99773 | 0.998 | 0.99761 | 0.99758 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230424 | 1.996 | -0.4985% | -1.09% | 1.989 | -0.4006% | -1.486% | 1.0035 | 0.47 | 0.49 | 0.18 | 0.23 | 0.84 | 0.19 | 0.22 | 1.996 | 9.1 | 9.2 | 8.9 | 8.5 | 119.6 | 120 | 0.28 | 0.17 | 0.06 | -0.02 | 0.56 | 0.1 | 0.02 | 1.989 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | AR, AJ | AC,YM | GTR1_20230424_R15M10_01, GTR1_20230424_R31M10_01 |  | 100 |  | 100 |  | 2023-04-24 | 0.99501 | 0.99771 | 0.99772 | 0.99599 | 0.99758 | 0.99757 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230425 | 2.001 | -0.2493% | -0.8424% | 1.989 | -0.4006% | -1.486% | 1.006 | 0.48 | 0.43 | 0.24 | 0.18 | 0.88 | 0.2 | 0.19 | 2.001 | 9.2 | 9 | 9 | 8.3 | 119.6 | 120 | 0.32 | 0.17 | 0.11 | 0.02 | 0.5 | 0.12 | 0.01 | 1.989 | 11.5 | 11.5 | 11.6 | 11.7 | 119.4 | 119.7 | OJ, AR | CVD | GTR1_20230425_R15M10_01, GTR1_20230425_R31M10_01 |  | 99.7 | 100 | 100 |  | 2023-04-25 | 0.99751 | 0.99766 | 0.9977 | 0.99599 | 0.99753 | 0.99754 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230426 | 2.012 | 0.2991% | -0.2973% | 2.002 | 0.2504% | -0.842% | 1.005 | 0.43 | 0.4 | 0.18 | 0.17 | 0.91 | 0.22 | 0.26 | 2.012 | 9.3 | 9 | 8.6 | 8.6 | 119.6 | 120 | 0.28 | 0.18 | 0.02 | -0.01 | 0.54 | 0.1 | 0 | 2.002 | 11.3 | 11.4 | 11.3 | 11.2 | 119.4 | 119.8 | ON, SM | YM | GTR1_20230426_R15M10_01, GTR1_20230426_R31M10_01 |  | 92.3 | 100 | 95 | 100 | 2023-04-26 | 1.003 | 0.99765 | 0.99776 | 1.0025 | 0.99753 | 0.9976 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230427 | 2.001 | -0.2493% | -0.8424% | 1.992 | -0.2504% | -1.337% | 1.0045 | 0.49 | 0.44 | 0.24 | 0.24 | 0.84 | 0.22 | 0.24 | 2.001 | 9.1 | 9.2 | 8.6 | 8.8 | 119.6 | 120 | 0.35 | 0.19 | 0.13 | 0.04 | 0.61 | 0.15 | 0.03 | 1.992 | 11.5 | 11.6 | 11.5 | 11.3 | 119.4 | 119.7 | SM, JR | CVD | GTR1_20230427_R15M10_01, GTR1_20230427_R31M10_01 |  | 100 |  | 100 |  | 2023-04-27 | 0.99751 | 0.99768 | 0.99775 | 0.9975 | 0.99757 | 0.9976 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230428 | 2.003 | -0.1496% | -0.7433% | 1.994 | -0.1502% | -1.238% | 1.0045 | 0.51 | 0.52 | 0.22 | 0.25 | 0.93 | 0.21 | 0.27 | 2.003 | 9.3 | 8.9 | 8.8 | 8.6 | 119.6 | 120 | 0.29 | 0.12 | 0.06 | -0.01 | 0.58 | 0.1 | 0.02 | 1.994 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | OJ, SM | AC | GTR1_20230428_R15M10_01, GTR1_20230428_R31M10_01 |  | 91 | 100 | 100 |  | 2023-04-28 | 0.9985 | 0.9977 | 0.99775 | 0.9985 | 0.99759 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230502 | 1.999 | -0.349% | -0.9415% | 1.99 | -0.3505% | -1.436% | 1.0045 | 0.44 | 0.49 | 0.19 | 0.25 | 0.86 | 0.18 | 0.23 | 1.999 | 9.1 | 9.1 | 8.9 | 8.4 | 119.6 | 120 | 0.26 | 0.2 | 0.05 | 0.04 | 0.6 | 0.1 | 0.05 | 1.99 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | EN,AJ | CVD | GTR1_20230502_R15M10_01, GTR1_20230502_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-02 | 0.99651 | 0.99775 | 0.99776 | 0.99649 | 0.99764 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230503 | 1.998 | -0.3988% | -0.9911% | 1.987 | -0.5008% | -1.585% | 1.0055 | 0.49 | 0.64 | 0.31 | 0.28 | 1.01 | 0.21 | 0.31 | 1.998 | 9.2 | 9.1 | 8.2 | 9.1 | 119.6 | 119.9 | 0.28 | 0.14 | 0.09 | 0.01 | 0.53 | 0.11 | 0.02 | 1.987 | 11.5 | 11.5 | 11.7 | 11.3 | 119.4 | 119.8 | JR | CVD | GTR1_20230503_R15M10_01, GTR1_20230503_R31M10_01 |  | 93.6 | 100 | 100 | 100 | 2023-05-03 | 0.99601 | 0.99774 | 0.99775 | 0.99499 | 0.99764 | 0.99763 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230504 | 2.003 | -0.1496% | -0.7433% | 1.997 | 0% | -1.09% | 1.003 | 0.39 | 0.45 | 0.13 | 0.23 | 0.87 | 0.2 | 0.27 | 2.003 | 9 | 9.2 | 8.7 | 8.6 | 119.6 | 120 | 0.27 | 0.14 | 0.06 | 0.01 | 0.59 | 0.1 | 0.02 | 1.997 | 11.5 | 11.5 | 11.6 | 11.3 | 119.4 | 119.7 | GK,AR,EF | YM | GTR1_20230504_R15M10_01, GTR1_20230504_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-04 | 0.9985 | 0.99772 | 0.99774 | 1 | 0.99763 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230504 | 1.997 | -0.4487% | -1.041% | 1.998 | 0.05008% | -1.04% | 0.9995 | 0.52 | 0.41 | 0.21 | 0.23 | 0.89 | 0.16 | 0.27 | 1.997 | 9.6 | 8.4 | 8.6 | 8.8 | 119.6 | 119.9 | 0.37 | 0.13 | 0.11 | 0.01 | 0.55 | 0.11 | 0.01 | 1.998 | 11.8 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | JR | AC | GTR1_20230505_R15M10_01, GTR1_20230505_R31M10_01, GTR1_20230504_R15M10_01 |  | 100 |  | 100 |  | 2023-05-04 | 0.99551 | 0.99771 | 0.99775 | 1.0005 | 0.99763 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230504 | 1.997 | -0.1496% | -0.7433% | 1.998 | 0.05008% | -1.04% | 1.0025 | 0.52 | 0.41 | 0.21 | 0.23 | 0.89 | 0.16 | 0.27 | 1.997 | 9.6 | 8.4 | 8.6 | 8.8 | 119.6 | 119.9 | 0.37 | 0.13 | 0.11 | 0.01 | 0.55 | 0.11 | 0.01 | 1.998 | 11.8 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | JR | AC | GTR1_20230505_R15M10_01, GTR1_20230505_R31M10_01, GTR1_20230504_R15M10_01 |  | 100 |  | 100 |  | 2023-05-05 | 0.9985 | 0.99772 | 0.99773 | 1.0005 | 0.99763 | 0.99761 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230505 | 2.003 | -0.1496% | -0.7433% | 1.998 | 0.05008% | -1.04% | 1.0025 | 0.46 | 0.42 | 0.24 | 0.23 | 0.85 | 0.2 | 0.23 | 2.003 | 8.8 | 9.4 | 8.8 | 8.5 | 119.6 | 120 | 0.37 | 0.13 | 0.11 | 0.01 | 0.55 | 0.11 | 0.01 | 1.998 | 11.8 | 11.7 | 11.5 | 11.3 | 119.3 | 119.7 | OJ | AC | GTR1_20230505_R15M10_01, GTR1_20230505_R31M10_01 |  | 100 |  | 100 |  | 2023-05-05 | 0.9985 | 0.99769 | 0.99771 | 1.0005 | 0.99762 | 0.99758 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230508 | 2 | -0.2991% | -0.892% | 1.994 | -0.1502% | -1.238% | 1.003 | 0.49 | 0.4 | 0.26 | 0.22 | 0.88 | 0.21 | 0.22 | 2 | 9 | 9.3 | 8.9 | 8.4 | 119.6 | 120 | 0.26 | 0.13 | 0.09 | 0 | 0.54 | 0.1 | 0.02 | 1.994 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.7 | SM, AJ | AC | GTR1_20230508_R15M10_01, GTR1_20230508_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-08 | 0.99701 | 0.99773 | 0.99766 | 0.9985 | 0.99765 | 0.99753 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230509 | 1.998 | -0.3988% | -0.9911% | 1.988 | -0.4507% | -1.535% | 1.005 | 0.53 | 0.47 | 0.25 | 0.26 | 0.95 | 0.23 | 0.26 | 1.998 | 9.1 | 9.2 | 8.7 | 8.6 | 119.6 | 120 | 0.29 | 0.15 | 0.1 | -0.04 | 0.55 | 0.1 | -0.03 | 1.988 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | EN | GB | GTR1_20230509_R15M10_01, GTR1_20230509_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-09 | 0.99601 | 0.99773 | 0.99765 | 0.99549 | 0.99767 | 0.99753 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230510 | 2.004 | -0.0997% | -0.6938% | 1.992 | -0.2504% | -1.337% | 1.006 | 0.19 | 0.61 | 0.09 | 0.27 | 0.87 | 0.1 | 0.29 | 2.004 | 9.5 | 8.7 | 8.5 | 8.8 | 119.6 | 120 | 0.2 | 0.15 | 0.02 | 0.02 | 0.48 | 0.06 | 0.07 | 1.992 | 11.6 | 11.5 | 11.5 | 11.3 | 119.4 | 119.8 | AR | GB | GTR1_20230510_R15M10_01, GTR1_20230510_R31M10_01 |  | 90.1 | 100 | 98.4 | 100 | 2023-05-10 | 0.999 | 0.99765 | 0.99768 | 0.9975 | 0.9976 | 0.99757 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230511 | 2.007 | 0.04985% | -0.5451% | 1.993 | -0.2003% | -1.288% | 1.007 | 0.4 | 0.45 | 0.15 | 0.21 | 0.83 | 0.17 | 0.25 | 2.007 | 9.4 | 8.8 | 8.6 | 8.7 | 119.6 | 120 | 0.21 | 0.13 | 0.05 | -0.01 | 0.52 | 0.09 | 0.02 | 1.993 | 11.5 | 11.5 | 11.4 | 11.2 | 119.4 | 119.7 | JR | CVD | GTR1_20230511_R15M10_01, GTR1_20230511_R31M10_01 |  | 91.1 | 100 | 99.7 | 100 | 2023-05-11 | 1.0005 | 0.99765 | 0.9977 | 0.998 | 0.9976 | 0.99759 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230512 | 1.999 | -0.349% | -0.9415% | 1.99 | -0.3505% | -1.436% | 1.0045 | 0.47 | 0.4 | 0.21 | 0.24 | 0.85 | 0.2 | 0.23 | 1.999 | 9.2 | 9 | 8.7 | 8.6 | 119.6 | 120 | 0.32 | 0.17 | 0.12 | 0.02 | 0.58 | 0.11 | 0.03 | 1.99 | 11.4 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EN,GK,ON | EAL | GTR1_20230512_R15M10_01, GTR1_20230512_R31M10_01 |  | 99 | 100 | 100 | 100 | 2023-05-12 | 0.99651 | 0.99764 | 0.99775 | 0.99649 | 0.99759 | 0.99764 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230515 | 1.992 | -0.6979% | -1.288% | 1.985 | -0.6009% | -1.684% | 1.0035 | 0.41 | 0.51 | 0.19 | 0.28 | 0.92 | 0.21 | 0.27 | 1.992 | 9.1 | 9.1 | 8.7 | 8.8 | 119.6 | 119.9 | 0.34 | 0.21 | 0.1 | 0.06 | 0.57 | 0.11 | 0.04 | 1.985 | 11.5 | 11.5 | 11.4 | 11.2 | 119.4 | 119.8 | AR | GB | GTR1_20230515_R15M10_01, GTR1_20230515_R31M10_01 |  | 99.7 | 100 | 100 | 100 | 2023-05-15 | 0.99302 | 0.99765 | 0.99774 | 0.99399 | 0.9976 | 0.99764 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230516 | 2.001 | -0.2493% | -0.8424% | 1.991 | -0.3005% | -1.387% | 1.005 | 0.4 | 0.42 | 0.13 | 0.23 | 0.85 | 0.19 | 0.25 | 2.001 | 9.2 | 9 | 8.8 | 8.6 | 119.6 | 120 | 0.22 | 0.14 | 0.03 | 0.02 | 0.55 | 0.09 | 0.05 | 1.991 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.8 | GF, AE, EF | ON, YM | GTR1_20230516_R15M10_01, GTR1_20230516_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-16 | 0.99751 | 0.99768 | 0.99772 | 0.997 | 0.99765 | 0.99763 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230517 | 2.004 | -0.0997% | -0.6938% | 1.992 | -0.2504% | -1.337% | 1.006 | 0.42 | 0.47 | 0.2 | 0.21 | 0.88 | 0.19 | 0.24 | 2.004 | 9.1 | 9.2 | 8.8 | 8.7 | 119.6 | 120 | 0.26 | 0.15 | 0.09 | 0 | 0.52 | 0.1 | -0.01 | 1.992 | 11.4 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | SM, GK | YM ON | GTR1_20230517_R15M10_01, GTR1_20230517_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-17 | 0.999 | 0.99767 | 0.99771 | 0.9975 | 0.99761 | 0.99763 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230519 | 2.012 | 0.2991% | -0.2973% | 2.002 | 0.2504% | -0.842% | 1.005 | 0.32 | 0.58 | 0.18 | 0.28 | 0.94 | 0.15 | 0.3 | 2.012 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.29 | 0.17 | 0.07 | 0.04 | 0.51 | 0.08 | 0.03 | 2.002 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | EF, AJ, OJ | AD | GTR1_20230519_R15M10_01, GTR1_20230519_R31M10_01 |  | 86.3 | 100 | 89.4 | 100 | 2023-05-19 | 1.003 | 0.99771 | 0.99772 | 1.0025 | 0.99756 | 0.99763 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230522 | 2.002 | -0.1994% | -0.7929% | 1.994 | -0.1502% | -1.238% | 1.004 | 0.39 | 0.64 | 0.26 | 0.32 | 1.08 | 0.2 | 0.31 | 2.002 | 9.4 | 8.9 | 8.6 | 9 | 119.6 | 119.9 | 0.31 | 0.18 | 0.1 | 0.03 | 0.55 | 0.11 | 0.01 | 1.994 | 11.6 | 11.5 | 11.6 | 11.3 | 119.4 | 119.7 | AJ, OJ | GB | GTR1_20230522_R15M10_01, GTR1_20230522_R31M10_01 |  | 91.1 | 100 | 100 |  | 2023-05-22 | 0.99801 | 0.99769 | 0.99769 | 0.9985 | 0.99751 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230523 | 2 | -0.2991% | -0.892% | 1.991 | -0.3005% | -1.387% | 1.0045 | 0.42 | 0.43 | 0.17 | 0.19 | 0.88 | 0.2 | 0.23 | 2 | 9.4 | 8.7 | 8.7 | 8.7 | 119.6 | 119.9 | 0.28 | 0.14 | 0.1 | 0.01 | 0.53 | 0.11 | -0.01 | 1.991 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | AE, EF | CVD | GTR1_20230523_R15M10_01, GTR1_20230523_R31M10_01 |  | 90.8 | 100 | 100 | 100 | 2023-05-23 | 0.99701 | 0.99768 | 0.99773 | 0.997 | 0.99745 | 0.99765 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230524 | 1.997 | -0.2991% | -0.892% | 1.986 | -0.3005% | -1.387% | 1.0045 | 0.52 | 0.41 | 0.21 | 0.23 | 0.89 | 0.16 | 0.27 | 1.997 | 9.6 | 8.4 | 8.6 | 8.8 | 119.6 | 119.9 | 0.23 | 0.16 | 0.04 | 0.04 | 0.54 | 0.05 | 0.07 | 1.986 | 11.6 | 11.6 | 11.5 | 11.2 | 119.4 | 119.7 | JR | AD | GTR1_20230524_R15M10_01, GTR1_20230524_R31M10_01 |  | 95.8 | 100 | 96.7 | 100 | 2023-05-24 | 0.99701 | 0.99769 | 0.99773 | 0.99449 | 0.99743 | 0.99767 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230525 | 2.004 | -0.2991% | -0.892% | 4.854 | -0.3005% | -1.387% | 1.0045 | 0.5 | 0.5 | 0.21 | 0.28 | 0.78 | 0.18 | 0.26 | 2.004 | 9.1 | 9.1 | 8.7 | 8.8 | 119.6 | 119.9 | 1.21 | 0.47 | -1 | 0.47 | 1.68 | -0.88 | 0.14 | 4.854 | 13.6 | 13.5 | 15.8 | 13.5 | 105.4 | 100.8 | GK,JR | x | GTR1_20230525_R15M10_01, GTR1_20230525_R31M10_01 | Checked and ok with Matrixx3 /GB - Weird valus on R31, did the test 2 times. | 100 | 100 | 37.9
20230526 | 2.002 | -0.1994% | -0.7929% | 1.993 | -0.2003% | -1.288% | 1.0045 | 0.45 | 0.48 | 0.21 | 0.26 | 0.92 | 0.21 | 0.29 | 2.002 | 9.3 | 8.9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.31 | 0.15 | 0.11 | 0 | 0.61 | 0.13 | 0.02 | 1.993 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.7 | GK,SF | AC | GTR1_20230526_R15M10_01, GTR1_20230526_R31M10_01 |  | 91.1 | 100 | 100 | 100 | 2023-05-26 | 0.99801 | 0.9977 | 0.99765 | 0.998 | 0.99747 | 0.9976 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230529 | 1.995 | -0.5484% | -1.14% | 1.986 | -0.5508% | -1.634% | 1.0045 | 0.4 | 0.46 | 0.17 | 0.24 | 0.88 | 0.21 | 0.27 | 1.995 | 9.2 | 9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.29 | 0.16 | 0.09 | 0.03 | 0.57 | 0.11 | 0.03 | 1.986 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | GK,SF | CVD | GTR1_20230529_R15M10_01, GTR1_20230529_R31M10_01 |  | 99.7 | 100 | 100 | 100 | 2023-05-29 | 0.99452 | 0.99766 | 0.99765 | 0.99449 | 0.99748 | 0.9976 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230530 | 1.997 | -0.4487% | -1.041% | 1.988 | -0.4507% | -1.535% | 1.0045 | 0.5 | 0.51 | 0.17 | 0.27 | 1.01 | 0.2 | 0.27 | 1.997 | 9.3 | 8.9 | 8.7 | 8.7 | 119.6 | 120 | 0.28 | 0.09 | 0.05 | -0.02 | 0.52 | 0.08 | -0.01 | 1.988 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | SM, JR | YM | GTR1_20230530_R15M10_01, GTR1_20230530_R31M10_01 |  | 98.7 | 100 | 100 | 100 | 2023-05-30 | 0.99551 | 0.99776 | 0.99768 | 0.99549 | 0.99755 | 0.99765 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230531 | 2.003 | -0.1496% | -0.7433% | 1.992 | -0.2504% | -1.337% | 1.0055 | 0.42 | 0.36 | 0.17 | 0.21 | 0.87 | 0.19 | 0.24 | 2.003 | 9.4 | 8.8 | 8.7 | 8.9 | 119.6 | 119.9 | 0.27 | 0.16 | 0.05 | 0.01 | 0.52 | 0.09 | 0.02 | 1.992 | 11.5 | 11.4 | 11.4 | 11.3 | 119.4 | 119.7 | GF, SF | YM | GTR1_20230531_R15M10_01, GTR1_20230531_R31M10_01 |  | 91.1 | 100 | 100 | 100 | 2023-05-31 | 0.9985 | 0.99776 | 0.99767 | 0.9975 | 0.99756 | 0.99761 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230601 | 2.003 | -0.1496% | -0.7433% | 1.991 | -0.3005% | -1.387% | 1.006 | 0.47 | 0.52 | 0.19 | 0.29 | 0.93 | 0.21 | 0.29 | 2.003 | 9 | 9.2 | 8.8 | 8.8 | 119.6 | 119.9 | 0.29 | 0.12 | 0.08 | -0.02 | 0.56 | 0.1 | 0.02 | 1.991 | 11.4 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ON | AD | GTR1_20230601_R15M10_01, GTR1_20230601_R31M10_01 |  | 100 |  | 100 |  | 2023-06-01 | 0.9985 | 0.99774 | 0.99771 | 0.997 | 0.99756 | 0.99756 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230602 | 1.984 | -1.097% | -1.685% | 1.976 | -1.052% | -2.13% | 1.004 | 0.47 | 0.51 | 0.2 | 0.25 | 0.8 | 0.15 | 0.22 | 1.984 | 9.2 | 9 | 9 | 8.4 | 119.6 | 120 | 0.27 | 0.13 | 0.07 | 0 | 0.53 | 0.09 | 0.01 | 1.976 | 11.5 | 11.5 | 11.7 | 11.6 | 119.4 | 119.7 | EF | AC | GTR1_20230602_R15M10_01, GTR1_20230602_R31M10_01 |  | 90.1 | 100 | 99.7 | 100 | 2023-06-02 | 0.98903 | 0.99762 | 0.99769 | 0.98948 | 0.99745 | 0.99751 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230605 | 2.005 | -0.04985% | -0.6442% | 1.994 | -0.1502% | -1.238% | 1.0055 | 0.46 | 0.4 | 0.23 | 0.19 | 0.94 | 0.22 | 0.21 | 2.005 | 9.3 | 9 | 8.9 | 8.5 | 119.6 | 119.9 | 0.26 | 0.1 | 0.07 | -0.01 | 0.59 | 0.12 | 0.03 | 1.994 | 11.5 | 11.4 | 11.6 | 11.4 | 119.4 | 119.7 | AR | AC/KW | GTR1_20230605_R15M10_01, GTR1_20230605_R31M10_01 |  | 92.9 | 100 | 100 | 100 | 2023-06-05 | 0.9995 | 0.99762 | 0.99768 | 0.9985 | 0.99743 | 0.99745 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230607 | 2 | -0.2991% | -0.892% | 1.99 | -0.3505% | -1.436% | 1.005 | 0.21 | 0.66 | 0.09 | 0.29 | 0.9 | 0.12 | 0.34 | 2 | 9.3 | 8.8 | 8.3 | 9 | 119.6 | 119.9 | 0.3 | 0.19 | 0.07 | 0.05 | 0.55 | 0.07 | 0.06 | 1.99 | 11.4 | 11.4 | 11.5 | 11.3 | 119.4 | 119.7 | AJ, EF | YM | GTR1_20230607_R15M10_01, GTR1_20230607_R31M10_01 |  | 85 | 100 | 99.1 |  | 2023-06-07 | 0.99701 | 0.99763 | 0.99769 | 0.99649 | 0.99744 | 0.99743 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230608 | 2 | -0.2991% | -0.892% | 1.993 | -0.2003% | -1.288% | 1.0035 | 0.35 | 0.51 | 0.11 | 0.25 | 0.8 | 0.18 | 0.29 | 2 | 9.2 | 9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.25 | 0.2 | 0.07 | 0.04 | 0.56 | 0.08 | 0.03 | 1.993 | 11.4 | 11.4 | 11.4 | 11.2 | 119.3 | 119.7 | EF | GB | GTR1_20230608_R15M10_01, GTR1_20230608_R31M10_01 |  | 100 |  | 100 |  | 2023-06-08 | 0.99701 | 0.99764 | 0.99772 | 0.998 | 0.99751 | 0.99747 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230609 | 1.991 | -0.7478% | -1.338% | 1.983 | -0.7011% | -1.783% | 1.004 | 0.48 | 0.49 | 0.17 | 0.24 | 0.92 | 0.17 | 0.23 | 1.991 | 9.3 | 8.9 | 8.8 | 8.6 | 119.6 | 119.9 | 0.26 | 0.19 | 0.04 | 0.05 | 0.56 | 0.1 | 0.06 | 1.983 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | ON | AC | GTR1_20230609_R15M10_01, GTR1_20230609_R31M10_01 |  | 93.6 | 100 | 100 |  | 2023-06-09 | 0.99252 | 0.99764 | 0.9977 | 0.99299 | 0.99751 | 0.99747 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230612 | 1.994 | -0.5982% | -1.189% | 1.986 | -0.5508% | -1.634% | 1.004 | 0.44 | 0.5 | 0.23 | 0.24 | 0.86 | 0.22 | 0.27 | 1.994 | 9.2 | 9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.37 | 0.11 | 0.14 | -0.01 | 0.6 | 0.17 | 0.02 | 1.986 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | EF | AC | GTR1_20230612_R15M10_01, GTR1_20230612_R31M10_01 |  | 95.5 | 100 | 100 |  | 2023-06-12 | 0.99402 | 0.99764 | 0.99764 | 0.99449 | 0.9975 | 0.99746 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230613 | 2.001 | -0.2493% | -0.8424% | 1.992 | -0.2504% | -1.337% | 1.0045 | 0.54 | 0.43 | 0.18 | 0.22 | 0.92 | 0.21 | 0.24 | 2.001 | 9.1 | 9.2 | 8.8 | 8.6 | 119.6 | 119.9 | 0.24 | 0.14 | 0.05 | 0.02 | 0.56 | 0.09 | 0.02 | 1.992 | 11.4 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | EF | GB | GTR1_20230613_R15M10_01, GTR1_20230613_R31M10_01 |  | 100 |  | 100 |  | 2023-06-13 | 0.99751 | 0.99764 | 0.99766 | 0.9975 | 0.9975 | 0.99748 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230614 | 2.002 | -0.1994% | -0.7929% | 1.994 | -0.1502% | -1.238% | 1.004 | 0.41 | 0.72 | 0.25 | 0.34 | 1.07 | 0.2 | 0.35 | 2.002 | 9.3 | 9 | 8.6 | 8.9 | 119.6 | 119.9 | 0.24 | 0.13 | 0.06 | -0.01 | 0.54 | 0.1 | 0.03 | 1.994 | 11.5 | 11.5 | 11.4 | 11.2 | 119.4 | 119.7 | GK,ON | YM | GTR1_20230614_R15M10_01, GTR1_20230614_R31M10_01 |  | 95.2 | 100 | 100 | 100 | 2023-06-14 | 0.99801 | 0.99771 | 0.99776 | 0.9985 | 0.99757 | 0.99755 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230615 | 2.002 | -0.1994% | -0.7929% | 1.992 | -0.2504% | -1.337% | 1.005 | 0.49 | 0.5 | 0.2 | 0.26 | 0.92 | 0.18 | 0.3 | 2.002 | 9.3 | 8.8 | 8.9 | 8.7 | 119.6 | 119.9 | 0.23 | 0.14 | 0.04 | 0.02 | 0.59 | 0.1 | 0.08 | 1.992 | 11.5 | 11.4 | 11.6 | 11.5 | 119.4 | 119.7 | ON | YM | GTR1_20230615_R15M10_01, GTR1_20230615_R31M10_01 |  | 91.1 | 100 | 100 |  | 2023-06-15 | 0.99801 | 0.99776 | 0.99776 | 0.9975 | 0.99762 | 0.99756 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230616 | 1.994 | -0.5982% | -1.189% | 1.986 | -0.5508% | -1.634% | 1.004 | 0.47 | 0.49 | 0.21 | 0.23 | 0.87 | 0.22 | 0.28 | 1.994 | 9.2 | 9.1 | 8.9 | 8.6 | 119.6 | 119.9 | 0.36 | 0.11 | 0.14 | -0.03 | 0.57 | 0.14 | 0.01 | 1.986 | 11.5 | 11.6 | 11.5 | 11.2 | 119.3 | 119.7 | EN | AC | GTR1_20230616_R15M10_01, GTR1_20230616_R31M10_01 |  | 98.4 | 100 | 100 |  | 2023-06-16 | 0.99402 | 0.99774 | 0.99774 | 0.99449 | 0.99762 | 0.99756 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230619 | 2.033 | 1.346% | 0.7433% | 2.017 | 1.002% | -0.09906% | 1.0079 | 0.5 | 0.68 | 0.29 | 0.3 | 1.09 | 0.24 | 0.32 | 2.033 | 9.5 | 8.8 | 8.2 | 9.2 | 119.5 | 119.8 | 0.3 | 0.16 | 0.11 | 0 | 0.59 | 0.15 | 0.05 | 2.017 | 11.6 | 11.7 | 11.6 | 11.3 | 119.3 | 119.7 | GK | ML | GTR1_20230619_R15M10_01, GTR1_20230619_R31M10_01 | After LLRF calibration /ML | 19.2 | 95.8 | 37.2 | 98.4 | 2023-06-19 | 1.0135 | 0.99772 | 0.99762 | 1.01 | 0.99764 | 0.99745 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230620 | 1.998 | -0.3988% | -0.9911% | 1.986 | -0.5508% | -1.634% | 1.006 | 0.42 | 0.51 | 0.22 | 0.29 | 0.96 | 0.23 | 0.31 | 1.998 | 9.1 | 9.2 | 8.9 | 8.7 | 119.6 | 119.9 | 0.32 | 0.15 | 0.13 | 0.03 | 0.58 | 0.13 | 0.03 | 1.986 | 11.5 | 11.6 | 11.6 | 11.7 | 119.4 | 119.7 | EN | GB | GTR1_20230620_R15M10_01, GTR1_20230620_R31M10_01 |  | 100 |  | 100 |  | 2023-06-20 | 0.99601 | 0.99795 | 0.99762 | 0.99449 | 0.99785 | 0.99743 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230621 | 2.002 | -0.1994% | -0.7929% | 1.99 | -0.3505% | -1.436% | 1.006 | 0.43 | 0.45 | 0.19 | 0.22 | 0.95 | 0.22 | 0.28 | 2.002 | 9.3 | 8.9 | 8.8 | 8.7 | 119.6 | 119.9 | 0.35 | 0.1 | 0.16 | 0.01 | 0.65 | 0.16 | 0.07 | 1.99 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | ON | CVD | GTR1_20230621_R15M10_01, GTR1_20230621_R31M10_01 |  | 95.5 | 100 | 100 |  | 2023-06-21 | 0.99801 | 0.99791 | 0.99763 | 0.99649 | 0.99783 | 0.99744 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230622 | 2.001 | -0.2493% | -0.8424% | 1.992 | -0.2504% | -1.337% | 1.0045 | 0.46 | 0.49 | 0.16 | 0.22 | 0.92 | 0.19 | 0.2 | 2.001 | 9 | 9.2 | 8.8 | 8.6 | 119.6 | 119.9 | 0.33 | 0.16 | 0.09 | 0.04 | 0.53 | 0.11 | 0.03 | 1.992 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | OJ | ML | GTR1_20230622_R15M10_01, GTR1_20230622_R31M10_01 |  | 100 |  | 100 |  | 2023-06-22 | 0.99751 | 0.99794 | 0.99764 | 0.9975 | 0.99787 | 0.99751 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230626 | 2.008 | 0.0997% | -0.4955% | 2 | 0.1502% | -0.9411% | 1.004 | 0.29 | 0.61 | 0.13 | 0.31 | 0.79 | 0.15 | 0.3 | 2.008 | 9.3 | 8.8 | 8.7 | 8.8 | 119.6 | 119.9 | 0.28 | 0.12 | 0.11 | 0.02 | 0.53 | 0.12 | 0 | 2 | 11.5 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | OJ | ML, YM | GTR1_20230626_R15M10_01, GTR1_20230626_R31M10_01 |  | 88.2 | 100 | 75.3 | 100 | 2023-06-26 | 1.001 | 0.99796 | 0.99764 | 1.0015 | 0.99787 | 0.99751 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230627 | 1.999 | -0.349% | -0.9415% | 1.992 | -0.2504% | -1.337% | 1.0035 | 0.47 | 0.56 | 0.28 | 0.32 | 0.95 | 0.25 | 0.32 | 1.999 | 9.4 | 8.7 | 8.9 | 8.6 | 119.6 | 119.9 | 0.38 | 0.17 | 0.17 | 0.07 | 0.57 | 0.15 | 0.05 | 1.992 | 11.5 | 11.4 | 11.5 | 11.2 | 119.3 | 119.7 | GK | AC | GTR1_20230627_R15M10_01, GTR1_20230627_R31M10_01 |  | 91.1 | 100 | 100 | 100 | 2023-06-27 | 0.99651 | 0.99812 | 0.99764 | 0.9975 | 0.99801 | 0.9975 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230628 | 2.002 | -0.1994% | -0.7929% | 1.993 | -0.2003% | -1.288% | 1.0045 | 0.37 | 0.52 | 0.19 | 0.33 | 0.96 | 0.25 | 0.31 | 2.002 | 9.2 | 8.9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.36 | 0.2 | 0.16 | 0.1 | 0.6 | 0.15 | 0.03 | 1.993 | 11.4 | 11.3 | 11.3 | 11.2 | 119.3 | 119.7 | EN | AD | GTR1_20230628_R15M10_01, GTR1_20230628_R31M10_01 |  | 94.2 | 100 | 98.8 | 100 | 2023-06-28 | 0.99801 | 0.99825 | 0.99771 | 0.998 | 0.99812 | 0.99757 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230629 | 1.997 | -0.4487% | -1.041% | 1.99 | -0.3505% | -1.436% | 1.0035 | 0.45 | 0.58 | 0.26 | 0.32 | 0.98 | 0.26 | 0.3 | 1.997 | 9 | 9.1 | 8.8 | 8.6 | 119.6 | 119.9 | 0.42 | 0.16 | 0.19 | 0.06 | 0.67 | 0.18 | 0.03 | 1.99 | 11.3 | 11.5 | 11.3 | 11.1 | 119.4 | 119.7 | OJ | ML | GTR1_20230629_R15M10_01, GTR1_20230629_R31M10_01 |  | 100 |  | 100 |  | 2023-06-29 | 0.99551 | 0.99827 | 0.99776 | 0.99649 | 0.99814 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230630 | 2.004 | -0.0997% | -0.6938% | 1.995 | -0.1002% | -1.189% | 1.0045 | 0.4 | 0.72 | 0.22 | 0.32 | 1.06 | 0.2 | 0.34 | 2.004 | 9.3 | 8.9 | 8.5 | 8.9 | 119.6 | 119.9 | 0.38 | 0.29 | 0.12 | 0.12 | 0.64 | 0.11 | 0.07 | 1.995 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | AR | ML | GTR1_20230630_R15M10_01, GTR1_20230630_R31M10_01 |  | 94.2 | 100 | 100 |  | 2023-06-30 | 0.999 | 0.99828 | 0.99774 | 0.999 | 0.99813 | 0.99762 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230703 | 1.999 | -0.349% | -0.9415% | 1.989 | -0.4006% | -1.486% | 1.005 | 0.39 | 0.5 | 0.2 | 0.25 | 0.86 | 0.19 | 0.26 | 1.999 | 9 | 9.2 | 8.9 | 8.4 | 119.6 | 119.9 | 0.31 | 0.16 | 0.15 | 0.02 | 0.55 | 0.13 | 0.02 | 1.989 | 11.4 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | OJ | AC YM | GTR1_20230703_R15M10_01, GTR1_20230703_R31M10_01 |  | 100 |  | 100 |  | 2023-07-03 | 0.99651 | 0.99829 | 0.99772 | 0.99599 | 0.99815 | 0.99764 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230704 | 2 | -0.2991% | -0.892% | 1.991 | -0.3005% | -1.387% | 1.0045 | 0.52 | 0.46 | 0.2 | 0.23 | 0.98 | 0.21 | 0.27 | 2 | 9.2 | 9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.33 | 0.15 | 0.13 | 0.03 | 0.55 | 0.12 | 0.04 | 1.991 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | AR | AC | GTR1_20230704_R15M10_01, GTR1_20230704_R31M10_01 |  | 98.7 | 100 | 100 |  | 2023-07-04 | 0.99701 | 0.99844 | 0.99795 | 0.997 | 0.99827 | 0.99785 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230705 | 1.993 | -0.6481% | -1.239% | 1.985 | -0.6009% | -1.684% | 1.004 | 0.49 | 0.52 | 0.29 | 0.28 | 0.94 | 0.25 | 0.32 | 1.993 | 9 | 9.3 | 8.8 | 8.7 | 119.6 | 119.9 | 0.3 | 0.13 | 0.11 | 0.03 | 0.54 | 0.12 | 0.03 | 1.985 | 11.6 | 11.6 | 11.5 | 11.5 | 119.4 | 119.7 | ON | AD | GTR1_20230705_R15M10_01, GTR1_20230705_R31M10_01 |  | 100 |  | 100 |  | 2023-07-05 | 0.99352 | 0.9979 | 0.99791 | 0.99399 | 0.99785 | 0.99783 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230706 | 1.997 | -0.4487% | -1.041% | 1.99 | -0.3505% | -1.436% | 1.0035 | 0.45 | 0.48 | 0.27 | 0.26 | 0.96 | 0.25 | 0.29 | 1.997 | 9.1 | 9.1 | 8.8 | 8.7 | 119.6 | 119.9 | 0.36 | 0.16 | 0.17 | 0.04 | 0.59 | 0.14 | 0.04 | 1.99 | 11.5 | 11.5 | 11.5 | 11.3 | 119.4 | 119.7 | ON | AD | GTR1_20230706_R15M10_01, GTR1_20230706_R31M10_01 |  | 100 |  | 100 |  | 2023-07-06 | 0.99551 | 0.99797 | 0.99794 | 0.99649 | 0.99798 | 0.99787 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230707 | 1.987 | -0.9472% | -1.536% | 1.981 | -0.8012% | -1.882% | 1.003 | 0.41 | 0.53 | 0.27 | 0.27 | 1.07 | 0.26 | 0.32 | 1.987 | 9.3 | 8.9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.37 | 0.15 | 0.15 | 0.04 | 0.61 | 0.14 | 0.05 | 1.981 | 11.5 | 11.5 | 11.5 | 11.2 | 119.3 | 119.7 | AR | ML | GTR1_20230707_R15M10_01, GTR1_20230707_R31M10_01 |  | 97.1 | 100 | 100 |  | 2023-07-07 | 0.99053 | 0.99797 | 0.99796 | 0.99199 | 0.99804 | 0.99787 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230710 | 2 | -0.2991% | -0.892% | 1.99 | -0.3505% | -1.436% | 1.005 | 0.38 | 0.67 | 0.27 | 0.33 | 1.05 | 0.26 | 0.31 | 2 | 9.6 | 8.6 | 8.6 | 8.9 | 119.6 | 119.9 | 0.4 | 0.11 | 0.16 | -0.01 | 0.58 | 0.13 | -0.01 | 1.99 | 11.5 | 11.3 | 11.4 | 11.2 | 119.3 | 119.7 | AR | AC | GTR1_20230710_R15M10_01, GTR1_20230710_R31M10_01 |  | 90.4 | 100 | 97.8 | 100 | 2023-07-10 | 0.99701 | 0.99799 | 0.99812 | 0.99649 | 0.99806 | 0.99801 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230711 | 2.005 | -0.04985% | -0.6442% | 1.995 | -0.1002% | -1.189% | 1.005 | 0.41 | 0.53 | 0.17 | 0.28 | 0.97 | 0.18 | 0.28 | 2.005 | 9.3 | 8.9 | 8.7 | 8.8 | 119.6 | 119.9 | 0.28 | 0.16 | 0.09 | 0.02 | 0.63 | 0.1 | 0.05 | 1.995 | 11.5 | 11.5 | 11.5 | 11.3 | 119.3 | 119.7 | EN | YM | GTR1_20230711_R15M10_01, GTR1_20230711_R31M10_01 |  | 94.2 | 100 | 100 |  | 2023-07-11 | 0.9995 | 0.99786 | 0.99825 | 0.999 | 0.99791 | 0.99812 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230712 | 1.998 | -0.3988% | -0.9911% | 1.988 | -0.4507% | -1.535% | 1.005 | 0.25 | 0.67 | 0.12 | 0.32 | 1.02 | 0.14 | 0.35 | 1.998 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.27 | 0.17 | 0.11 | 0.03 | 0.57 | 0.1 | 0.05 | 1.988 | 11.5 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | OJ | YM | GTR1_20230712_R15M10_01, GTR1_20230712_R31M10_01 |  | 93.6 | 100 | 100 |  | 2023-07-12 | 0.99601 | 0.99792 | 0.99827 | 0.99549 | 0.99793 | 0.99814 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230713 | 2 | -0.2991% | -0.892% | 1.989 | -0.4006% | -1.486% | 1.0055 | 0.31 | 0.69 | 0.15 | 0.31 | 0.95 | 0.14 | 0.33 | 2 | 9.3 | 8.9 | 8.6 | 8.7 | 119.6 | 120 | 0.23 | 0.16 | 0.08 | 0.04 | 0.45 | 0.08 | 0.03 | 1.989 | 11.4 | 11.4 | 11.4 | 11.2 | 119.4 | 119.7 | OJ | AD | GTR1_20230713_R15M10_01, GTR1_20230713_R31M10_01 |  | 91.1 | 100 | 100 |  | 2023-07-13 | 0.99701 | 0.99792 | 0.99828 | 0.99599 | 0.99793 | 0.99813 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230714 | 2.011 | 0.2493% | -0.3469% | 2.004 | 0.3505% | -0.7429% | 1.0035 | 0.43 | 0.6 | 0.26 | 0.32 | 0.99 | 0.24 | 0.32 | 2.011 | 9.2 | 9 | 8.3 | 8.9 | 119.6 | 119.9 | 0.4 | 0.11 | 0.2 | -0.01 | 0.53 | 0.17 | -0.02 | 2.004 | 11.3 | 11.3 | 11.4 | 11.1 | 119.4 | 119.8 | OJ | YM | GTR1_20230714_R15M10_01, GTR1_20230714_R31M10_01 |  | 90.4 | 100 | 85 | 100 | 2023-07-14 | 1.0025 | 0.99803 | 0.99829 | 1.0035 | 0.998 | 0.99815 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230717 | 2.002 | -0.1994% | -0.7929% | 1.991 | -0.3005% | -1.387% | 1.0055 | 0.35 | 0.49 | 0.21 | 0.29 | 0.9 | 0.21 | 0.29 | 2.002 | 9.2 | 9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.32 | 0.13 | 0.13 | -0.04 | 0.57 | 0.14 | 0.01 | 1.991 | 11.4 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | EN | AD | GTR1_20230717_R15M10_01, GTR1_20230717_R31M10_01 |  | 100 |  | 100 |  | 2023-07-17 | 0.99801 | 0.99798 | 0.99844 | 0.997 | 0.99795 | 0.99827 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230718 | 2.013 | 0.349% | -0.2478% | 2.005 | 0.4006% | -0.6934% | 1.004 | 0.56 | 0.8 | 0.33 | 0.48 | 1.23 | 0.31 | 0.38 | 2.013 | 9.4 | 9.1 | 9 | 8.6 | 119.6 | 120 | 0.48 | 0.43 | 0.26 | 0.3 | 0.9 | 0.24 | 0.22 | 2.005 | 11.6 | 11.6 | 11.5 | 11.5 | 119.4 | 119.8 | AR | AC,AD | GTR1_20230718_R15M10_01, GTR1_20230718_R31M10_01 |  | 90.4 | 100 | 92.5 | 100 | 2023-07-18 | 1.0035 | 0.99806 | 0.9979 | 1.004 | 0.99805 | 0.99785 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230719 | 2.007 | 0.04985% | -0.5451% | 1.996 | -0.05008% | -1.139% | 1.0055 | 0.43 | 0.64 | 0.28 | 0.31 | 1.02 | 0.22 | 0.32 | 2.007 | 9.3 | 8.9 | 8.4 | 8.8 | 119.6 | 120 | 0.34 | 0.17 | 0.13 | 0.01 | 0.57 | 0.12 | 0.04 | 1.996 | 11.4 | 11.3 | 11.4 | 11.2 | 119.4 | 119.7 | AR | AC,YM | GTR1_20230719_R15M10_01, GTR1_20230719_R31M10_01 |  | 91.1 | 100 | 98.8 | 100 | 2023-07-19 | 1.0005 | 0.99812 | 0.99797 | 0.9995 | 0.99811 | 0.99798 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230720 | 2.002 | -0.1994% | -0.7929% | 1.993 | -0.2003% | -1.288% | 1.0045 | 0.33 | 0.71 | 0.18 | 0.33 | 1.03 | 0.18 | 0.35 | 2.002 | 9.2 | 9.1 | 8.4 | 9 | 119.6 | 119.9 | 0.27 | 0.11 | 0.12 | 0.01 | 0.58 | 0.13 | 0.04 | 1.993 | 11.5 | 11.5 | 11.5 | 11.2 | 119.4 | 119.7 | ON | YM | GTR1_20230720_R15M10_01, GTR1_20230720_R31M10_01 |  | 97.8 |  | 100 |  | 2023-07-20 | 0.99801 | 0.99839 | 0.99797 | 0.998 | 0.99835 | 0.99804 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230721 | 2.005 | -0.04985% | -0.6442% | 1.995 | -0.1002% | -1.189% | 1.005 | 0.43 | 0.62 | 0.16 | 0.34 | 0.99 | 0.17 | 0.3 | 2.005 | 9.2 | 9 | 8.7 | 8.7 | 119.6 | 119.9 | 0.26 | 0.21 | 0.11 | 0.09 | 0.61 | 0.11 | 0.07 | 1.995 | 11.4 | 11.4 | 11.5 | 11.2 | 119.4 | 119.7 | EN | YM, NIC | GTR1_20230721_R15M10_01, GTR1_20230721_R31M10_01 |  | 100 |  | 100 |  | 2023-07-21 | 0.9995 | 0.99857 | 0.99799 | 0.999 | 0.99847 | 0.99806 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230724 | 2.002 | -0.1994% | -0.7929% | 1.995 | -0.1002% | -1.189% | 1.0035 | 0.39 | 0.62 | 0.25 | 0.31 | 1.07 | 0.26 | 0.35 | 2.002 | 9 | 9.3 | 8.6 | 8.8 | 119.6 | 119.9 | 0.33 | 0.15 | 0.15 | -0.01 | 0.62 | 0.15 | 0.02 | 1.995 | 11.5 | 11.5 | 11.4 | 11.2 | 119.3 | 119.7 | ON | YM KW | GTR1_20230724_R15M10_01, GTR1_20230724_R31M10_01 |  | 100 |  | 100 |  | 2023-07-24 | 0.99801 | 0.9991 | 0.99786 | 0.999 | 0.9989 | 0.99791 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230725 | 2.004 | -0.0997% | -0.6938% | 1.997 | 0% | -1.09% | 1.0035 | 0.41 | 0.49 | 0.22 | 0.24 | 0.8 | 0.18 | 0.26 | 2.004 | 9.2 | 9 | 8.6 | 8.8 | 119.6 | 120 | 0.24 | 0.13 | 0.07 | 0.02 | 0.43 | 0.07 | 0.01 | 1.997 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | AR | YM KW | GTR1_20230725_R15M10_01, GTR1_20230725_R31M10_01 |  | 99.4 | 100 | 99.7 | 100 | 2023-07-25 | 0.999 | 0.99925 | 0.99792 | 1 | 0.99907 | 0.99793 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230726 | 2.006 | 0% | -0.5946% | 1.996 | -0.05008% | -1.139% | 1.005 | 0.32 | 0.73 | 0.19 | 0.34 | 1.09 | 0.2 | 0.37 | 2.006 | 8.9 | 9.3 | 8.6 | 8.9 | 119.6 | 119.9 | 0.28 | 0.19 | 0.08 | 0.03 | 0.49 | 0.1 | 0.03 | 1.996 | 11.4 | 11.5 | 11.4 | 11.2 | 119.4 | 119.8 | EN | YM | GTR1_20230726_R15M10_01, GTR1_20230726_R31M10_01 |  | 100 |  | 100 |  | 2023-07-26 | 1 | 0.99923 | 0.99792 | 0.9995 | 0.99908 | 0.99793 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230727 | 2.003 | -0.1496% | -0.7433% | 1.993 | -0.2003% | -1.288% | 1.005 | 0.41 | 0.54 | 0.21 | 0.3 | 0.96 | 0.21 | 0.31 | 2.003 | 9 | 9.2 | 8.8 | 8.6 | 119.6 | 119.9 | 0.28 | 0.23 | 0.13 | 0.09 | 0.63 | 0.13 | 0.06 | 1.993 | 11.5 | 11.5 | 11.6 | 11.5 | 119.4 | 119.7 | EN | YM KW | GTR1_20230727_R15M10_01, GTR1_20230727_R31M10_01 |  | 100 |  | 100 |  | 2023-07-27 | 0.9985 | 0.9995 | 0.99803 | 0.998 | 0.99937 | 0.998 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230728 | 2.005 | -0.04985% | -0.6442% | 1.995 | -0.1002% | -1.189% | 1.005 | 0.44 | 0.61 | 0.28 | 0.3 | 1.04 | 0.22 | 0.36 | 2.005 | 8.9 | 9.3 | 8.5 | 8.9 | 119.6 | 119.9 | 0.29 | 0.21 | 0.1 | 0.07 | 0.61 | 0.1 | 0.06 | 1.995 | 11.6 | 11.6 | 11.5 | 11.2 | 119.4 | 119.8 | EN | YM CMN | GTR1_20230728_R15M10_01, GTR1_20230728_R31M10_01 |  | 100 |  | 100 |  | 2023-07-28 | 0.9995 | 0.99973 | 0.99798 | 0.999 | 0.99968 | 0.99795 | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
Formulas:
Plot = DATE(LEFT([Plan ID],4), MID([Plan ID],5,2), RIGHT([Plan ID],2))

## Safety and lasers
|  | Ljusgardin: | Mån, ons och fredag |  |  |  | Lasers: | distance in mm between the middle of the laser line and the marker on the wall (0.5 mm precision)
|  | Dörrinterlock: | Tis och torsdag
|  | Avsyningsprocedur: | Dagligen
|  | Nödstopp: | Första fredagen i månaden
Datum |  | G1
|  | Safety |  |  |  |  | Lasers [mm] |  |  |  | Dosimetry Manager |  | All
|  | Avsyning | varningslampor | ljusgardin | dörrfunktion | RS | Red1 | Green1 | Red2 | Green2 | T [K] | p [hPa] | Kommentar
2023-02-24 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.1 | 993.9
2023-02-27 | Måndag | ok | ok | ok |  | ok | 0.5 | -0.5 | -0.5 | 0 | 296.3 | 1020.5
2023-02-28 | Tisdag |  | ok |  | ok | ok | -0.5 | 0 | 1 | -0.5 | 296.3 | 1021.8
2023-03-01 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.3 | 1022.1
2023-03-02 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 1 | -0.5 | 295.9 | 1017.5
2023-03-03 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.3 | 1009.1
2023-03-06 | Måndag | ok | ok | ok |  | ok | 0 | 0 | 0.5 | -1 | 296.4 | 993.2
2023-03-07 | Tisdag | ok | ok |  | ok | ok | 0 | 0 | 0.5 | -0.5 | 296.1 | 990.7
2023-03-08 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.3 | 979.7 | SM JR YM After beam alignment
2023-03-09 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 1 | -0.5 | 295.9 | 992.2
2023-03-10 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 295.9 | 994.3
2023-03-13 | Måndag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 295.9 | 994.4
2023-03-14 | Tisdag | ok | ok |  | ok | ok | 0 | -0.5 | 1 | -0.5 | 296.1 | 968.2
2023-03-15 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.4 | 989.5
2023-03-16 | Torsdag | ok | ok |  | ok | ok | -0.5 | -0.5 | 1 | -0.5 | 296.6 | 1010.1 | DM updated to T=296.4 K p=1011.14 hPa at 7:30 /ML
| Fredag | [11 empty] | Cyclo PM, no treatment
2023-03-20 | Måndag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.9 | 1002.6 | OK/AD
2023-03-21 | Tisdag | ok | ok |  | ok | ok | 0.5 | 0 | 1 | -0.5 | 296.1 | 1007.3
2023-03-22 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 295.9 | 996.3
2023-03-23 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 1 | -0.5 | 295.9 | 988.2
2023-03-24 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 1 | -0.5 | 296.1 | 988.7 | OK/KW
2023-03-27 | Måndag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.4 | 997.6 | ok/AC
2023-03-28 | Tisdag | ok | ok |  | ok | ok | -0.5 | 0 | 1 | -0.5 | 296.1 | 1004.9 | ok / YM
2023-03-29 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | 0 | 296.4 | 1012.2 | ok / YM ON
2023-03-30 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 0.5 | -0.5 | 296.1 | 1009.1 | ok /ON
2023-03-31 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.1 | 1001.5 | OK/GB
2023-04-03 | Måndag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 296.1 | 1032.2 | Ok/YM
2023-04-04 | Tisdag | ok | ok |  | ok | ok | 0 | 0 | 1 | -0.5 | 295.9 | 1024.7 | OK/GB
2023-04-05 | Onsdag | ok | ok | ok |  | ok | -0.5 | -0.5 | 1 | -0.5 | 295.9 | 1021.3 | OK/GB
2023-04-06 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1022.2 | ok/CVD
2023-04-07 | Fredag | ok | ok | ok |  | ok | -0.5 | -0.5 | 1 | -0.5 | 296.1 | 1023.1 | ok/cvd (se kommentar på setup check)
| Måndag | [11 empty] | Påskhelg, ingen behandling
2023-04-11 | Tisdag | ok | ok | ok |  | ok | -0.5 | 0 | 0.5 | -0.5 | 296.1 | 1010.4 | OK/GB
2023-04-12 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1005.6 | ok/YM
2023-04-13 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.8 | 1005.8 | ok/YM
2023-04-14 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1014.1 | ok/GB
2023-04-17 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.8 | 1025.9 | OK/GB
2023-04-18 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -1 | 295.8 | 1032.7 | ok/GB
2023-04-19 | Onsdag | ok | ok | ok |  | ok | -1 | 0 | 1 | 0 | 295.9 | 1033.3 | ok/CVD (I did a quick laser check)
2023-04-20 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1029.4 | ok/CVD
2023-04-21 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 0.5 | -0.5 | 295.8 | 1022.2 | ok/CVD
2023-04-24 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 999.9 | ok/AC,YM
2023-04-25 | Tisdag | ok | ok |  | ok | ok | -1 | 0 | 1 | -0.5 | 295.8 | 995.8 | ok/CVD
2023-04-26 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 995.5 | ok/YM
2023-04-27 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1000.8 | ok/CVD
2023-04-28 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1007.8 | ok/AC
| Måndag | [11 empty] | Holiday
2023-05-02 | Tisdag | ok | ok | ok |  | ok | -1 | 0 | 1 | -0.5 | 295.9 | 1002.4 | ok/CVD
2023-05-03 | Onsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.4 | 1015.2 | ok/CVD
2023-05-04 | Torsdag | ok | ok |  | ok | ok | -1 | 0 | 0.5 | -0.5 | 295.9 | 1023.4 | ok/YM
2023-05-05 | Fredag | ok | ok | ok |  | ok | -1 | 0 | 1 | -0.5 | 295.9 | 1028 | ok/AC
2023-05-08 | Måndag | ok | ok | ok |  | ok | -1 | 0 | 1 | -0.5 | 296.1 | 1025.4 | ok/AC
2023-05-09 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 297.4 | 1020 | ok/GB
2023-05-10 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.3 | 1015.8 | ok/GB
2023-05-11 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.3 | 1017.7 | ok/CVD
2023-05-12 | Fredag | ok | ok | ok |  | ok | -1 | 0 | 1 | 0 | 295.9 | 1021.4 | ok/EAL
2023-05-15 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.8 | 1013.5 | ok/GB
2023-05-16 | Tisdag | ok | ok |  | ok | ok | -1 | 0 | 1 | -0.5 | 296.3 | 1003.3 | ok/ON, YM
2023-05-17 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 0.5 | 0 | 296.4 | 997.7 | ok / YM ON
| Torsdag | [11 empty] | Holiday
2023-05-19 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1025.3 | AD
2023-05-22 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.9 | 1020.8 | GB
2023-05-23 | Tisdag | ok | ok |  | ok | ok | -1 | 0 | 1 | -0.5 | 295.75 | 1010.6 | CVD
2023-05-24 | Onsdag | ok | ok | ok |  | ok | -1.5 | -0.5 | 1 | -0.5 | 295.8 | 1012.6 | AD
2023-05-25 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 0.5 | 0 | 295.6 | 1009.7 | GB
2023-05-26 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.8 | 1010.6 | AC
2023-05-09 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1016.6 | CVD
2023-05-30 | Tisdag | ok | ok |  | ok | ok | -1 | 0 | 1 | -0.5 | 296.6 | 1018.6 | ok/YM
2023-05-31 | Onsdag | ok | ok | ok |  | ok | -1 | 0 | 1 | -0.5 | 295.8 | 1011.7 | ok/YM GB
2023-06-01 | Torsdag | ok | ok |  | ok | ok | -1 | 0 | 0.5 | -0.5 | 295.8 | 1012.2 | AD
2023-06-02 | Fredag | ok | ok | ok |  | ok | 1 | 0 | 0.5 | -0.5 | 295.9 | 1015.9 | AC
2023-06-05 | Måndag | ok | ok | ok |  | ok | 1 | -0.5 | 1 | -0.5 | 296.1 | 1016.7 | AC/KW
Tisdag
2023-06-07 | Onsdag | ok | ok | ok |  | ok | -1 | 0 | 0.5 | -0.5 | 295.8 | 1018.3 | ok/YM
2023-06-08 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | 0.5 | 295.9 | 1013.8 | ok/GB
2023-06-09 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1021.6 | ok/AC
2023-06-12 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.8 | 1023.5 | ok/AC
2023-06-13 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1018 | OK/GB
2023-06-14 | Onsdag | ok | ok | ok |  | ok | -0.5 | 0 | 1 | -0.5 | 295.8 | 1018 | ok / YM
2023-06-15 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 1 | -0.5 | 295.8 | 1017 | ok / YM
2023-06-16 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1013.6 | ok/AC
2023-06-19 | Måndag | ok | ok | ok |  | ok | -0.5 | 0 | 0.5 | 0 | 295.8 | 1012.4 | ok/ML
2023-06-20 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 0.5 | -0.5 | 295.9 | 1012.7 | OK/GB
2023-06-21 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 0.5 | -0.5 | 295.8 | 1011.9 | ok/CVD
2023-06-22 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1006.1 | ok/ML
| Fredag | [11 empty] | Holiday
2023-06-26 | Måndag | ok | ok | ok |  | ok | -0.5 | 0 | 0.5 | 0 | 295.9 | 1011.7
2023-06-27 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1003.9 | ok/AC
2023-06-28 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1007.5 | AD
2023-06-29 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1004.9 | ML
2023-06-30 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 1001.1 | ML
2023-07-03 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 984.8 | ok/AC
2023-07-04 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.3 | 991.6 | ok/AC
2023-07-05 | Onsdag | ok | ok | ok |  | ok | -1 | -1 | 1 | 0 | 296.3 | 1000.6 | AD
2023-07-06 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1005.6 | AD
2023-07-07 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | 0.5 | 295.6 | 1007.9 | ML
2023-07-10 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | 0.5 | 295.9 | 1017.6 | AC
2023-07-11 | Tisdag | ok | ok |  | ok | ok | -0.5 | -0.5 | 1 | -0.5 | 296.1 | 1013 | ok / YM
2023-07-12 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1008.1 | ok / YM
2023-07-13 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.3 | 1000.7 | AD
2023-07-14 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | 0 | 296.1 | 1003.8 | ok/YM
2023-07-17 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296 | 1003.4 | AD/YM
2023-07-18 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.4 | 999.5 | AC,AD
2023-07-19 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 1000 | AC,YM
2023-07-20 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 997 | ok/YM
2023-07-21 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.25 | 997.4 | YM, NIC
2023-07-24 | Måndag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 998.4 | ok/YM KW
2023-07-25 | Tisdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 296.4 | 990.7 | ok/YM KW
2023-07-26 | Onsdag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.1 | 994 | ok/YM
2023-07-27 | Torsdag | ok | ok |  | ok | ok | -1 | -0.5 | 1 | -0.5 | 295.9 | 998.8 | ok / YM KW
2023-07-28 | Fredag | ok | ok | ok |  | ok | -1 | -0.5 | 1 | -0.5 | 296.2 | 997.4 | ok/ YM CMN
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag

## Signatures
Signature | First name | Last name | Profession | Note
AA | Ali | Alkhiat | Medical Physicist
AC | Athanasia | Christou | Summer physicist
AD | Alexandru | Dasu | Medical physicist
AE | Albin | Eriksson | Uppstuk physicist
AHL | Annika | Hall | Medical physicist
AJ | Anton | Jansson | Uppstuk physicist
AR | Adam | Röjvall | Uppstuk physicist
CVD | Christina | Vallhagen Dahlgren | Medical Physicist
DJ | Dan | Josefsson | Medical physicist
EA | Erik | Almhagen | Medical physicist
EF | Emil | Forsberg | Uppstuk physicist
EL | Elias | Lindbäck | Medical physicist
EN | Erik | Näslund | Uppstuk physicist
EPe | Erik | Pettersson | Medical physicist
GB | Gloria | Bäckström | Medical physicist
GF | Gustav | Furubjelke | Uppstuk physicist
JM | Joakim | Medin | Medical physicist
JR | Johan | Rensfeldt | Uppstuk physicist
GK | Gurpreet | Kaur | Uppstuk physicist
KW | Kenneth | Wikström | Medical physicist
LL | Linnea | Lund | Medical physicist
MB | Mikael | Blomqvist | Medical physicist
ME | Marika | Enmark | Medical physicist
MF | Marcus | Fager | Medical physicist
MJ | Mikael | Johansson | Uppstuk physicist
MK | Marcus | Krantz | Medical physicist
Ml | Malgorzata | Liszka | Medical physicist
NIC | Carl Magnus | Nilsson | Medical physicist
OJ | Olof | Jardfelt | Uppstuk physicist
ON | Olivier | Näslund | Uppstuk physicist
PL | Peter | Larsson | Medical physicist
SF | Stina | Fredriksson | Uppstuk physicist
SM | Sebastian | Mikkelsen | Uppstuk physicist
TH | Thomas | Henry | Medical physicist
UG | Ulf | Granlund | Medical physicist
YM | Younes | Majeddam | Medical physicist

## Updates
No | Date | Responsible | Change Description
1 | 2023-02-23 | ML | The document released for users
2
3
4
5
6