## GTR2_Setup check
ID.nr:
Titel: | [30 empty] | Dokumenttyp:
Protocol for Daily QA: Setup preparation fast check | [30 empty] | Strålskyddsinstruktion
Godkänt av: | [29 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [29 empty] | Skapat den:
Reviderat av: | [29 empty] | Reviderat den:
GTR2
Setup check: GTR2 MatriXX PT_1
Set-up: |  | 2.3 cm SW + MatrixPT position adjusted using X-rays (chambers at iso)
Detector |  | IBA MatrixPT_1
Phantom |  | SW
MyQA Project: |  | GTR2_DailyQA_Matrixx
Patient in AdaptDelivery |  | G2_Daily_MatriXX Skandion
Plan ID: |  | Single layer_E150MeV_F10x10
Detector Calib. Factor: |  | 87.422 | GTR2_Reference Calib. 2022-11-04
Reference dose (TPS) [Gy] |  | 1.699
Calibration dose [Gy] |  | 1.716
Good: within 0.5%
1.682 | 1.716
1.6988 | 1.7332
Warning: within 2%
1.717 | 1.681
1.7342 | 1.6978
Warning: over 2%
1.734 | 1.664
1.7513 | 1.6807
|  | Conditions in room |  |  | Detector condition | [6 empty] | IBA laser off CV (mm) | [6 empty] | Table position saved in Physic QA
Date | Responsible | p [hPa] | Troom [0C] | Troom [K] | Calibration Factor | TMatrixx [0C] | pMatriXX [hPa] | MyQA TMatrixx [0C] | MyQA pMatriXX [hPa] | MyQA Offset TMatrixx [0C] | MyQA Offset pMatriXX [hPa] | Δx (cm) | Δy (cm) | Δz (cm) | Rot (o) | Pitcht (o) | Roll (o) | CV (cm) | X (cm) | Y (cm) | Z (cm) | Rot (o) | Pitcht (o) | Roll (o) | CV (mm) | Z (cm) after applying Offset -3 cm | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from Calibration dose (%) | GI (1%, 1 mm) | GI (2%, 2 mm) | Notes
2022-11-18 | AC | 1017 | 23.7 | 296.85 | 87.422 | 22.9 | 1017 | 23.5 | 1018 | -0.6 | -1 | [15 empty] | 1.704 | 0.2943% | -0.6993% | 100 |  | run 04
2022-11-18 | AC | 1017 | 23.7 | 296.85 | 87.422 | 22.9 | 1017 | 23.5 | 1018 | -0.6 | -1 | [15 empty] | 1.706 | 0.412% | -0.5828% |  |  | run 05, reference image
2022-11-18 | AC | 1017 | 23.7 | 296.85 | 87.422 | 22.9 | 1017 | 23.5 | 1018 | -0.6 | -1 | [15 empty] | 1.706 | 0.412% | -0.5828% | 100 |  | run 06
2023-02-16 | x | 1009 | 23.6 | 296.75 | 87.422 | 22.6 | 1009 | 23.7 | 1009 | -1.1 | 0 | [14 empty] | -2.95 | 1.6912 | -0.4603% | -1.446% | 100
2023-02-17 | x | 1000.9 | 23.7 | 296.85 | 87.422 | 22.1 | 1000.9 | 24.1 | 1001 | -2 | -0.1 | 0.07 | 0.07 | 0 | 0 | 0 | 0 | 0.098995 | -0.03 | 0.05 | 0.06 | 0 | -0.1 | 0 | 0.083666 | -2.94 | 1.6923 | -0.3955% | -1.382% | 83.2 | 99.6 | Used wrong project, reference is from old matrix procedure project
2023-02-20 | AJ | 992.6 | 23.1 | 296.25 | 87.422 | 21.8 | 992.6 | 23.6 | 991 | -1.8 | 1.6 | 0.07 | 0.15 | 0.08 | -0.2 | 0 | 0 | 0.18385 | 0.07 | 0.15 | 0.08 | 359.8 | 0 | 0 | 0.18385 | -2.92 | 1.6909 | -0.4768% | -1.463% | 83.5 | 100
2023-02-21 | AE | 1009.2 | 23.7 | 296.85 | 87.422 | 22.1 | 1009.2 | 23 | 1010 | -0.9 | -0.8 | 0.05 | 0.05 | 0.03 | 0 | 0 | 0 | 0.076811 | 0.05 | 0.06 | 0.08 | 0 | 0 | 0 | 0.1118 | -2.92 | 1.6854 | -0.8005% | -1.783% | 58.5 | 100 | run_01/Training session
2023-02-21 | AR | 1009.2 | 23.7 | 296.85 | 87.422 | 22.4 | 1009.2 | 23.6 | 1012 | -1.2 | -2.8 | 0.05 | 0.05 | -0.03 | 0 | 0 | 0 | 0.076811 | 0.05 | 0.06 | 0.08 | 0 | 0 | 0 | 0.1118 | -2.92 | 1.68 | -1.118% | -2.098% | 54.6 | 100 | run_02/Training session
2023-02-24 | OJ | 993.8 | 23.4 | 296.55 | 87.422 | 22.2 | 993.8 | 23.5 | 994 | -1.3 | -0.2 | 0.05 | 0.08 | -0.03 | 0 | 0 | 0 | 0.098995 | 0.05 | 0.08 | 0.05 | 0 | 0 | 0 | 0.10677 | -2.95 | 1.685 | -0.824% | -1.807% | 57.2 | 100
2023-02-27 | GK,EH,AR,AJ,OJ | 1019.9 | 23.5 | 296.65 | 87.422 | 22.7 | 1019.9 | 23.7 | 1021 | -1 | -1.1 | 0.03 | 0.07 | -0.02 | 0 | 0 | 0 | 0.07874 | 0.03 | 0.07 | 0.08 | 0 | 0 | 0 | 0.11045 | -2.92 | 1.676 | -1.354% | -2.331% | 54 | 93.4
2023-02-28 | GF | 1021.8 | 23.6 | 296.75 | 87.422 | 22.8 | 1021.8 | 23.5 | 1023 | -0.7 | -1.2 | 0.04 | 0.07 | 0 | -0.1 | 0 | 0 | 0.080623 | 0.04 | 0.07 | 0.1 | 359.9 | 0 | 0 | 0.12845 | -2.9 | 1.683 | -0.9417% | -1.923% | 55.8 | 100
2023-03-01 | SM, AE | 1022 | 23.3 | 296.45 | 87.422 | 22.4 | 1022 | 23.4 | 1023 | -1 | -1 | 0.05 | 0.05 | 0 | 0 | 0 | 0 | 0.070711 | 0.05 | 0.05 | 0.04 | 0 | 0 | 0 | 0.08124 | -2.96 | 1.684 | -0.8829% | -1.865% | 54.8 | 100
2023-03-02 | JR,AR | 1017.7 | 23.4 | 296.55 | 87.422 | 22.5 | 1017.7 | 23.1 | 1018 | -0.6 | -0.3 | 0.02 | 0.05 | 0 | -0.1 | 0 | 0 | 0.053852 | 0.02 | 0.05 | 0.09 | -0.01 | 0 | 0 | 0.10488 | -2.91 | 1.6887 | -0.6051% | -1.59% | 70 | 100
2023-03-03 | SM, AR | 1009.2 | 23.6 | 296.75 | 87.422 | 22.6 | 1009.2 | 23.3 | 1010 | -0.7 | -0.8 | 0.06 | 0.06 | -0.02 | -0.1 | 0 | 0 | 0.087178 | 0.06 | 0.06 | 0.08 | 359.9 | 0 | 0 | 0.11662 | -2.92 | 1.688 | -0.6474% | -1.632% | 85.1 | 100
2023-03-06 | AR, GK | 993.1 | 23.3 | 296.45 | 87.422 | 22.7 | 993.1 | 23.7 | 994 | -1 | -0.9 | 0.05 | 0.08 | -0.02 | 0 | 0 | 0 | 0.096437 | 0.05 | 0.07 | 0.06 | 0 | 0 | 0 | 0.10488 | -2.94 | 1.7009 | 0.1107% | -0.8811% | 100 | 100 | 1,685 vid upprepning efter jk-byte
2023-03-07 | AE, GK | 990.7 | 23.4 | 296.55 | 87.422 | 22.7 | 990.7 | 23.5 | 992 | -0.8 | -1.3 | 0.07 | 0.11 | 0.08 | 0 | 0 | 0 | 0.15297 | 0.07 | 0.11 | 0.08 | 0 | 0 | 0 | 0.15297 | -2.92 | 1.6923 | -0.3972% | -1.384% | 97.8 | 100
2023-03-08 | JR,AR | 979.6 | 23.5 | 296.65 | 87.422 | 22.9 | 990.7 | 23.5 | 981 | -0.6 | 9.7 | 0.06 | 0.06 | 0 | 0 | 0 | 0 | 0.084853 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -2.91 | 1.6842 | -0.8705% | -1.853% | 64.4 | 100 | Odd to have only 0 on physics QA
2023-03-09 | EH | 992.3 | 23.7 | 296.85 | 87.422 | 22.7 | 992.3 | 23.2 | 993 | -0.5 | -0.7 | 0.05 | 0.05 | 0 | 0 | 0 | 0 | 0.070711 | 0.05 | 0.06 | 0.09 | 0 | 0 | 0 | 0.11916 | -2.91 | 1.6981 | -0.05062% | -1.041% | 100 | 100 | YM
2023-03-10 | GK,AR | 994.5 | 22.6 | 295.75 | 87.422 | 22.8 | 994.5 | 24.2 | 995 | -1.4 | -0.5 | 0.04 | 0.05 | 0 | -0.1 | 0 | 0 | 0.064031 | 0.04 | 0.05 | 0.09 | 359.9 | 0 | 0 | 0.11045 | -2.9 | 1.7027 | 0.219% | -0.7739% | 100 | 100

## GTR2_150MeV
ID.nr:
Titel: | [24 empty] | Dokumenttyp:
Protocol for Daily QA: Consistency check of single layer and cubic plans dosimetric parameters | [24 empty] | Strålskyddsinstruktion
Godkänt av: | [23 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [23 empty] | Skapat den:
Reviderat av: | [23 empty] | Reviderat den:
GTR2
Daily single layer parameters
MatriXX PT_1
Parameters and its criterias
Parameter | Value (Gy) | Good: within 1% |  | Warning: within 2% |  | Action level: over 2%
Ref. dose (SPS) | 1.699 | 1.716 | 1.682 | 1.717 | 1.681 | 1.734 | 1.664
Calibration dose | 1.716 | 1.7332 | 1.6988 | 1.7342 | 1.6978 | 1.7513 | 1.6807
Parameter | Value (mm) |  |  | +/- 1.001 mm |  | +/- 2.001 mm
Penumbra (average) | 7.7975 | 8.7975 | 6.7975 | 8.7985 | 6.7965 | 9.7985 | 5.7965
FWHM (average) | 105.05 | 106.05 | 104.05 | 106.051 | 104.049 | 107.051 | 103.049
Parameter | Value (Gy) | Good: within +/-0.5% |  | Warning: within +/-1% |  | Action level: over 1%
Symmetry | x | 0.5 | -0.5 | 0.5 | -0.5 | 1 | 1
Parameter | Value (Gy) | Good: within +/-1% |  | Warning: within +/-2% |  | Action level: over 3%
Flatness | x | 1 | -1 | 2 | -2 | 3 | -3
| Output Calibration (150 MeV 10x10 cm2, 1MU per spot) | [19 empty] | Plot
Date | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from Calibration dose (%) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Sign | Checked | Notes | Date | Difference from reference (a.u.) | Medelvärde +/- 10 dygn | Medelvärde +/- 20 dygn | Ref | 1% | -1% | 2% | -2%
20221118 | 1.704 | 0.2943% | -0.6993% | 0.72 | 0.86 | -0.34 | 0.49 | 1.33 | -0.35 | 0.43 | 1.704 | 8.2 | 8.3 | 7.2 | 7.5 | 104.8 | 105.2 | ML | x | GTR2_20221118_OF_04 | 2022-11-18 | 1.0029 | 0.9947 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221118 | 1.706 | 0.412% | -0.5828% | 0.87 | 0.83 | -0.43 | 0.46 | 1.27 | -0.37 | 0.38 | 1.706 | 8.1 | 8.4 | 7.5 | 7.3 | 104.8 | 105.2 | ML | x | GTR2_20221118_OF_05, REFERENCE IMAGE | 2022-11-18 | 1.0041 | 0.99461 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221118 | 1.706 | 0.412% | -0.5828% | 0.79 | 0.82 | -0.42 | 0.48 | 1.38 | -0.39 | 0.4 | 1.706 | 8 | 8.5 | 7.2 | 7.6 | 104.8 | 105.2 | ML | x | GTR2_20221118_OF_06 | 2022-11-18 | 1.0041 | 0.99508 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230216 | 1.691 | -0.4709% | -1.457% | 0.8 | 0.78 | -0.4 | 0.41 | 1.27 | -0.32 | 0.38 | 1.691 | 8.4 | 8.1 | 7.2 | 7.6 | 104.9 | 105.3 | JR | ok/ML | GTR2_20230216_OF_01 | 2023-02-16 | 0.99529 | 0.99517 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230220 | 1.691 | -0.4709% | -1.457% | 0.66 | 0.84 | -0.34 | 0.47 | 1.17 | -0.26 | 0.41 | 1.691 | 7.5 | 8.8 | 7.6 | 7.2 | 104.8 | 105.2 | AJ | ok/ML | GTR2_20230220_OF_01 | 2023-02-20 | 0.99529 | 0.99492 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230221 | 1.685 | -0.824% | -1.807% | 0.76 | 0.9 | -0.41 | 0.56 | 1.28 | -0.33 | 0.44 | 1.685 | 7.8 | 8.6 | 7.6 | 7.4 | 104.9 | 105.2 | AE | ok/ML | GTR2_20230221_OF_01 | 2023-02-21 | 0.99176 | 0.99491 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230221 | 1.68 | -1.118% | -2.098% | 0.75 | 0.74 | -0.37 | 0.49 | 1.21 | -0.29 | 0.43 | 1.68 | 7.3 | 8.9 | 7.8 | 7 | 104.9 | 105.3 | AR | ok/ML | GTR2_20230221_OF_02 | 2023-02-21 | 0.98882 | 0.99516 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230224 | 1.685 | -0.824% | -1.807% | 0.72 | 0.79 | -0.35 | 0.5 | 1.2 | -0.29 | 0.42 | 1.685 | 8 | 8.4 | 7.5 | 7.2 | 104.9 | 105.3 | OJ | AD | GTR2_20230224_OF_01 | 2023-02-24 | 0.99176 | 0.99554 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230227 | 1.676 | -1.354% | -2.331% | 0.59 | 0.83 | -0.29 | 0.39 | 1.03 | -0.21 | 0.37 | 1.676 | 7.5 | 8.8 | 7.6 | 7.2 | 104.9 | 105.2 | GK,AJ,AR,OJ,EH | AH,AC | GTR2_20230227_OF_01 | 2023-02-27 | 0.98646 | 0.99554 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230228 | 1.683 | -0.9417% | -1.923% | 0.84 | 0.86 | -0.42 | 0.48 | 1.19 | -0.33 | 0.4 | 1.683 | 7.9 | 8.5 | 7.6 | 7.1 | 104.9 | 105.3 | GF | YM | GTR2_20230228_OF_01 | 2023-02-28 | 0.99058 | 0.99554 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230301 | 1.684 | -0.8829% | -1.865% | 0.76 | 0.82 | -0.34 | 0.43 | 1.17 | -0.3 | 0.36 | 1.684 | 8 | 8.4 | 7.4 | 7.3 | 104.9 | 105.2 | SM, AE | GB | GTR2_20230301_OF_01 | 2023-03-01 | 0.99117 | 0.99513 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230302 | 1.689 | -0.5886% | -1.573% | 0.74 | 0.89 | -0.41 | 0.47 | 1.19 | -0.33 | 0.4 | 1.689 | 7.8 | 8.6 | 7.5 | 7.3 | 104.9 | 105.2 | JR | YM | GTR2_20230302_OF_01 | 2023-03-02 | 0.99411 | 0.9946 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230303 | 1.688 | -0.6474% | -1.632% | 0.68 | 0.74 | -0.36 | 0.44 | 1.13 | -0.29 | 0.37 | 1.688 | 8.1 | 8.4 | 7 | 7.6 | 104.9 | 105.3 | SM, AR | AC | GTR2_20230303_OF_01 | 2023-03-03 | 0.99353 | 0.994 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230306 | 1.701 | 0.1177% | -0.8741% | 0.76 | 0.84 | -0.35 | 0.5 | 1.18 | -0.32 | 0.39 | 1.701 | 8 | 8.3 | 7.4 | 7.1 | 104.9 | 105.3 | CVD | CVD | GTR2_20230306_OF_01 | 2023-03-06 | 1.0012 | 0.99392 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230307 | 1.693 | -0.3531% | -1.34% | 0.73 | 0.81 | -0.37 | 0.4 | 1.21 | -0.27 | 0.37 | 1.693 | 8 | 8.4 | 7.3 | 7.4 | 104.9 | 105.2 | AE | YM | GTR2_20230307_OF_01 | 2023-03-07 | 0.99647 | 0.99382 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230308 | 1.684 | -0.8829% | -1.865% | 0.68 | 0.87 | -0.33 | 0.5 | 1.09 | -0.25 | 0.41 | 1.684 | 7.8 | 8.5 | 7.4 | 7.3 | 104.9 | 105.2 | JR | YM | GTR2_20230308_OF_01 | 2023-03-08 | 0.99117 | 0.99398 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230308 | 1.69 | -0.5297% | -1.515% | 0.56 | 0.77 | -0.29 | 0.47 | 1.09 | -0.27 | 0.36 | 1.69 | 7.5 | 8.8 | 7.2 | 7.5 | 104.9 | 105.3 | AC | x | GTR2_20230308_OF_02 | 2023-03-08 | 0.9947 | 0.99441 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230309 | 1.698 | -0.05886% | -1.049% | 0.64 | 0.82 | -0.31 | 0.47 | 1.15 | -0.27 | 0.39 | 1.698 | 7.7 | 8.7 | 7.4 | 7.4 | 104.9 | 105.2 | EH | YM | GTR2_20230309_OF_01 | 2023-03-09 | 0.99941 | 0.99465 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230310 | 1.703 | 0.2354% | -0.7576% | 0.76 | 0.81 | -0.39 | 0.42 | 1.13 | -0.24 | 0.38 | 1.703 | 8.2 | 8.2 | 7.2 | 7.5 | 104.9 | 105.2 | AR, GK | KW | GTR2_20230310_OF_01 | 2023-03-10 | 1.0024 | 0.99547 | 0.99554 | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
Formulas:
Date = DATE(LEFT([Parameter],4), MID([Parameter],5,2), RIGHT([Parameter],2))

## GTR2_Cubes
ID.nr:
Titel: | [52 empty] | Dokumenttyp:
Protocol for Daily QA: Consistency check of single layer and cubic plans dosimetric parameters | [52 empty] | Strålskyddsinstruktion
Godkänt av: | [51 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [51 empty] | Skapat den:
Reviderat av: | [51 empty] | Reviderat den:
GTR2
Daily cubes parameters
MatriXX PT_1
Set-up: |  | MatriXX (MatrixPT_1 - position adjusted using X-rays (chambers at iso))
Detector: |  | IBA MatrixPT_1
Phantom: |  | SW
MyQA Project: |  | GTR2_DailyQA_Matrixx
Patient in AdaptDelivery: |  | G2_Daily_MatriXX Skandion
Plan ID: |  | Cube_R15_SOBP10; Cube_R31_SOBP10
Parameters and its criterias
Plan ID | Parameter | Value (Gy) | Good: within 2% |  | Warning: over 2% |  | Action level: over 3% |  |  | Roos IC vs. TPS |  | Roos IC | MatriXX_1 (Mean)
R15M10 | Ref. dose (Roos IC SW)* | 1.987 | 2.0267 | 1.9473 | 2.0277 | 1.9463 | 2.0476 | 1.9264 |  |  | -0.8483% | 1.987 | 1.977 | -0.5033%
R31M10 | Ref. dose (Roos IC SW)* | 2.007 | 2.0471 | 1.9669 | 2.0481 | 1.9659 | 2.0682 | 1.9458 |  |  | -0.9378% | 2.007 | 1.999 | -0.3986%
R15M10 | Dose from TPS | 2.004 | 2.0441 | 1.9639 | 2.0451 | 1.9629 | 2.0651 | 1.9429
R31M10 | Dose from TPS | 2.026 | 2.0441 | 1.9639 | 2.0451 | 1.9629 | 2.0651 | 1.9429
Plan ID | Parameter | Value (mm) | Good |  | +/- 1.001 mm |  | +/- 2.001 mm
R15M10 | Penumbra (average) | 8.4722 | 9.4722 | 7.4722 | 9.4732 | 7.4712 | 10.4732 | 6.4712
R31M10 | Penumbra (average) | 11.4306 | 12.4306 | 10.4306 | 12.4316 | 10.4296 | 13.4316 | 9.4296
Plan ID | Parameter | Value (mm) | Good |  | +/- 1.001 mm |  | +/- 2.001 mm
R15M10 | FWHM (average) | 120.0167 | 121.0167 | 119.0167 | 121.0177 | 119.0157 | 122.0177 | 118.0157
R31M10 | FWHM (average) | 119.8 | 120.8 | 118.8 | 120.801 | 118.799 | 121.801 | 117.799
Plan ID | Parameter |  | Good: within +/-0.5% |  | Warning: within +/-1% |  | Action level: over 1%
| Symmetry | x | 0.5 | -0.5 | 0.5 | -0.5 | 1 | 1
Plan ID |  |  | Good: within +/-1% |  | Warning: within +/-2% |  | Action level: over 3%
| Flatness | x | 1 | -1 | 2 | -2 | 3 | -3
* reference measurements, Roos IC #2519, 18.11.2022
| R15M10 |  |  | R31M10 |  |  |  | R15_SOBP10_2Gy, 10 cm WET, 9.2 SW | [13 empty] | R31_SOBP10_2Gy , 26 cm WET, 24.8 SW | [17 empty] | R15_SOBP10_2Gy |  | R31_SOBP10_2Gy |  | Plot
Date | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | R31M10 vs. R15M10 | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Sign | Checked | File name | Notes | GI (1%, 1 mm) | GI (2%, 2 mm) | GI (1%, 1 mm) | GI (2%, 2 mm) | Date | Difference from reference (a.u.) R15M10 | Medelvärde +/- 10 dygn R15M10 | Medelvärde +/- 20 dygn R15M10 | Difference from reference (a.u.) R31M10 | Medelvärde +/- 10 dygn R31M10 | Medelvärde +/- 20 dygn R31M10 | Ref | 2% | -2% | 3% | -3%
20221118 | 1.985 | -0.1007% | -0.9481% | 2.007 | 0% | -0.9378% | 0.98904 | 0.86 | 0.82 | -0.42 | 0.57 | 1.31 | -0.36 | 0.48 | 1.985 | 8.7 | 9 | 8 | 8.1 | 119.8 | 120.2 | 0.3 | 0.49 | -0.11 | 0.23 | 0.73 | -0.14 | 0.18 | 2.007 | 11.4 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | ML | x | GTR2_20221118_R15M10_01, GTR2_20221118_R31M10_01 | REFERENCE IMAGES |  |  |  |  | 2022-11-18 | 0.99899 | 0.99104 | 0.99123 | 1 | 0.99043 | 0.98778 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20221118 | 1.982 | -0.2516% | -1.098% | 2.008 | 0.04983% | -0.8885% | 0.98705 | 0.81 | 0.8 | -0.41 | 0.51 | 1.23 | -0.36 | 0.47 | 1.982 | 8.7 | 9 | 8.1 | 8.1 | 119.8 | 120.2 | 0.3 | 0.47 | -0.12 | 0.12 | 0.74 | -0.12 | 0.13 | 2.008 | 11.3 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | ML | x | GTR2_20221118_R15M10_02, GTR2_20221118_R31M10_02 |  | 98.4 |  | 100 |  | 2022-11-18 | 0.99748 | 0.99085 | 0.99104 | 1.0005 | 0.99026 | 0.98767 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20221118 | 1.985 | -0.1007% | -0.9481% | 2.006 | -0.04983% | -0.9872% | 0.98953 | 0.84 | 0.75 | -0.45 | 0.53 | 1.24 | -0.35 | 0.45 | 1.985 | 8.7 | 9 | 8 | 8.2 | 119.8 | 120.2 | 0.35 | 0.44 | -0.12 | 0.17 | 0.7 | -0.12 | 0.15 | 2.006 | 11.4 | 11.3 | 11.3 | 11.1 | 119.6 | 120 | ML | x | GTR2_20221118_R15M10_03, GTR2_20221118_R31M10_03 |  | 96.7 |  | 100 |  | 2022-11-18 | 0.99899 | 0.99073 | 0.99108 | 0.9995 | 0.98954 | 0.98783 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20221213 | 1.981 | -0.302% | -1.148% | 1.997 | -0.4983% | -1.431% | 0.99199 | 0.91 | 0.73 | -0.41 | 0.44 | 1.31 | -0.37 | 0.45 | 1.981 | 9 | 8.6 | 7.9 | 8.2 | 119.9 | 120.2 | 0.39 | 0.41 | -0.14 | 0.08 | 0.83 | -0.15 | 0.11 | 1.997 | 11.3 | 11.3 | 11.1 | 11.2 | 119.6 | 120.1 | ML | x | GTR2_20221213_R15M10_01, GTR2_20221213_R31M10_01 |  | 99.7 |  | 94.4 |  | 2022-12-13 | 0.99698 | 0.99063 | 0.9911 | 0.99502 | 0.98938 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230216 | 1.971 | -0.8052% | -1.647% | 1.988 | -0.9467% | -1.876% | 0.99145 | 0.76 | 0.87 | -0.37 | 0.55 | 1.27 | -0.31 | 0.48 | 1.971 | 8.7 | 9 | 7.9 | 8.5 | 119.8 | 120.2 | 0.32 | 0.39 | -0.13 | -0.03 | 0.53 | -0.11 | -0.01 | 1.988 | 11.3 | 11.4 | 11.6 | 11.5 | 119.6 | 120 | JR | ok/ML | GTR2_20230216_R15M10_02, GTR2_20230216_R31M10_02 |  | 100 |  | 99.4 |  | 2023-02-16 | 0.99195 | 0.99047 | 0.9911 | 0.99053 | 0.98882 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230220 | 1.968 | -0.9562% | -1.796% | 1.984 | -1.146% | -2.073% | 0.99194 | 0.81 | 0.77 | -0.41 | 0.5 | 1.2 | -0.3 | 0.44 | 1.968 | 8.3 | 9.2 | 8 | 8.3 | 119.9 | 120.2 | 0.35 | 0.42 | -0.14 | 0.02 | 0.58 | -0.11 | 0.01 | 1.984 | 11.6 | 11.5 | 11.8 | 11.5 | 119.6 | 120 | AJ | ok/ML | GTR2_20230220_R15M10_01, GTR2_20230220_R31M10_01 |  | 91.2 |  | 74.4 | 98.4 | 2023-02-20 | 0.99044 | 0.99118 | 0.9911 | 0.98854 | 0.98871 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230221 | 1.956 | -1.56% | -2.395% | 1.977 | -1.495% | -2.419% | 0.98938 | 0.84 | 0.75 | -0.42 | 0.47 | 1.3 | -0.34 | 0.41 | 1.956 | 8.5 | 9.1 | 8.2 | 8.1 | 119.9 | 120.2 | 0.28 | 0.36 | -0.07 | 0.11 | 0.58 | -0.09 | 0.09 | 1.977 | 11.6 | 11.5 | 11.8 | 11.6 | 119.6 | 120 | AE | ok/ML | GTR2_20230221_R15M10_01, GTR2_20230221_R31M10_01 | Training session, beam not tunned | 54 | 100 | 56.3 | 93.8 | 2023-02-21 | 0.9844 | 0.99179 | 0.9911 | 0.98505 | 0.98857 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230221 | 1.955 | -1.61% | -2.445% | 1.973 | -1.694% | -2.616% | 0.99088 | 0.86 | 0.79 | -0.45 | 0.47 | 1.2 | -0.36 | 0.4 | 1.955 | 7.9 | 9.5 | 8.3 | 8.2 | 119.8 | 120.2 | 0.37 | 0.39 | -0.16 | -0.01 | 0.54 | -0.13 | -0.02 | 1.973 | 11.8 | 11.6 | 11.7 | 11.5 | 119.6 | 120 | AR | ok/ML | GTR2_20230221_R15M10_02 GTR2_20230221_R31M10_02 | Training session, beam not tunned | 51.9 | 100 | 56.3 | 87.5 | 2023-02-21 | 0.9839 | 0.99159 | 0.9911 | 0.98306 | 0.98833 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230224 | 1.957 | -1.51% | -2.345% | 1.975 | -1.594% | -2.517% | 0.99089 | 0.97 | 0.73 | -0.46 | 0.43 | 1.24 | -0.35 | 0.38 | 1.957 | 8.6 | 9.1 | 8 | 8.3 | 119.8 | 120.2 | 0.4 | 0.45 | -0.19 | -0.04 | 0.61 | -0.14 | 0.01 | 1.975 | 11.4 | 11.4 | 11.7 | 11.6 | 119.6 | 119.9 | OJ | AD | GTR2_20230224_R15M10_01, GTR2_20230224_R31M10_01 |  | 51.8 | 100 | 55.9 | 88.4 | 2023-02-24 | 0.9849 | 0.99142 | 0.9911 | 0.98406 | 0.98812 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230227 | 1.952 | -1.761% | -2.595% | 1.963 | -2.192% | -3.11% | 0.9944 | 1.04 | 0.77 | -0.57 | 0.45 | 1.29 | -0.38 | 0.33 | 1.952 | 8.6 | 9 | 8 | 8.3 | 119.9 | 120.2 | 0.2 | 0.41 | -0.06 | -0.09 | 0.51 | -0.03 | -0.03 | 1.963 | 11.5 | 11.4 | 11.7 | 11.5 | 119.6 | 120 | GK,AJ,AR,OJ,EH | AC | GTR2_20230227_R15M10_01, GTR2_20230227_R31M10_01 |  | 52.3 | 96.4 | 51.6 | 80 | 2023-02-27 | 0.98239 | 0.99123 | 0.9911 | 0.97808 | 0.98778 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230228 | 1.965 | -1.107% | -1.946% | 1.984 | -1.146% | -2.073% | 0.99042 | 0.89 | 0.77 | -0.47 | 0.53 | 1.26 | -0.36 | 0.43 | 1.965 | 8.5 | 9.1 | 8.3 | 7.9 | 119.9 | 120.3 | 0.55 | 0.53 | -0.25 | -0.13 | 0.7 | -0.14 | -0.09 | 1.984 | 11.4 | 11.4 | 11.6 | 11.6 | 119.5 | 119.9 | GF | YM | GTR2_20230228_R15M10_01, GTR2_20230228_R31M10_01 |  | 89.1 | 100 | 75.6 | 95.6 | 2023-02-28 | 0.98893 | 0.99062 | 0.9911 | 0.98854 | 0.98702 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230301 | 1.966 | -1.057% | -1.896% | 1.97 | -1.844% | -2.764% | 0.99797 | 0.84 | 0.72 | -0.42 | 0.45 | 1.23 | -0.33 | 0.4 | 1.966 | 8.8 | 8.9 | 7.7 | 8.4 | 119.9 | 120.3 | 0.28 | 0.39 | -0.12 | -0.05 | 0.65 | -0.12 | -0.06 | 1.97 | 11.3 | 11.3 | 11.5 | 11.3 | 119.7 | 120.1 | SM, AE | GB | GTR2_20230301_R15M10_01, GTR2_20230301_R31M10_01 |  | 91.3 | 100 | 60.6 | 90 | 2023-03-01 | 0.98943 | 0.99033 | 0.9911 | 0.98156 | 0.98652 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230302 | 1.966 | -1.057% | -1.896% | 1.982 | -1.246% | -2.172% | 0.99193 | 0.82 | 0.83 | -0.36 | 0.53 | 1.19 | -0.26 | 0.46 | 1.966 | 8.7 | 8.9 | 7.8 | 8.4 | 119.8 | 120.2 | 0.36 | 0.36 | -0.12 | 0.16 | 0.56 | -0.09 | 0.16 | 1.982 | 11.3 | 11.3 | 11.6 | 11.4 | 119.6 | 120 | JR | YM | GTR2_20230302_R15M10_01, GTR2_20230302_R31M10_01 | Remeber to record the GI vaules/ML |  |  |  |  | 2023-03-02 | 0.98943 | 0.98993 | 0.9911 | 0.98754 | 0.98592 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230303 | 1.964 | -1.158% | -1.996% | 1.97 | -1.844% | -2.764% | 0.99695 | 0.86 | 0.87 | -0.46 | 0.57 | 1.34 | -0.35 | 0.45 | 1.964 | 8.6 | 9 | 7.8 | 8.4 | 119.9 | 120.2 | 0.29 | 0.39 | -0.11 | -0.07 | 0.59 | -0.1 | -0.07 | 1.97 | 11.3 | 11.3 | 11.7 | 11.5 | 119.6 | 120 | SM, AR | AC | GTR2_20230303_R15M10_01, GTR2_20230303_R31M10_01 |  | 70.9 | 100 | 61.3 | 89.1 | 2023-03-03 | 0.98842 | 0.98954 | 0.9911 | 0.98156 | 0.98541 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230306 | 1.989 | 0.1007% | -0.7485% | 1.981 | -1.295% | -2.221% | 1.004 | 0.78 | 0.74 | -0.38 | 0.43 | 1.19 | -0.32 | 0.4 | 1.989 | 8.4 | 9.1 | 8 | 8 | 119.9 | 120.3 | 0.27 | 0.37 | -0.09 | 0.02 | 0.51 | -0.08 | 0.02 | 1.981 | 11.5 | 11.5 | 11.7 | 11.5 | 119.6 | 120 | ML | x | GTR2_20230306_R15M10_01, GTR2_20230306_R31M10_01 |  | 100 |  | 55.6 | 92.5 | 2023-03-06 | 1.001 | 0.9894 | 0.9911 | 0.98705 | 0.98511 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230306 | 1.989 | 0.1007% | -0.7485% | 1.98 | -1.345% | -2.27% | 1.0045 | [14 empty] | 0.3 | 0.34 | -0.12 | 0.04 | 0.55 | -0.09 | 0.03 | 1.98 | 11.5 | 11.5 | 11.7 | 11.4 | 119.5 | 120 | ML | x | GTR2_20230306_R31M10_02 | only large cube was repeated |  |  | 56.6 | 92.5 | 2023-03-06 | 1.001 | 0.98934 | 0.9911 | 0.98655 | 0.9849 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230307 | 1.964 | -1.158% | -1.996% | 1.976 | -1.545% | -2.468% | 0.99393 | 0.89 | 0.89 | -0.48 | 0.56 | 1.29 | -0.38 | 0.48 | 1.964 | 9 | 8.7 | 7.8 | 8.4 | 119.9 | 120.2 | 0.46 | 0.45 | 0.08 | 0.07 | 0.79 | 0.02 | 0.14 | 1.976 | 11.1 | 11.1 | 11 | 11.2 | 119.7 | 120 | ML | x | GTR2_20230307_R15M10_01, GTR2_20230307_R31M10_01 |  | 94.3 |  | 58.1 | 90.3 | 2023-03-07 | 0.98842 | 0.98967 | 0.9911 | 0.98455 | 0.98489 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230307 | 1.964 | -1.158% | -1.996% | 1.976 | -1.545% | -2.468% | 0.99393 | [14 empty] | 0.45 | 0.43 | 0.08 | 0.1 | 0.82 | 0.03 | 0.16 | 1.976 | 11.1 | 11.1 | 11 | 11.2 | 119.7 | 120 | ML | x | GTR2_20230307_R31M10_02 | only large cube was repeated |  |  |  |  | 2023-03-07 | 0.98842 | 0.99008 | 0.9911 | 0.98455 | 0.98502 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230308 | 1.963 | -1.208% | -2.046% | 1.97 | -1.844% | -2.764% | 0.99645 | 0.78 | 0.75 | -0.39 | 0.49 | 1.12 | -0.32 | 0.41 | 1.963 | 8.6 | 9 | 8 | 8.2 | 119.9 | 120.2 | 0.18 | 0.41 | -0.03 | -0.09 | 0.56 | -0.05 | -0.06 | 1.97 | 11.4 | 11.4 | 11.6 | 11.5 | 119.6 | 120 | ML | x | GTR2_20230308_R15M10_01, GTR2_20230308_R31M10_01 |  | 70.2 | 100 | 59.4 | 87.2 | 2023-03-08 | 0.98792 | 0.99048 | 0.9911 | 0.98156 | 0.98509 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230308 | 1.962 | -1.258% | -2.096% | 1.978 | -1.445% | -2.369% | 0.99191 | 0.68 | 0.82 | -0.34 | 0.46 | 1.09 | -0.3 | 0.39 | 1.962 | 8.5 | 9.1 | 7.7 | 8.4 | 119.8 | 120.3 | 0.27 | 0.38 | -0.1 | -0.04 | 0.51 | -0.08 | -0.03 | 1.978 | 11.3 | 11.4 | 11.6 | 11.4 | 119.6 | 120 | ML | x | GTR2_20230308_R15M10_02, GTR2_20230308_R31M10_02 | Large cube: region approx. 3x3cm (right, down corner of the field) where dose deviates from reference in av. -2.5% | 66.6 | 100 | 58.1 | 90.3 | 2023-03-08 | 0.98742 | 0.99115 | 0.9911 | 0.98555 | 0.98568 | 0.98784 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230309 | 1.971 | -0.8052% | -1.647% | 1.989 | -0.8969% | -1.826% | 0.99095 | 0.85 | 0.69 | -0.43 | 0.46 | 1.18 | -0.32 | 0.39 | 1.971 | 8.5 | 9.1 | 8.1 | 8.1 | 119.9 | 120.2 | 0.28 | 0.33 | -0.07 | 0.05 | 0.48 | -0.08 | 0.02 | 1.989 | 11.4 | 11.4 | 11.6 | 11.5 | 119.6 | 119.9 | EH | YM | GTR2_20230309_R15M10_01, GTR2_20230309_R31M10_01 |  | 100 | 100 | 80.3 | 99.1 | 2023-03-09 | 0.99195 | 0.99135 | 0.99073 | 0.99103 | 0.98541 | 0.98726 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230310 | 1.97 | -0.8556% | -1.697% | 1.983 | -1.196% | -2.122% | 0.99344 | 0.81 | 0.63 | -0.41 | 0.38 | 1.21 | -0.32 | 0.41 | 1.97 | 8.9 | 8.5 | 8.2 | 7.9 | 119.9 | 120.3 | 0.27 | 0.49 | -0.13 | -0.22 | 0.74 | -0.12 | -0.16 | 1.983 | 11.3 | 11.4 | 11.6 | 11.4 | 119.6 | 120 | AR, GK | KW | GTR2_20230310_R15M10_01, GTR2_20230310_R31M10_01 |  | 99.4 | 100 | 80.6 | 98.4 | 2023-03-10 | 0.99144 | 0.99155 | 0.99039 | 0.98804 | 0.9858 | 0.9866 | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
Formulas:
Plot = DATE(LEFT([Plan ID],4), MID([Plan ID],5,2), RIGHT([Plan ID],2))

## Safety and lasers
|  | Ljusgardin: | Mån, ons och fredag |  |  |  | Lasers: | distance in mm between the middle of the laser line and the marker on the wall (0.5 mm precision)
|  | Dörrinterlock: | Tis och torsdag
|  | Avsyningsprocedur: | Dagligen
|  | Nödstopp: | Första fredagen i månaden
Datum |  | G2
|  | Safety |  |  |  |  | Lasers [mm] |  |  |  | Dosimetry Manager |  | All
|  | Avsyning | varningslampor | ljusgardin | dörrfunktion | RS | Red1 | Green1 | Red2 | Green2 | T [K] | p [hPa] | Kommentar
2023-02-24 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 993.8
2023-02-27 | Måndag | ok | ok | ok |  | ok | 0 | 0 | 0 | 1 | 296.6 | 1019.9
2023-02-28 | Tisdag | ok | ok |  | ok | ok | 0 | 0 | -0.5 | 1 | 296.8 | 1021.6
2023-03-01 | Onsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1022
2023-03-02 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 0 | 1 | 296.6 | 1017.7
2023-03-03 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 0 | 1 | 296.8 | 1009.2
2023-03-06 | Måndag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.5 | 993.1
2023-03-07 | Tisdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.6 | 990.7
2023-03-08 | Onsdag | ok | ok | ok |  | ok | 0 | -0.5 | -0.5 | 1 | 296.6 | 979.6
2023-03-09 | Torsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.9 | 992.3
2023-03-10 | Fredag | ok | ok | ok |  | ok | 0 | 0 | 0 | 1 | 295.8 | 994.5
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag

## Signatures
Signature | First name | Last name | Profession | Note
AA | Ali | Alkhiat | Medical Physicist
AC | Athanasia | Christou | Summer physicist
AD | Alexandru | Dasu | Medical physicist
AE | Albin | Eriksson | Uppstuk physicist
AHL | Annika | Hall | Medical physicist
AJ | Anton | Jansson | Uppstuk physicist
AR | Adam | Röjvall | Uppstuk physicist
CVD | Christina | Vallhagen Dahlgren | Medical Physicist
DJ | Dan | Josefsson | Medical physicist
EA | Erik | Almhagen | Medical physicist
EF | Emil | Forsberg | Uppstuk physicist
EL | Elias | Lindbäck | Medical physicist
EN | Erik | Näslund | Uppstuk physicist
EPe | Erik | Pettersson | Medical physicist
GB | Gloria | Bäckström | Medical physicist
GF | Gustav | Furubjelke | Uppstuk physicist
JM | Joakim | Medin | Medical physicist
JR | Johan | Rensfeldt | Uppstuk physicist
GK | Gurpreet | Kaur | Uppstuk physicist
KW | Kenneth | Wikström | Medical physicist
LL | Linnea | Lund | Medical physicist
MB | Mikael | Blomqvist | Medical physicist
ME | Marika | Enmark | Medical physicist
MF | Marcus | Fager | Medical physicist
MJ | Mikael | Johansson | Uppstuk physicist
MK | Marcus | Krantz | Medical physicist
Ml | Malgorzata | Liszka | Medical physicist
NIC | Carl Magnus | Nilsson | Medical physicist
OJ | Olof | Jardfelt | Uppstuk physicist
ON | Olivier | Näslund | Uppstuk physicist
PL | Peter | Larsson | Medical physicist
SF | Stina | Fredriksson | Uppstuk physicist
SM | Sebastian | Mikkelsen | Uppstuk physicist
TH | Thomas | Henry | Medical physicist
UG | Ulf | Granlund | Medical physicist
YM | Younes | Majeddam | Medical physicist

## Updates
No | Date | Responsible | Change Description
1 | 2023-02-23 | ML | The document released for users
2
3
4
5
6