## GTR2_Setup check
ID.nr:
Titel: | [30 empty] | Dokumenttyp:
Protocol for Daily QA: Setup preparation fast check | [30 empty] | Strålskyddsinstruktion
Godkänt av: | [29 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [29 empty] | Skapat den:
Reviderat av: | [29 empty] | Reviderat den:
GTR2
Setup check: GTR2 MatriXX PT_3
Set-up: |  | 2.3 cm SW + MatrixPT position adjusted using X-rays (chambers at iso)
Detector |  | IBA MatrixPT_3
Phantom |  | SW
MyQA Project: |  | GTR2_DailyQA_Matrixx
Patient in AdaptDelivery |  | G2_Daily_MatriXX Skandion
Plan ID: |  | Single layer_E150MeV_F10x10
Detector Calib. Factor: |  | 85.483 | GTR2_Reference Calib. 2022-11-04
Reference dose (TPS) [Gy] |  | 1.699
Calibration dose [Gy] |  | 1.716
Good: within 0.5%
1.682 | 1.716
1.6988 | 1.7332
Warning: within 2%
1.717 | 1.681
1.7342 | 1.6978
Warning: over 2%
1.734 | 1.664
1.7513 | 1.6807
|  | Conditions in room |  |  | Detector condition | [6 empty] | IBA laser off CV (mm) | [6 empty] | Table position saved in Physic QA
Date | Responsible | p [hPa] | Troom [0C] | Troom [K] | Calibration Factor | TMatrixx [0C] | pMatriXX [hPa] | MyQA TMatrixx [0C] | MyQA pMatriXX [hPa] | MyQA Offset TMatrixx [0C] | MyQA Offset pMatriXX [hPa] | Δx (cm) | Δy (cm) | Δz (cm) | Rot (o) | Pitcht (o) | Roll (o) | CV (cm) | X (cm) | Y (cm) | Z (cm) | Rot (o) | Pitcht (o) | Roll (o) | CV (mm) | Z (cm) after applying Offset -3 cm | Dcentral ROI [1 cm x 1 cm] | Difference from reference TPS (%) | Difference from Calibration dose (%) | GI (1%, 1 mm) | GI (2%, 2 mm) | Notes
2022-11-18 | AC | 1017 | 23.7 | 296.85 | 85.483 | 23.3 | 1017 | 23.9 | 1028 | -0.6 | -11 | [15 empty] | 1.707 | 0.4709% | -0.5245% | 100 |  | run 01
2022-11-18 | [26 empty] | 1.709 | 0.5886% | -0.4079% | 100 |  | run 02
2022-11-18 | [26 empty] | 1.708 | 0.5297% | -0.4662% |  |  | run03, REFERENCE IMAGE
2023-03-13 | AJ | 993.8 | 23.3 | 296.45 | 85.483 | 22.7 | 993.8 | 23.8 | 1005 | -1.1 | -11.2 | -0.11 | 0.08 | 0.07 | 0.1 | 0 | 0 | 0.15297 | -0.11 | 0.09 | 0.06 | 0.1 | 0 | 0 | 0.15427 | -2.94 | 1.6875 | -0.6769% | -1.661% | 64.3 | 100 | MyQA offset applied/ Very good ML
2023-03-14 | SM, EH | 968 | 24.1 | 297.25 | 85.483 | 22.7 | 968 | 23.2 | 979 | -0.5 | -11 | 0.07 | 0.08 | 0.03 | 0 | 0 | 0 | 0.11045 | 0.07 | 0.08 | 0.03 | 0 | 0 | 0 | 0.11045 | -2.97 | 1.688 | -0.6474% | -1.632% | 72.6 | 100
2023-03-15 | GF | 989.5 | 24.1 | 297.25 | 85.483 | 22.9 | 989.5 | 23.7 | 1001 | -0.8 | -11.5 | 0.08 | 0.06 | 0.05 | 0 | 0 | 0 | 0.1118 | 0.08 | 0.06 | 0.05 | 0 | 0 | 0 | 0.1118 | -2.95 | 1.6892 | -0.5774% | -1.562% | 76.6 | 100
2023-03-16 | EH, SM | 1010 | 23.6 | 296.75 | 85.483 | 22.6 | 1010 | 23.1 | 1021 | -0.5 | -11 | 0.07 | 0.07 | 0.02 | -0.2 | 0 | 0 | 0.101 | 0.07 | 0.07 | 0.05 | 359.8 | 0 | 0 | 0.11091 | -2.95 | 1.69 | -0.5297% | -1.515% | 80.5 | 100
2023-03-20 | AC,ML | 1001.7 | 23.7 | 296.85 | 85.483 | 22.7 | 1001.7 | 23.1 | 1012 | -0.4 | -10.3 | 0.07 | 0.06 | 0.03 | 0 | 0 | 0 | 0.096954 | 0.07 | 0.06 | 0.05 | 0 | 0 | 0 | 0.10488 | -2.95 | 1.6896 | -0.5533% | -1.538% | 74.9 | 100
2023-03-21 | GF,AE | 1007.3 | 23.8 | 296.95 | 85.483 | 22.7 | 1007.3 | 23.2 | 1018 | -0.5 | -10.7 | 0.07 | 0.08 | 0.03 | 0 | 0 | 0 | 0.11045 | 0.07 | 0.08 | 0.03 | 0 | 0 | 0 | 0.11045 | -2.97 | 1.6926 | -0.379% | -1.366% | 90.7 | 100 | 10 min wait time for MyQA calibration…
2023-03-22 | SF, AE | 996 | 23.9 | 297.05 | 85.483 | 22.7 | 996 | 23.4 | 1006 | -0.7 | -10 | -0.09 | 0.19 | 0.05 | 0.1 | 0 | 0 | 0.2161 | -0.09 | 0.2 | 0.04 | 0.1 | 0 | 0 | 0.22293 | -2.96 | 1.6929 | -0.3567% | -1.344% | 94.9 | 100
2023-03-23 | SM, OJ | 988.2 | 23.7 | 296.85 | 85.483 | 22.6 | 988.2 | 23.3 | 999 | -0.7 | -10.8 | 0.07 | 0.07 | 0.04 | -0.1 | 0 | 0 | 0.10677 | 0.07 | 0.07 | 0.05 | 359.9 | 0 | 0 | 0.11091 | -2.95 | 1.693 | -0.3531% | -1.34% | 95.2 | 100
2023-03-24 | SM, EH | 988.6 | 23.2 | 296.35 | 85.483 | 22.4 | 988.6 | 23.2 | 999 | -0.8 | -10.4 | 0.07 | 0.08 | 0.02 | 0 | 0 | 0 | 0.10817 | 0.07 | 0.08 | 0.07 | 0 | 0 | 0 | 0.12728 | -2.93 | 1.6903 | -0.5121% | -1.498% | 80 | 100
2023-03-27 | AJ, GK | 997.7 | 23.8 | 296.95 | 85.483 | 22.6 | 997.7 | 23.6 | 1008 | -1 | -10.3 | -0.05 | 0.07 | 0.03 | 0 | 0 | 0 | 0.091104 | -0.05 | 0.07 | 0.03 | 0 | 0 | 0 | 0.091104 | -2.97 | 1.6835 | -0.9147% | -1.896% | 59.5 | 100
2023-03-28 | AE, GF | 1005 | 24.1 | 297.25 | 85.483 | 23 | 1005 | 23.3 | 1016 | -0.3 | -11 | -0.05 | 0.05 | 0.02 | -0.1 | 0 | 0 | 0.073485 | -0.05 | 0.05 | 0.07 | 359.9 | 0 | 0 | 0.099499 | -2.93 | 1.6863 | -0.7451% | -1.728% | 58.9 | 100
2023-03-28 | GK,ON,EH | 1012.1 | 23.6 | 296.75 | 85.483 | 23.2 | 1012.1 | 23.7 | 1023 | -0.5 | -10.9 | 0.04 | 0.08 | 0.03 | 0 | 0 | 0 | 0.09434 | 0.04 | 0.08 | 0.08 | 0 | 0 | 0 | 0.12 | -2.92 | 1.6949 | -0.2413% | -1.23% | 99.1 | 100
2023-03-30 | JR,ON | 1009.1 | 23.5 | 296.65 | 85.483 | 22.8 | 1009.1 | 22.8 | 1020 | 0 | -10.9 | 0.09 | 0.07 | 0.15 | 0 | 0 | 0 | 0.18841 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0.01 | -2.93 | 1.6886 | -0.6092% | -1.594% | 74.9 | 100
2023-03-31 | AR,SM | 1001.5 | 23.7 | 296.85 | 85.483 | 22.6 | 1001.5 | 23.2 | 1013 | -0.6 | -11.5 | 0.07 | 0.07 | 0.05 | 0 | 0 | 0 | 0.11091 | 0.07 | 0.07 | 0.07 | 0 | 0 | 0 | 0.12124 | -2.93 | 1.6903 | -0.5097% | -1.495% | 72.3 | 100
2023-04-03 | SM, AR | 1032.1 | 23.6 | 296.75 | 85.483 | 22.7 | 1032.1 | 23.5 | 1043 | -0.8 | -10.9 | 0.01 | 0.06 | 0 | -0.1 | 0 | 0 | 0.060828 | 0.01 | 0.07 | 0.08 | 359.9 | 0 | 0 | 0.10677 | -2.92 | 1.6928 | -0.3643% | -1.351% | 98.2 | 100
2023-04-04 | ON, AR, AJ | 1024.6 | 23.7 | 296.85 | 85.483 | 22.9 | 1024.6 | 23.2 | 1036 | -0.3 | -11.4 | -0.04 | 0.07 | 0 | -0.3 | 0 | 0 | 0.080623 | -0.04 | 0.08 | 0.12 | 359.7 | 0 | 0 | 0.14967 | -2.88 | 1.6893 | -0.5709% | -1.556% | 78.9 | 100
2023-04-05 | AE, OJ | 1021.3 | 23.6 | 296.75 | 85.483 | 22.8 | 1021.3 | 23.2 | 1033 | -0.4 | -11.7 | 0.08 | 0.06 | 0 | -0.1 | 0 | 0 | 0.1 | 0.08 | 0.06 | 0.1 | 359.9 | 0 | 0 | 0.14142 | -2.89 | 1.698 | -0.06062% | -1.051% | 100 | 100
2023-04-06 | EH, GK | 1022.3 | 23.6 | 296.75 | 85.483 | 22.8 | 1022.3 | 23.1 | 1034 | -0.3 | -11.7 | 0.06 | 0.08 | 0.05 | -0.1 | 0 | 0 | 0.1118 | 0.06 | 0.08 | 0.12 | 359.9 | 0 | 0 | 0.1562 | -2.88 | 1.6951 | -0.2272% | -1.216% | 100
2023-04-07 | SM, JR | 1023.2 | 23.6 | 296.75 | 85.483 | 22.5 | 1023.2 | 22.9 | 1034 | -0.4 | -10.8 | 0.08 | 0.05 | 0.05 | 0 | 0 | 0 | 0.10677 | 0.08 | 0.06 | 0.14 | 0 | 0 | 0 | 0.17205 | -2.86 | 1.6953 | -0.2178% | -1.206% | 96.3 | 100
2023-04-11 | GF,EH,ON | 1010.5 | 23.2 | 296.35 | 85.483 | 22.9 | 1010.5 | 23.5 | 1021 | -0.6 | -10.5 | 0.05 | 0.08 | 0 | 0 | 0 | 0 | 0.09434 | 0.05 | 0.08 | 0.11 | 0 | 0 | 0 | 0.14491 | -2.89 | 1.6905 | -0.4974% | -1.483% | 80.9 | 100
2023-04-12 | JR,GF | 1005.5 | 23.6 | 296.75 | 85.483 | 22.7 | 1005.5 | 23.2 | 1016 | -0.5 | -10.5 | -0.01 | 0.06 | 0 | 0 | 0 | 0 | 0.060828 | -0.01 | 0.06 | 0.11 | 0 | 0 | 0 | 0.1257 | -2.89 | 1.693 | -0.3531% | -1.34% | 93.2 | 100 | ML corrected the value of Dcentral
2023-04-13 | SM, AE | 1005.8 | 23.8 | 296.95 | 85.483 | 22.4 | 1005.8 | 23.4 | 1017 | -1 | -11.2 | 0.08 | 0.08 | 0.02 | 0 | 0 | 0 | 0.11489 | 0.07 | 0.08 | 0.15 | 0 | 0 | 0 | 0.18385 | -2.85 | 1.686 | -0.7663% | -1.749% | 57.2 | 100
2023-04-14 | OJ, SM | 1014 | 23.7 | 296.85 | 85.483 | 22.6 | 1014 | 23.1 | 1025 | -0.5 | -11 | 0.07 | 0.07 | -0.03 | 0 | 0 | 0 | 0.10344 | 0.07 | 0.07 | 0.13 | 0 | 0 | 0 | 0.1634 | -2.87 | 1.686 | -0.7652% | -1.748% | 70.3 | 100
2023-04-17 | OJ, AJ | 1026 | 23.4 | 296.55 | 85.483 | 22.4 | 1026 | 23.2 | 1037 | -0.8 | -11 | -0.02 | 0.1 | 0.12 | -0.1 | 0 | 0 | 0.15748 | 0.02 | 0.1 | 0.12 | 359.9 | 0 | 0 | 0.15748 | -2.88 | 1.689 | -0.5886% | -1.573% | 75.7 | 100
2023-04-18 | GK,GF | 1032.4 | 23.6 | 296.75 | 85.483 | 22.7 | 1032.4 | 23.3 | 1044 | -0.6 | -11.6 | 0 | 0.07 | 0.03 | 0 | 0 | 0 | 0.076158 | 0 | 0.07 | 0.1 | 0 | 0 | 0 | 0.12207 | -2.9 | 1.689 | -0.5886% | -1.573% | 73.1 | 100
2023-04-19 | OJ, AR | 1033.2 | 23.7 | 296.85 | 85.483 | 22.7 | 1033.2 | 23.1 | 1044 | -0.4 | -10.8 | -0.02 | 0.06 | 0.02 | -0.1 | 0 | 0 | 0.066332 | -0.02 | 0.06 | 0.14 | 359.9 | 0 | 0 | 0.15362 | -2.86 | 1.687 | -0.7063% | -1.69% | 60.5 | 100
2023-04-20 | ON, EH | 1029.3 | 23.2 | 296.35 | 85.483 | 22.6 | 1029.3 | 22.9 | 1041 | -0.3 | -11.7 | 0.08 | 0.07 | 0.08 | 0 | 0 | 0 | 0.13304 | 0.08 | 0.06 | 0.12 | 0 | 0 | 0 | 0.1562 | -2.88 | 1.689 | -0.5886% | -1.573% | 99.5 | 100
2023-04-21 | SM, ON | 1022.1 | 23.7 | 296.85 | 85.483 | 22.4 | 1022.1 | 23.5 | 1033 | -1.1 | -10.9 | 0.04 | 0.07 | 0 | 0 | 0 | 0 | 0.080623 | 0.04 | 0.07 | 0.11 | 0 | 0 | 0 | 0.13638 | -2.89 | 1.689 | -0.5886% | -1.573% | 66.4 | 100
2023-04-24 | GK,AE | 999.9 | 23.4 | 296.55 | 85.483 | 23.2 | 999.9 | 23.5 | 1012 | -0.3 | -12.1 | 0.05 | 0.04 | 0.05 | 0 | 0 | 0 | 0.08124 | 0.05 | 0.05 | 0.07 | 0 | 0 | 0 | 0.099499 | -2.93 | 1.6903 | -0.5097% | -1.495% | 80.3 | 100 | ok/YM
2023-04-25 | AE, GF | 995.9 | 23.7 | 296.85 | 85.483 | 22.9 | 995.9 | 23.3 | 1007 | -0.4 | -11.1 | -0.05 | 0.05 | -0.02 | -0.1 | 0 | 0 | 0.073485 | -0.05 | 0.05 | 0.12 | 359.9 | 0 | 0 | 0.13928 | -2.88 | 1.6902 | -0.5209% | -1.506% | 80 | 100
2023-04-26 | AR, GK | 995.4 | 23.2 | 296.35 | 85.483 | 22.9 | 995.4 | 23.4 | 1007 | -0.5 | -11.6 | 0.18 | 0 | 0 | 0 | 0 | 0 | 0.18 | 0.11 | 0.07 | 0.12 | 359.9 | 0 | 0 | 0.1772 | -2.88 | 1.704 | 0.2943% | -0.6993% | 100 | 100 | ok/YM
2023-04-27 | GK,OJ | 1000.7 | 23.7 | 296.85 | 85.483 | 24 | 1000.7 | 23.1 | 1012 | 0.9 | -11.3 | 0.1 | 0.08 | 0.02 | 0 | 0 | 0 | 0.12961 | 0.1 | 0.08 | 0.08 | 0 | 0 | 0 | 0.151 | -2.91 | 1.7027 | 0.2166% | -0.7762% | 100 | 100 | ok/CVD
2023-04-28 | GK,JR | 1007.7 | 23.3 | 296.45 | 85.483 | 22.8 | 1007.7 | 23.4 | 1019 | -0.6 | -11.3 | 0.08 | 0.08 | 0.07 | 0 | 0 | 0 | 0.13304 | 0.08 | 0.08 | 0.07 | 0 | 0 | 0 | 0.13304 | -2.92 | 1.692 | -0.4114% | -1.398% | 86.1 | 100 | ok/AC
2023-05-02 | AE, AR, EF | 1002.1 | 23 | 296.15 | 85.483 | 22.4 | 1002.1 | 23.4 | 1014 | -1 | -11.9 | -0.09 | 0.09 | 0.02 | 0 | 0 | 0 | 0.12884 | 0.09 | 0.1 | 0.11 | 0 | 0 | 0 | 0.17378 | -2.89 | 1.6932 | -0.3402% | -1.328% | 89.4 | 100 | ok/CVD
2023-05-03 | ON, AJ, EF | 1015.2 | 23.3 | 296.45 | 85.483 | 22.6 | 1015.2 | 23.5 | 1027 | -0.9 | -11.8 | -0.09 | 0.09 | 0.07 | 0 | 0 | 0 | 0.14526 | -0.16 | 0.2 | 0.12 | 0 | -0.1 | 0 | 0.28284 | -2.88 | 1.6922 | -0.4002% | -1.387% | 89.4 | 100 | ok/CVD
2023-05-04 | JR | 1024.3 | 23.4 | 296.55 | 85.483 | 22.8 | 1024.3 | 23.4 | 1036 | -0.6 | -11.7 | 0.08 | 0.07 | 0.03 | 0 | 0 | 0 | 0.11045 | 0.08 | 0.07 | 0.11 | 0 | 0 | 0 | 0.15297 | -2.89 | 1.6942 | -0.2802% | -1.268% | 98.6 | 100 | ok/YM
2023-05-04 | ON, EF | 1025.5 | 23.9 | 297.05 | 85.483 | 22.9 | 1025.5 | 23.6 | 1037 | -0.7 | -11.5 | 0.07 | 0.07 | 0.12 | -0.1 | 0 | 0 | 0.15556 | 0.07 | 0.07 | 0.14 | 359.9 | 0 | 0 | 0.17146 | -2.86 | 1.696 | -0.1766% | -1.166% | 98.6 | 100
2023-05-05 | SM, GK | 1027.9 | 23.3 | 296.45 | 85.483 | 22.7 | 1027.9 | 23.1 | 1039 | -0.4 | -11.1 | 0.08 | 0.07 | 0.08 | 0 | 0 | 0 | 0.13304 | 0.08 | 0.07 | 0.12 | 0 | 0 | 0 | 0.16031 | -2.88 | 1.695 | -0.2354% | -1.224% | 99.1 | 100 | ok/AC
2023-05-08 | OJ, SF | 1025.3 | 23.6 | 296.75 | 85.483 | 23.1 | 1025.3 | 23.7 | 1036 | -0.6 | -10.7 | 0.07 | 0.08 | 0.07 | 0.1 | 0 | 0 | 0.12728 | 0.07 | 0.08 | 0.07 | 0.1 | 0 | 0 | 0.12728 | -2.93 | 1.69 | -0.5297% | -1.515% | 75.4 | 100 | ok/AC
2023-05-09 | GF, OJ, EN | 1019.9 | 23.6 | 296.75 | 85.483 | 22.8 | 1019.9 | 24.4 | 1031 | -1.6 | -11.1 | 0.08 | 0.04 | 0 | -0.1 | 0 | 0 | 0.089443 | 0.08 | 0.04 | 0 | 359.9 | 0 | 0 | 0.089443 | -3 | 1.69 | -0.5297% | -1.515% | 79.3 | 100 | ok/GB
2023-05-10 | ON, OJ | 1015.8 | 23.5 | 296.65 | 85.483 | 22.4 | 1015.8 | 23.7 | 1027 | -1.3 | -11.2 | 0.09 | 0.07 | 0 | -0.1 | 0 | 0 | 0.11402 | 0.09 | 0.07 | 0.05 | 359.9 | 0 | 0 | 0.1245 | -2.95 | 1.693 | -0.3531% | -1.34% | 93.1 | 100 | ok/GB
2023-05-11 | AR, GK | 1017.8 | 23.3 | 296.45 | 85.483 | 22.7 | 1017.8 | 24.2 | 1027 | -1.5 | -9.2 | 0.08 | 0.09 | 0.03 | 0 | 0 | 0 | 0.1241 | 0.08 | 0.09 | 0.04 | 0 | 0 | 0 | 0.12689 | -2.96 | 1.695 | -0.2354% | -1.224% | 100 | 100 | ok/CVD
2023-05-12 | SF, JR | 1021.4 | 23.6 | 296.75 | 85.483 | 22.9 | 1021.3 | 23.4 | 1033 | -0.5 | -11.7 | 0.09 | 0.08 | -0.02 | 0 | 0 | 0 | 0.12207 | 0.09 | 0.08 | 0.1 | 0 | 0 | 0 | 0.15652 | -2.9 | 1.697 | -0.1177% | -1.107% | 100 | 100 | ok/EAL
2023-05-15 | OJ, ON, EN | 1013.5 | 23.6 | 296.75 | 85.483 | 22.7 | 1013.5 | 23.5 | 1024 | -0.8 | -10.5 | 0.08 | 0.07 | 0.03 | -0.01 | 0 | 0 | 0.11045 | 0.08 | 0.07 | 0.1 | 359.9 | 0 | 0 | 0.14595 | -2.9 | 1.69 | -0.5297% | -1.515% | 76 | 100 | ok/GB
2023-05-16 | AJ,AR,EN | 1003.4 | 23.8 | 296.95 | 85.483 | 22.9 | 1003.4 | 23.4 | 1013 | -0.5 | -9.6 | 0.07 | 0.04 | 0.07 | -0.1 | 0 | 0 | 0.10677 | 0.07 | 0.05 | 0.09 | 359.9 | 0 | 0 | 0.1245 | -2.91 | 1.6926 | -0.3761% | -1.363% | 91.7 | 100 | ok/ON YM
2023-05-17 | AE, OJ, EF | 997.7 | 23.2 | 296.35 | 85.483 | 22.9 | 997.7 | 23.3 | 1009 | -0.4 | -11.3 | 0.07 | 0.08 | 0 | -0.1 | 0 | 0 | 0.1063 | 0.07 | 0.08 | 0.11 | 359.9 | 0 | 0 | 0.15297 | -2.89 | 1.7012 | 0.1313% | -0.8607% | 100 | 100 | ok / YM ON
2023-05-19 | ON, SM | 1025.2 | 23.4 | 296.55 | 85.483 | 22.7 | 1025.2 | 23.5 | 1036 | -0.8 | -10.8 | 0.07 | 0.07 | 0.2 | -0.2 | 0 | 0 | 0.22316 | 0.07 | 0.07 | 0.05 | 359.8 | 0 | 0 | 0.11091 | -2.95 | 1.707 | 0.4709% | -0.5245% | 100 |  | AD
2023-05-22 | AE, EF | 1020.8 | 23.4 | 296.55 | 85.483 | 22.8 | 1020.8 | 24.1 | 1032 | -1.3 | -11.2 | -0.03 | 0 | 0.05 | 0 | 0 | 0 | 0.05831 | -0.09 | 0.24 | 0.05 | 0 | -0.1 | 0 | 0.26115 | -2.95 | 1.6913 | -0.455% | -1.441% | 88.9 | 100 | GB
2023-05-23 | GF, ON | 1010.6 | 23.7 | 296.85 | 85.483 | 22.8 | 1010.6 | 23.3 | 1021 | -0.5 | -10.4 | 0.06 | 0.08 | 0.03 | -0.1 | 0 | 0 | 0.1044 | 0.04 | 0.06 | 0.03 | 359.9 | 0 | 0 | 0.078102 | -2.97 | 1.6917 | -0.4303% | -1.417% | 88.3 | 100 | CVD
2023-05-24 | AE, SM | 1012.5 | 24.1 | 297.25 | 85.483 | 22.8 | 1012.5 | 23.2 | 1023 | -0.4 | -10.5 | 0.07 | 0.06 | 0.03 | -0.1 | 0 | 0 | 0.096954 | 0.07 | 0.06 | 0.07 | 359.9 | 0 | 0 | 0.11576 | -2.93 | 1.6922 | -0.3985% | -1.385% | 87.2 | 100 | AD
2023-05-25 | AR, SM | 1009.8 | 24 | 297.15 | 85.483 | 22.8 | 1009.8 | 23.2 | 1021 | -0.4 | -11.2 | 0.07 | 0.07 | 0 | -0.1 | 0 | 0 | 0.098995 | 0.07 | 0.08 | 0.04 | 359.9 | 0 | 0 | 0.11358 | -2.96 | 1.6891 | -0.5827% | -1.568% | 73 | 100 | GB
2023-05-26 | OJ, EN | 1010.6 | 23.6 | 296.75 | 85.483 | 23.1 | 1010.6 | 23.7 | 1022 | -0.6 | -11.4 | 0.05 | 0.04 | 0.03 | 0 | 0 | 0 | 0.070711 | -0.02 | 0.18 | 0.07 | 0 | -0.1 | 0 | 0.19416 | -2.93 | 1.69 | -0.5297% | -1.515% | 80.7 | 100 | AC
2023-05-29 | OJ | 1016.4 | 23.2 | 296.35 | 85.483 | 22.6 | 1016.4 | 23.8 | 1027 | -1.2 | -10.6 | 0.07 | 0.07 | 0.08 | -0.1 | 0 | 0 | 0.12728 | ,0,7 | 0.07 | 0.07 | 359.9 | 0 | 0 |  | -2.94 | 1.688 | -0.6474% | -1.632% | 65.9 | 100 | CVD
2023-05-30 | AE, ON | 1018.5 | 23.2 | 296.35 | 85.483 | 22.6 | 1018.5 | 23 | 1029 | -0.4 | -10.5 | 0.05 | 0.04 | 0.03 | -0.1 | 0 | 0 | 0.070711 | 0.05 | 0.04 | 0.04 | 359.9 | 0 | 0 |  | -2.95 | 1.6904 | -0.5074% | -1.493% | 83 | 100 | ok/YM
2023-05-31 | JR | 1011.8 | 23.5 | 296.65 | 85.483 | 22.8 | 1011.8 | 22.2 | 1022 | 0.6 | -10.2 | 0.07 | 0.08 | 0.05 | -0.1 | 0 | 0 | 0.11747 | 0.06 | 0.05 | 0.05 | 395.9 | 0 | 0 |  | -2.95 | 1.69 | -0.528% | -1.513% | 76 | 100 | ok/YM
2023-06-01 | EF, AR | 1012.2 | 23.3 | 296.45 | 85.483 | 23.3 | 1012.2 | 23.7 | 1023 | -0.4 | -10.8 | -0.04 | 0.11 | 0.03 | -0.1 | 0 | 0 | 0.12083 | -0.1 | 0.24 | 0.06 | 359.9 | -0.1 | 0 |  | -2.94 | 1.6986 | -0.02472% | -1.015% | 100 |  | AD
2023-06-02 | AJ | 1015.9 | 23.7 | 296.85 | 85.483 | 22.7 | 1015.9 | 23.5 | 1027 | -0.8 | -11.1 | 0.07 | 0.1 | 0.05 | -0.1 | 0 | 0 | 0.13191 | 0.07 | 0.1 | 0.04 | 359.9 | 0 | 0 |  | -2.96 | 1.6887 | -0.6045% | -1.589% | 80.1 | 100 | AC
2023-06-05 | EF, AJ | 1016.5 | 23.4 | 296.55 | 85.483 | 22.9 | 1016.5 | 23.9 | 1027 | -1 | -10.5 | -0.02 | 0.12 | 0.05 | 0 | 0 | 0 | 0.13153 | -0.02 | -0.12 | 0.05 | 0 | 0 | 0 |  | -2.95 | 1.6936 | -0.3184% | -1.306% | 94.1 | 100
2023-06-07 | ON, AR | 1018.4 | 23.8 | 296.95 | 85.483 | 22.9 | 1018.4 | 23.3 | 1028 | -0.4 | -9.6 | 0.07 | 0.07 | 0 | -0.2 | 0 | 0 | 0.098995 | 0.07 | 0.07 | 0.06 | 359.8 | 0 | 0 |  | -2.94 | 1.701 | 0.1177% | -0.8741% | 100 |  | ok/YM
2023-06-08 | ON,EN | 1013.7 | 23.9 | 297.05 | 85.483 | 22.8 | 1013.7 | 23.6 | 1025 | -0.8 | -11.3 | 0.08 | 0.07 | 0.05 | -0.1 | 0 | 0 | 0.11747 | 0.08 | 0.08 | 0.09 | 359.9 | 0 | 0 |  | -2.91 | 1.686 | -0.7652% | -1.748% | 100 |  | ok/GB
2023-06-09 | AR, EF | 1021.5 | 23.3 | 296.45 | 85.483 | 23.1 | 1021.5 | 23.7 | 1032 | -0.6 | -10.5 | -0.05 | 0.09 | 0.02 | 0 | 0 | 0 | 0.10488 | -0.05 | 0.09 | 0.11 | 0 | 0 | 0 |  | -2.89 | 1.7035 | 0.2631% | -0.7302% | 100 |  | AC
2023-06-12 | GK | 1023.4 | 23.5 | 296.65 | 85.483 | 22.5 | 1023.4 | 23.9 | 1034 | -1.4 | -10.6 | 0.08 | 0.08 | -0.02 | 0 | 0 | 0 | 0.11489 | 0.08 | 0.08 | 0.07 | 0 | 0 | 0 |  | -2.93 | 1.7052 | 0.3632% | -0.6311% | 100 | 100 | AC
2023-06-13 | GK, OJ | 1018 | 23.7 | 296.85 | 85.483 | 22.8 | 1018 | 23.6 | 1028 | -0.8 | -10 | 0.09 | 0.07 | 0.02 | 0 | 0 | 0 | 0.11576 | 0.09 | 0.09 | 0.1 | 0 | 0 | 0 |  | -2.9 | 1.699 | 0% | -0.9907% | 100 |  | GB
2023-06-14 | EF, EN | 1017.8 | 23.9 | 297.05 | 85.483 | 22.8 | 1017.8 | 23.6 | 1028 | -0.8 | -10.2 | 0.06 | 0.06 | 0 | -0.1 | 0 | 0 | 0.084853 | -0.02 | 0.19 | 0.08 | 359.9 | -0.1 | 0 |  | -2.92 | 1.7041 | 0.2996% | -0.6941% | 100 |  | ok / YM
2023-06-15 | EF, EN | 1017 | 23.7 | 296.85 | 85.483 | 22.8 | 1017 | 23.5 | 1028 | -0.7 | -11 | 0.07 | 0.06 | 0 | 0 | 0 | 0 | 0.092195 | 0.07 | 0.06 | 0.11 | 0 | 0 | 0 |  | -2.89 | 1.7007 | 0.09947% | -0.8922% | 100 |  | ok / YM
2023-06-16 | EN, EF | 1013.7 | 23.2 | 296.35 | 85.483 | 22.7 | 1013.7 | 23.7 | 1024 | -1 | -10.3 | 0.06 | 0.08 | 0 | 0 | 0 | 0 | 0.1 | 0.06 | 0.08 | 0.07 | 0 | 0 | 0 |  | -2.93 | 1.6943 | -0.279% | -1.267% | 97.7 | 100 | AC
2023-06-19 | EN, OJ | 1012.5 | 23.2 | 296.35 | 85.483 | 22.6 | 1012.5 | 23.1 | 1023 | -0.5 | -10.5 | 0.08 | 0.07 | 0.02 | -0.01 | 0 | 0 | 0.10817 | 0.08 | 0.07 | 0.12 | 359.9 | 0 | 0 |  | -2.88 | 1.721 | 1.295% | 0.2914% | 90.6 | 100 | ML corrected the value of Dcentral
2023-06-20 | GK | 1012.8 | 23.5 | 296.65 | 85.483 | 23.2 | 1012.8 | 23.5 | 1023 | -0.3 | -10.2 | 0.09 | 0.09 | 0.08 | 0 | 0 | 0 | 0.15033 | 0.09 | 0.09 | 0.01 | 0 | 0 | 0 |  | 2.99 | 1.697 | -0.1177% | -1.107% | 89.4 | 100 | GB
2023-06-21 | OJ, GK | 1011.9 | 23.6 | 296.75 | 85.483 | 23.1 | 1011.9 | 23.3 | 1023 | -0.2 | -11.1 | 0.04 | 0.08 | -0.02 | 0 | 0 | 0 | 0.091652 | 0.04 | 0.07 | 0.02 | 0 | 0 | 0 |  | -2.98 | 1.699 | 0% | -0.9907% | 100 |  | CVD
2023-06-22 | ON | 1006.1 | 23.7 | 296.85 | 85.483 | 22.7 | 1006.1 | 23.2 | 1016 | -0.5 | -9.9 | 0.08 | 0.08 | 0.05 | -0.1 | 0 | 0 | 0.12369 | 0 | 0.21 | 0.04 | 359.9 | 0 | 0 |  | -2.96 | 1.703 | 0.2354% | -0.7576% | 100 |  | ok / ML YM
2023-06-26 | AR, GK | 1011.7 | 23.6 | 296.75 | 85.483 | 22.6 | 1011.7 | 23.2 | 1022 | -0.6 | -10.3 | 0.07 | 0.04 | 0.02 | -0.2 | 0 | 0 | 0.083066 | 0.07 | 0.04 | 0.04 | 359.8 | 0 | 0 |  | -2.96 | 1.704 | 0.2943% | -0.6993% | 100 |  | ok/ML
2023-06-27 | EN, OJ | 1003.9 | 23.6 | 296.75 | 85.483 | 22.6 | 1003.9 | 22.7 | 1014 | -0.1 | -10.1 | 0.06 | 0.07 | 0.01 | -0.1 | 0 | 0 | 0.092736 | 0.06 | 0.07 | 0.02 | 359.9 | 0 | 0 |  | -2.98 | 1.707 | 0.4709% | -0.5245% | 100 |  | ok/AC
2023-06-28 | AR, GK | 1007.5 | 23.7 | 296.85 | 85.483 | 22.6 | 1007.5 | 22.8 | 1018 | -0.2 | -10.5 | 0.08 | 0.08 | -0.02 | -0.1 | 0 | 0 | 0.11489 | 0.08 | 0.08 | -0.05 | 359.9 | 0 | 0 |  | -3.05 | 1.704 | 0.2943% | -0.6993% | 95.6 | 100 | AD
2023-06-29 | ON, EN | 1004.8 | 23.7 | 296.85 | 85.483 | 22.6 | 1004.8 | 22.9 | 1015 | -0.3 | -10.2 | 0.07 | 0.05 | 0.1 | 0 | 0 | 0 | 0.13191 | 0.07 | 0.06 | 0.03 | 0 | 0 | 0 |  | -2.97 | 1.706 | 0.412% | -0.5828% | 100
2023-06-30 | EN, ON | 1001.1 | 23.5 | 296.65 | 85.483 | 22.4 | 1001.1 | 23.1 | 1012 | -0.7 | -10.9 | 0.08 | 0.08 | 0.06 | 0 | 0 | 0 | 0.12806 | 0.08 | 0.08 | 0.07 | 0 | 0 | 0 |  | -2.93 | 1.702 | 0.1766% | -0.8159% | 100
2023-07-03 | AR | 984.8 | 23.4 | 296.55 | 85.483 | 22.3 | 984.8 | 23.2 | 995 | -0.9 | -10.2 | 0.08 | 0.09 | -0.03 | 0 | 0 | 0 | 0.1241 | 0.08 | 0.08 | 0.08 | 0 | 0 | 0 |  | -2.92 | 1.698 | -0.05886% | -1.049% | 100 |  | ok/AC
2023-07-04 | EN, OJ | 991.5 | 22.9 | 296.05 | 85.483 | 22.2 | 991.5 | 23 | 1002 | -0.8 | -10.5 | 0.07 | 0.07 | 0 | -0.1 | 0 | 0 | 0.098995 | 0.07 | 0.07 | 0.07 | 359.9 | 0 | 0 |  | -2.93 | 1.7 | 0.05886% | -0.9324% | 100 |  | ok/AC
2023-07-05 | AR, OJ | 1000.6 | 23.3 | 296.45 | 85.483 | 22.4 | 1000.6 | 23.2 | 1012 | -0.8 | -11.4 | 0.08 | 0.09 | 0.03 | -0.1 | 0 | 0 | 0.1241 | 0.08 | 0.09 | 0.1 | 359.9 | 0 | 0 |  | -2.9 | 1.697 | -0.1177% | -1.107% | 100 |  | AD
2023-07-06 | OJ | 1005.5 | 23.4 | 296.55 | 85.483 | 22.6 | 1005.5 | 23.5 | 1016 | -0.9 | -10.5 | 0.06 | 0.08 | 0.02 | 0 | 0 | 0 | 0.10198 | 0.06 | 0.08 | 0.1 | 0 | 0 | 0 |  | -2.9 | 1.694 | -0.2943% | -1.282% | 96.9 | 100 | AD
2023-07-07 | AJ, OJ | 1007.9 | 23.6 | 296.75 | 85.483 | 22.4 | 1007.9 | 23.1 | 1018 | -0.7 | -10.1 | 0.06 | 0.09 | 0 | -0.1 | 0 | 0 | 0.10817 | 0.06 | 0.09 | 0.09 | 359.9 | 0 | 0 |  | -2.91 | 1.696 | -0.1766% | -1.166% | 100 |  | ML
2023-07-10 | EN, OJ | 1017.8 | 23.4 | 296.55 | 85.483 | 22.8 | 1017.8 | 23.4 | 1028 | -0.6 | -10.2 | 0.05 | 0.07 | 0.03 | 0 | 0 | 0 | 0.091104 | 0.05 | 0.07 | 0.07 | 0 | 0 | 0 |  | -2.93 | 1.695 | -0.2354% | -1.224% | 100 |  | AC
2023-07-11 | OJ, ON | 1012.9 | 23.1 | 296.25 | 85.483 | 22.6 | 1012.9 | 23.3 | 1023 | -0.7 | -10.1 | 0.08 | 0.07 | 0 | -0.1 | 0 | 0 | 0.1063 | 0.08 | 0.07 | 0.01 | 359.9 | 0 | 0 |  | -2.99 | 1.695 | -0.2354% | -1.224% | 98.6 | 100 | ok / YM
2023-07-12 | ON, AR | 1008.1 | 22.9 | 296.05 | 85.483 | 22.4 | 1008.1 | 23 | 1018 | -0.6 | -9.9 | 0.07 | 0.08 | 0 | -0.1 | 0 | 0 | 0.1063 | 0.07 | 0.08 | 0 | 359.9 | 0 | 0 |  | -3.01 | 1.701 | 0.1177% | -0.8741% | 100 |  | ok / YM
2023-07-13 | ON | 1000.6 | 23.1 | 296.25 | 85.483 | 22.4 | 1000.6 | 22.8 | 1006 | -0.4 | -5.4 | 0.09 | 0.07 | 0.05 | 0 | 0 | 0 | 0.1245 | 0.09 | 0.07 | 0.05 | 0 | 0 | 0 |  | -2.95 | 1.702 | 0.1766% | -0.8159% | 100 |  | ME
2023-07-14 | ON | 1003.8 | 23.4 | 296.55 | 85.483 | 22.4 | 1003.8 | 23.1 | 1015 | -0.7 | -11.2 | 0.1 | 0.09 | 0.05 | 0 | 0 | 0 | 0.14353 | 0.1 | 0.1 | 0.04 | 0 | 0 | 0 |  | -2.96 | 1.711 | 0.7063% | -0.2914% | 100 |  | ok/YM
2023-07-17 | ON | 1003.5 | 23.1 | 296.25 | 85.483 | 22.3 | 1004.5 | 23.2 | 1014 | -0.9 | -9.5 | 0.09 | 0.06 | 0.03 | 0.1 | 0 | 0 | 0.11225 | 0.09 | 0.06 | 0.03 | 359.9 | 0 | 0 |  | -2.97 | 1.705 | 0.3531% | -0.641% | 100 |  | ok/YM
2023-07-18 | AC, EN | 999.5 | 23 | 296.15 | 85.483 | 22.8 | 999.6 | 24.4 | 1011 | -1.6 | -11.4 | 0.1 | 0.08 | 0.05 | 0 | 0 | 0 | 0.13748 | 0.1 | 0.09 | 0.05 | 0 | 0 | 0 |  | -2.95 | 1.6948 | -0.249% | -1.237% | 98.2 | 100 | ok/AC
2023-07-19 | ON, EN | 999.9 | 23.5 | 296.65 | 85.483 | 22.2 | 999.9 | 22.7 | 1011 | -0.5 | -11.1 | 0.06 | 0.05 | 0.05 | -0.1 | 0 | 0 | 0.092736 | 0.06 | 0.05 | 0.04 | 359.9 | 0 | 0 |  | -2.96 | 1.707 | 0.4709% | -0.5245% | 100 |  | ok/AC,YM
2023-07-20 | EN, AR | 997 | 23.3 | 296.45 | 85.483 | 22.3 | 997 | 22.7 | 1008 | -0.4 | -11 | 0.09 | 0.05 | 0.03 | -0.1 | 0 | 0 | 0.10724 | 0.09 | 0.05 | 0.11 | 359.9 | 0 | 0 |  | -2.89 | 1.7003 | 0.07652% | -0.9149% | 89.4 | 100 | ok/YM
2023-07-21 | AR | 997.4 | 23.4 | 296.55 | 85.483 | 22.7 | 997.4 | 23 | 1008 | -0.3 | -10.6 | 0.06 | 0.07 | 0.02 | 0 | 0 | 0 | 0.09434 | 0.06 | 0.07 | 0.14 | 0 | 0 | 0 |  | -2.86 | 1.704 | 0.2943% | -0.6993% | 100 |  | ok, YM, MCN
2023-07-24 | AR | 998.4 | 23.4 | 296.55 | 85.483 | 22.6 | 998.4 | 24 | 1009 | -1.4 | -10.6 | 0.1 | 0.06 | 0.05 | 0 | 0 | 0 | 0.12689 | 0.1 | 0.06 | 0.11 | 0 | 0 | 0 |  | -2.89 | 1.696 | -0.1766% | -1.166% | 94.4 |  | ok / YM KW
2023-07-25 | EN | 990.7 | 23.4 | 296.55 | 85.483 | 22.4 | 990.7 | 23 | 1001 | -0.6 | -10.3 | 0.08 | 0.08 | 0.03 | -0.1 | 0 | 0 | 0.11705 | 0.08 | 0.08 | 0.07 | 359.9 | 0 | 0 |  | -2.93 | 1.693 | -0.3531% |  | 96.3 |  | ok / YM KW
2023-07-26 | AR, ON | 993.8 | 23.5 | 296.65 | 85.483 | 22.7 | 993.8 | 23.1 | 1005 | -0.4 | -11.2 | 0.06 | 0.07 | 0.07 | 0 | 0 | 0 | 0.11576 | 0.06 | 0.07 | 0.08 | 0 | 0 | 0 |  | -2.92 | 1.701 | 0.1177% |  | 100 |  | ok/YM
2023-07-27 | ON | 998.7 | 23 | 296.15 | 85.483 | 22.6 | 998.7 | 22.8 | 1009 | -0.2 | -10.3 | 0.04 | 0.07 | 0.05 | -0.1 | 0 | 0 | 0.094868 | 0.04 | 0.07 | 0.04 | 359.9 | 0 | 0 |  | -2.96 | 1.695 | -0.2354% |  | 99.1 |  | ok / YM KW
2023-07-28 | ON | 997.4 | 23.4 | 296.55 | 85.483 | 22.6 | 997.4 | 22.9 | 1008 | -0.3 | -10.6 | 0.04 | 0.06 | 0.07 | -0.1 | 0 | 0 | 0.1005 | 0.04 | 0.06 | 0.06 | 359.9 | 0 | 0 |  | -2.93 | 1.7005 | 0.08829% |  | 100 |  | ok/ YM MCN
IMPORTANT INFORMATIONS:
Before you start working with MatriXX detector #3:
1. Check the orientation of the detector.
2. Check the calibration of the detector (the correct value and the name of calibration factor is written in columns D-E row 23).
3. Apply correction for p & T (for this detector correction for pressure is around 10 hPa).
4. Readout the value of Dcentral (column F) from the Region of interest ROI (1 cm x 1 cm) not from the 'Position cursor'.
5. Record the GI (1mm, 1%) and (2mm, 2%) if the first one does not pass the rate of 90%.

## GTR2_150MeV
ID.nr:
Titel: | [24 empty] | Dokumenttyp:
Protocol for Daily QA: Consistency check of single layer and cubic plans dosimetric parameters | [24 empty] | Strålskyddsinstruktion
Godkänt av: | [23 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [23 empty] | Skapat den:
Reviderat av: | [23 empty] | Reviderat den:
GTR2
Daily single layer parameters
MatriXX PT_3
Parameters and its criterias
Parameter | Value (Gy) | Good: within 1% |  | Warning: within 2% |  | Action level: over 2%
Ref. dose (SPS) | 1.699 | 1.716 | 1.682 | 1.717 | 1.681 | 1.734 | 1.664
Calibration dose | 1.716 | 1.7332 | 1.6988 | 1.7342 | 1.6978 | 1.7513 | 1.6807
Parameter | Value (mm) |  |  | +/- 1.001 mm |  | +/- 2.001 mm
Penumbra (average) | 7.955 | 8.955 | 6.955 | 8.956 | 6.954 | 9.956 | 5.954
FWHM (average) | 104.975 | 105.975 | 103.975 | 105.976 | 103.974 | 106.976 | 102.974
Parameter | Value (Gy) | Good: within +/-0.5% |  | Warning: within +/-1% |  | Action level: over 1%
Symmetry | x | 0.5 | -0.5 | 0.5 | -0.5 | 1 | 1
Parameter | Value (Gy) | Good: within +/-1% |  | Warning: within +/-2% |  | Action level: over 3%
Flatness | x | 1 | -1 | 2 | -2 | 3 | -3
| Output Calibration (150 MeV 10x10 cm2, 1MU per spot) | [19 empty] | Plot
Date | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from Calibration dose (%) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Sign | Checked | Notes | Date | Difference from reference (a.u.) | Medelvärde +/- 10 dygn | Medelvärde +/- 20 dygn | Ref | 1% | -1% | 2% | -2%
20221118 | 1.707 | 0.4709% | -0.5245% | 0.66 | 0.7 | -0.4 | 0.13 | 1.25 | -0.36 | 0.11 | 1.707 | 8.1 | 8.7 | 7.2 | 7.7 | 104.7 | 105.2 | ML | x | GTR2_20221118_OF_01 | 2022-11-18 | 1.0047 | 0.99781 | 0.9965 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221118 | 1.709 | 0.5886% | -0.4079% | 0.69 | 0.61 | -0.41 | 0.16 | 1.29 | -0.35 | 0.1 | 1.709 | 8.2 | 8.6 | 7.4 | 7.7 | 104.7 | 105.1 | ML | x | GTR2_20221118_OF_02 | 2022-11-18 | 1.0059 | 0.99755 | 0.99655 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20221118 | 1.708 | 0.5297% | -0.4662% | 0.66 | 0.7 | -0.41 | 0.09 | 1.22 | -0.37 | 0.09 | 1.708 | 8.1 | 8.7 | 7.4 | 7.6 | 104.7 | 105.1 | ML | x | GTR2_20221118_OF_03 REFERENCE IMAGE | 2022-11-18 | 1.0053 | 0.99701 | 0.9966 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230313 | 1.687 | -0.7063% | -1.69% | 0.54 | 0.62 | -0.3 | 0.13 | 1.13 | -0.28 | 0.08 | 1.687 | 8.1 | 8.6 | 7.4 | 7.6 | 104.8 | 105.2 | AJ | AC | GTR2_20230313_OF_01 | 2023-03-13 | 0.99294 | 0.99668 | 0.99655 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230314 | 1.688 | -0.6474% | -1.632% | 0.6 | 0.55 | -0.37 | 0.13 | 1.06 | -0.32 | 0.09 | 1.688 | 8.2 | 8.6 | 7.1 | 7.8 | 104.8 | 105.2 | SM, EH | YM | GTR2_20230314_OF_01 | 2023-03-14 | 0.99353 | 0.99674 | 0.99654 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230315 | 1.689 | -0.5886% | -1.573% | 0.59 | 0.43 | -0.37 | 0.06 | 1.02 | -0.29 | 0.05 | 1.689 | 8.1 | 8.7 | 7.3 | 7.7 | 104.8 | 105.2 | GF | CVD | GTR2_20230315_OF_01 | 2023-03-15 | 0.99411 | 0.99658 | 0.99637 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230316 | 1.69 | -0.5297% | -1.515% | 0.42 | 0.42 | -0.24 | 0.03 | 0.96 | -0.24 | 0.01 | 1.69 | 8 | 8.8 | 7.6 | 7.4 | 104.8 | 105.2 | EH | ML | GTR2_20230316_OF_01 | 2023-03-16 | 0.9947 | 0.99647 | 0.99622 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230320 | 1.69 | -0.5297% | -1.515% | 0.46 | 0.63 | -0.29 | 0.16 | 1.07 | -0.24 | 0.1 | 1.69 | 8 | 8.8 | 7.7 | 7.5 | 104.8 | 105.2 | AC | AC,ML | GTR2_20230320_OF_01 | 2023-03-20 | 0.9947 | 0.99647 | 0.99614 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230321 | 1.693 | -0.3531% | -1.34% | 0.53 | 0.53 | -0.33 | 0.07 | 1.03 | -0.27 | 0.05 | 1.693 | 8.1 | 8.7 | 7.6 | 7.5 | 104.8 | 105.2 | GF,AE | GB | GTR2_20230321_OF_01 | 2023-03-21 | 0.99647 | 0.99634 | 0.99607 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230322 | 1.693 | -0.3531% | -1.34% | 0.44 | 0.56 | -0.29 | 0.09 | 1.08 | -0.27 | 0.07 | 1.693 | 8 | 8.8 | 7.4 | 7.7 | 104.8 | 105.2 | SF | CVD | GTR2_20230322_OF_01 | 2023-03-22 | 0.99647 | 0.9965 | 0.99596 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230323 | 1.694 | -0.2943% | -1.282% | 0.56 | 0.61 | -0.25 | 0.09 | 1.04 | -0.26 | 0.05 | 1.694 | 8.5 | 8.3 | 7.1 | 7.9 | 104.8 | 105.2 | SM, OJ | ML | GTR2_20230323_OF_01 | 2023-03-23 | 0.99706 | 0.99655 | 0.99598 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230324 | 1.69 | -0.5297% | -1.515% | 0.54 | 0.52 | -0.31 | 0.09 | 1.02 | -0.28 | 0.05 | 1.69 | 7.6 | 9 | 7.1 | 7.9 | 104.8 | 105.2 | SM, EH | KW | GTR2_20230324_OF_01 | 2023-03-24 | 0.9947 | 0.99622 | 0.99592 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230327 | 1.683 | -0.9417% | -1.923% | 0.39 | 0.57 | -0.24 | 0.07 | 0.97 | -0.23 | 0.03 | 1.683 | 8.1 | 8.7 | 6.8 | 7.9 | 104.8 | 105.2 | AJ, GK | AC | GTR2_20230327_OF_01 | 2023-03-27 | 0.99058 | 0.99571 | 0.99588 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230328 | 1.686 | -0.7652% | -1.748% | 0.62 | 0.59 | -0.31 | 0.07 | 1.2 | -0.28 | 0.08 | 1.686 | 8.4 | 8.4 | 7.1 | 7.9 | 104.8 | 105.2 | AE | YM | GTR2_20230328_OF_01 | 2023-03-28 | 0.99235 | 0.99529 | 0.99584 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230329 | 1.695 | -0.2354% | -1.224% | 0.46 | 0.57 | -0.24 | 0.1 | 0.97 | -0.22 | 0.09 | 1.695 | 8.4 | 8.4 | 7.4 | 7.7 | 104.8 | 105.2 | GK,EH,ON | YM ON | GTR2_20230329_OF_01 | 2023-03-29 | 0.99765 | 0.99526 | 0.99605 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230330 | 1.689 | -0.5886% | -1.573% | 0.41 | 0.57 | -0.2 | 0.12 | 0.98 | -0.21 | 0.1 | 1.689 | 8.3 | 8.5 | 7.4 | 7.6 | 104.8 | 105.2 | JR,ON | ON | GTR2_20230330_OF_01 | 2023-03-30 | 0.99411 | 0.99521 | 0.99623 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230331 | 1.69 | -0.5297% | -1.515% | 0.55 | 0.63 | -0.31 | 0.12 | 1.15 | -0.32 | 0.09 | 1.69 | 8 | 8.8 | 7.5 | 7.6 | 104.8 | 105.2 | AR, SM | GB | GTR2_20230331_OF_01 | 2023-03-31 | 0.9947 | 0.99521 | 0.99622 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230403 | 1.693 | -0.3531% | -1.34% | 0.6 | 0.57 | -0.31 | 0.13 | 1.18 | -0.3 | 0.09 | 1.693 | 8.4 | 8.4 | 7.5 | 7.5 | 104.8 | 105.2 | SM, AR | YM | GTR2_20230403_OF_01 | 2023-04-03 | 0.99647 | 0.99518 | 0.99623 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230404 | 1.689 | -0.5886% | -1.573% | 0.57 | 0.55 | -0.36 | 0.11 | 1.11 | -0.32 | 0.07 | 1.689 | 8.1 | 8.8 | 7.4 | 7.7 | 104.8 | 105.2 | ON | GB | GTR2_20230404_OF_01 | 2023-04-04 | 0.99411 | 0.9951 | 0.99622 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230405 | 1.698 | -0.05886% | -1.049% | 0.48 | 0.64 | -0.23 | 0.13 | 1.06 | -0.24 | 0.09 | 1.698 | 8.4 | 8.4 | 7.4 | 7.7 | 104.8 | 105.2 | AE | GB | GTR2_20230405_OF_01 | 2023-04-05 | 0.99941 | 0.9951 | 0.99624 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230406 | 1.695 | -0.2354% | -1.224% | 0.57 | 0.57 | -0.31 | 0.12 | 1.11 | -0.28 | 0.08 | 1.695 | 8.5 | 8.4 | 7.2 | 7.9 | 104.8 | 105.2 | EH | CVD | GTR2_20230406_OF_01 | 2023-04-06 | 0.99765 | 0.99498 | 0.99629 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230407 | 1.695 | -0.2354% | -1.224% | 0.52 | 0.67 | -0.28 | 0.18 | 1.03 | -0.29 | 0.13 | 1.695 | 8 | 8.8 | 7.4 | 7.8 | 104.8 | 105.2 | SM, JR | CVD | GTR2_20230407_OF_01 | 2023-04-07 | 0.99765 | 0.99487 | 0.99632 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230411 | 1.691 | -0.4709% | -1.457% | 0.55 | 0.56 | -0.35 | 0.07 | 1.08 | -0.28 | 0.07 | 1.691 | 8.2 | 8.6 | 7.2 | 8 | 104.8 | 105.2 | GF, EH, ON | GB | GTR2_20230411_OF_01 | 2023-04-11 | 0.99529 | 0.99487 | 0.99612 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230412 | 1.693 | -0.3531% | -1.34% | 0.5 | 0.63 | -0.31 | 0.11 | 1.09 | -0.29 | 0.07 | 1.693 | 7.8 | 8.9 | 7.4 | 7.8 | 104.8 | 105.2 | JR,GF | Ym | GTR2_20230412_OF_01 | 2023-04-12 | 0.99647 | 0.99546 | 0.99584 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230413 | 1.686 | -0.7652% | -1.748% | 0.57 | 0.65 | -0.34 | 0.17 | 1.09 | -0.3 | 0.12 | 1.686 | 7.8 | 8.9 | 7.6 | 7.5 | 104.8 | 105.2 | SM, AE | YM | GTR2_20230413_OF_01 | 2023-04-13 | 0.99235 | 0.99594 | 0.99557 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230414 | 1.686 | -0.7652% | -1.748% | 0.5 | 0.59 | -0.31 | 0.14 | 1.05 | -0.27 | 0.11 | 1.686 | 8.2 | 8.6 | 7 | 8 | 104.8 | 105.2 | OJ, SM | GB | GTR2_20230414_OF_01 | 2023-04-14 | 0.99235 | 0.99585 | 0.99567 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230417 | 1.689 | -0.5886% | -1.573% | 0.49 | 0.64 | -0.29 | 0.14 | 0.94 | -0.26 | 0.07 | 1.689 | 7.4 | 9.2 | 7.3 | 7.6 | 104.7 | 105.2 | AJ, OJ | GB | GTR2_20230417_OF_01 | 2023-04-17 | 0.99411 | 0.99596 | 0.99578 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230418 | 1.689 | -0.5886% | -1.573% | 0.53 | 0.55 | -0.32 | 0.12 | 1.05 | -0.3 | 0.08 | 1.689 | 7.5 | 9.1 | 7.2 | 7.8 | 104.8 | 105.2 | GK | GB | GTR2_20230418_OF_01 | 2023-04-18 | 0.99411 | 0.99602 | 0.99589 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230419 | 1.687 | -0.7063% | -1.69% | 0.59 | 0.57 | -0.31 | 0.11 | 1.2 | -0.28 | 0.1 | 1.687 | 8.4 | 8.4 | 7.4 | 7.6 | 104.8 | 105.2 | OJ, AR | CVD | GTR2_20230419_OF_01 | 2023-04-19 | 0.99294 | 0.99605 | 0.99589 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230420 | 1.693 | -0.3531% | -1.34% | 0.57 | 0.56 | -0.34 | 0.1 | 1.12 | -0.33 | 0.07 | 1.693 | 7.9 | 8.8 | 7.5 | 7.7 | 104.8 | 105.2 | ON, EH | CVD | GTR2_20230420_OF_01 | 2023-04-20 | 0.99647 | 0.99624 | 0.99594 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230421 | 1.689 | -0.5886% | -1.573% | 0.45 | 0.65 | -0.29 | 0.09 | 1.16 | -0.29 | 0.06 | 1.689 | 7.7 | 9 | 7.5 | 7.6 | 104.7 | 105.2 | SM, ON | CVD | GTR2_20230421_OF_01 | 2023-04-21 | 0.99411 | 0.99616 | 0.99606 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230424 | 1.69 | -0.5297% | -1.515% | 0.57 | 0.5 | -0.32 | 0.1 | 1.14 | -0.31 | 0.09 | 1.69 | 8.3 | 8.5 | 7.3 | 7.7 | 104.8 | 105.2 | GK,AE | YM | GTR2_20230424_OF_01 | 2023-04-24 | 0.9947 | 0.99609 | 0.99614 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230425 | 1.69 | -0.5297% | -1.515% | 0.49 | 0.69 | -0.3 | 0.15 | 1 | -0.25 | 0.08 | 1.69 | 7.8 | 8.9 | 7.6 | 7.3 | 104.8 | 105.2 | AE | CVD | GTR2_20230425_OF_01 | 2023-04-25 | 0.9947 | 0.99594 | 0.9961 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230426 | 1.704 | 0.2943% | -0.6993% | 0.56 | 0.58 | -0.34 | 0.11 | 1.07 | -0.31 | 0.06 | 1.704 | 7.9 | 8.8 | 6.9 | 7.9 | 104.8 | 105.2 | AR, GK | YM | GTR2_20230426_OF_01 | 2023-04-26 | 1.0029 | 0.99591 | 0.99613 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230427 | 1.703 | 0.2354% | -0.7576% | 0.49 | 0.63 | -0.3 | 0.12 | 1.12 | -0.28 | 0.08 | 1.703 | 8.1 | 8.7 | 7.2 | 7.8 | 104.8 | 105.2 | GK,OJ | CVD | GTR2_20230427_OF_01 | 2023-04-27 | 1.0024 | 0.99594 | 0.99626 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230428 | 1.692 | -0.412% | -1.399% | 0.52 | 0.69 | -0.23 | 0.03 | 1.03 | -0.19 | 0.07 | 1.692 | 8.6 | 8.2 | 6.9 | 8 | 104.8 | 105.2 | GK,JR | AC | GTR2_20230428_OF_01 | 2023-04-28 | 0.99588 | 0.9962 | 0.99636 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230502 | 1.693 | -0.3531% | -1.34% | 0.48 | 0.71 | -0.29 | 0.07 | 1.1 | -0.27 | 0.07 | 1.693 | 8.1 | 8.7 | 7.3 | 7.7 | 104.8 | 105.2 | AE | CVD | GTR2_20230502_OF_01 | 2023-05-02 | 0.99647 | 0.99653 | 0.99627 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230503 | 1.692 | -0.412% | -1.399% | 0.5 | 0.55 | -0.32 | 0.05 | 1.04 | -0.28 | 0.05 | 1.692 | 8 | 8.8 | 7.4 | 7.6 | 104.8 | 105.2 | ON | CVD | GTR2_20230503_OF_01 | 2023-05-03 | 0.99588 | 0.99656 | 0.99629 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230504 | 1.694 | -0.2943% | -1.282% | 0.54 | 0.52 | -0.35 | 0.1 | 1.07 | -0.31 | 0.06 | 1.694 | 7.6 | 9 | 7.2 | 7.8 | 104.7 | 105.2 | JR | YM | GTR2_20230504_OF_01 | 2023-05-04 | 0.99706 | 0.99667 | 0.99626 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230504 | 1.696 | -0.1766% | -1.166% | 0.63 | 0.61 | -0.37 | 0.09 | 1.18 | -0.31 | 0.06 | 1.696 | 8 | 8.8 | 7.5 | 7.6 | 104.8 | 105.2 | ON, EF | ML | GTR2_20230504_OF_02 | 2023-05-04 | 0.99823 | 0.99709 | 0.99621 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230505 | 1.695 | -0.2354% | -1.224% | 0.58 | 0.59 | -0.31 | 0.14 | 1.08 | -0.27 | 0.07 | 1.695 | 8.4 | 8.5 | 7.5 | 7.6 | 104.8 | 105.2 | SM, GK | AC | GTR2_20230505_OF_01 | 2023-05-05 | 0.99765 | 0.99726 | 0.99623 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230508 | 1.69 | -0.5297% | -1.515% | 0.65 | 0.64 | -0.37 | 0.08 | 1.19 | -0.35 | 0.05 | 1.69 | 8.2 | 8.6 | 7.2 | 7.9 | 104.8 | 105.2 | OJ, SF | AC | GTR2_20230508_OF_01 | 2023-05-08 | 0.9947 | 0.99738 | 0.99615 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230509 | 1.69 | -0.5297% | -1.515% | 0.43 | 0.58 | -0.28 | 0.11 | 0.99 | -0.24 | 0.09 | 1.69 | 8.1 | 8.6 | 6.7 | 7.9 | 104.8 | 105.2 | GF | GB | GTR2_20230509_OF_01 | 2023-05-09 | 0.9947 | 0.99744 | 0.99614 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230510 | 1.694 | -0.2943% | -1.282% | 0.47 | 0.68 | -0.28 | 0.16 | 1.03 | -0.28 | 0.08 | 1.694 | 8 | 8.8 | 7.3 | 7.6 | 104.8 | 105.2 | ON | GB | GTR2_20230510_OF_01 | 2023-05-10 | 0.99706 | 0.99715 | 0.9963 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230511 | 1.695 | -0.2354% | -1.224% | 0.6 | 0.59 | -0.33 | 0.08 | 1.09 | -0.3 | 0.05 | 1.695 | 8.4 | 8.4 | 7.5 | 7.5 | 104.8 | 105.2 | AR | CVD | GTR2_20230511_OF_01 | 2023-05-11 | 0.99765 | 0.99672 | 0.9965 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230512 | 1.697 | -0.1177% | -1.107% | 0.49 | 0.57 | -0.31 | 0.07 | 1.09 | -0.28 | 0.06 | 1.697 | 8 | 8.8 | 7.5 | 7.4 | 104.8 | 105.2 | SF, JR | EAL | GTR2_20230512_OF_01 | 2023-05-12 | 0.99882 | 0.99665 | 0.99676 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230515 | 1.69 | -0.5297% | -1.515% | 0.56 | 0.6 | -0.33 | 0.08 | 1.17 | -0.3 | 0.04 | 1.69 | 7.7 | 9 | 7.4 | 7.7 | 104.8 | 105.2 | OJ, ON, EN | GB | GTR2_20230515_OF_01 | 2023-05-15 | 0.9947 | 0.9965 | 0.99704 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230516 | 1.693 | -0.3531% | -1.34% | 0.47 | 0.64 | -0.29 | 0.09 | 1.08 | -0.28 | 0.08 | 1.693 | 7.9 | 8.8 | 7.4 | 7.7 | 104.8 | 105.2 | EN,AJ,AR | ON, YM | GTR2_20230516_OF_01 | 2023-05-16 | 0.99647 | 0.99644 | 0.99719 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230517 | 1.701 | 0.1177% | -0.8741% | 0.51 | 0.54 | -0.33 | 0.08 | 1.04 | -0.31 | 0.06 | 1.701 | 8.1 | 8.7 | 7.4 | 7.6 | 104.8 | 105.2 | AE | YM ON | GTR2_20230517_OF_01 | 2023-05-17 | 1.0012 | 0.99631 | 0.99742 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230519 | 1.699 | 0% | -0.9907% | 0.56 | 0.64 | -0.33 | 0.12 | 1.12 | -0.3 | 0.07 | 1.699 | 7.9 | 8.8 | 7.4 | 7.5 | 104.8 | 105.2 | ON | AD | GTR2_20230519_OF_01 | 2023-05-19 | 1 | 0.99641 | 0.99763 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230522 | 1.691 | -0.4709% | -1.457% | 0.58 | 0.53 | -0.31 | 0.08 | 1.07 | -0.28 | 0.06 | 1.691 | 8.4 | 8.4 | 7.7 | 7.5 | 104.8 | 105.2 | AE | GB | GTR2_20230522_OF_01 | 2023-05-22 | 0.99529 | 0.99622 | 0.99765 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230523 | 1.692 | -0.412% | -1.399% | 0.52 | 0.56 | -0.31 | 0.1 | 1.11 | -0.28 | 0.05 | 1.692 | 8.5 | 8.3 | 7 | 8 | 104.8 | 105.2 | GF, ON | CVD | GTR2_20230523_OF_01 | 2023-05-23 | 0.99588 | 0.99626 | 0.99813 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230524 | 1.692 | -0.412% | -1.399% | 0.62 | 0.7 | -0.36 | 0.13 | 1.16 | -0.31 | 0.08 | 1.692 | 8 | 8.8 | 7.3 | 7.8 | 104.8 | 105.1 | AE | AD | GTR2_20230524_OF_01 | 2023-05-24 | 0.99588 | 0.99662 | 0.9993 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230525 | 1.689 | -0.5886% | -1.573% | 0.5 | 0.66 | -0.3 | 0.09 | 1.14 | -0.29 | 0.08 | 1.689 | 7.8 | 8.9 | 7.6 | 7.5 | 104.8 | 105.2 | AR, SM | GB | GTR2_20230525_OF_01 | 2023-05-25 | 0.99411 | 0.99735 | 0.99931 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230526 | 1.69 | -0.5297% | -1.515% | 0.55 | 0.58 | -0.25 | 0.15 | 1.13 | -0.27 | 0.07 | 1.69 | 8.4 | 8.4 | 7.4 | 7.7 | 104.8 | 105.1 | OJ, EN | AC | GTR2_20230526_OF_01 | 2023-05-26 | 0.9947 | 0.99765 | 0.99926 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230529 | 1.688 | -0.6474% | -1.632% | 0.54 | 0.66 | -0.24 | 0.07 | 1.04 | -0.21 | 0.08 | 1.688 | 8.6 | 8.2 | 7.3 | 7.8 | 104.8 | 105.1 | OJ | CVD | GTR2_20230529_OF_01 | 2023-05-29 | 0.99353 | 0.9977 | 0.99932 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230530 | 1.69 | -0.5297% | -1.515% | 0.56 | 0.51 | -0.33 | 0.1 | 1.1 | -0.3 | 0.08 | 1.69 | 8.1 | 8.7 | 7.5 | 7.6 | 104.8 | 105.2 | AE | YM | GTR2_20230530_OF_01 | 2023-05-30 | 0.9947 | 0.99812 | 0.99936 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230531 | 1.69 | -0.5297% | -1.515% | 0.46 | 0.71 | -0.2 | 0.02 | 1.05 | -0.21 | 0.05 | 1.69 | 8.7 | 8 | 7.1 | 8 | 104.8 | 105.1 | JR | YM | GTR2_20230531_OF_01 | 2023-05-31 | 0.9947 | 0.99835 | 0.99942 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230601 | 1.699 | 0% | -0.9907% | 0.54 | 0.66 | -0.26 | 0.16 | 1.07 | -0.29 | 0.1 | 1.699 | 8.3 | 8.5 | 7.4 | 7.7 | 104.8 | 105.1 | EF | AD | GTR2_20230601_OF_01 | 2023-06-01 | 1 | 0.99815 | 0.99946 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230602 | 1.689 | -0.5886% | -1.573% | 0.39 | 0.66 | -0.16 | 0.01 | 0.97 | -0.16 | 0.03 | 1.689 | 8.6 | 8.2 | 7.1 | 7.9 | 104.8 | 105.2 | AJ | AC | GTR2_20230602_OF_01 | 2023-06-02 | 0.99411 | 0.99879 | 0.99948 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230605 | 1.694 | -0.2943% | -1.282% | 0.73 | 0.52 | -0.42 | 0.03 | 1.17 | -0.37 | 0.02 | 1.694 | 8.1 | 8.7 | 7.7 | 7.6 | 104.8 | 105.1 | EF | AC/KW | GTR2_20230605_OF_01 | 2023-06-05 | 0.99706 | 1 | 0.99951 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230607 | 1.702 | 0.1766% | -0.8159% | 0.47 | 0.68 | -0.3 | 0.11 | 1.16 | -0.25 | 0.07 | 1.702 | 8.2 | 8.7 | 7.6 | 7.6 | 104.7 | 105.1 | ON | YM | GTR2_20230607_OF_01 | 2023-06-07 | 1.0018 | 1.0001 | 0.99951 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230608 | 1.706 | 0.412% | -0.5828% | 0.53 | 0.47 | -0.31 | 0.09 | 1.07 | -0.29 | 0.05 | 1.706 | 8.5 | 8.3 | 7.1 | 7.9 | 104.8 | 105.1 | EN | GB | GTR2_20230608_OF_01 | 2023-06-08 | 1.0041 | 1.0002 | 0.9996 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230609 | 1.703 | 0.2354% | -0.7576% | 0.58 | 0.52 | -0.29 | 0.05 | 1.1 | -0.31 | 0.03 | 1.703 | 8.2 | 8.5 | 6.9 | 8 | 104.8 | 105.1 | AR | AC | GTR2_20230609_OF_01 | 2023-06-09 | 1.0024 | 1.0002 | 0.99969 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230612 | 1.705 | 0.3531% | -0.641% | 0.62 | 0.6 | -0.3 | 0.16 | 1.11 | -0.27 | 0.06 | 1.705 | 8.4 | 8.5 | 7.1 | 7.9 | 104.8 | 105.2 | AC | AC | GTR2_20230612_OF_01 | 2023-06-12 | 1.0035 | 1.0004 | 0.99974 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230613 | 1.699 | 0% | -0.9907% | 0.6 | 0.6 | -0.36 | 0.12 | 1.1 | -0.34 | 0.06 | 1.699 | 8.2 | 8.6 | 7.3 | 7.8 | 104.8 | 105.1 | OJ | GB | GTR2_20230613_OF_01 | 2023-06-13 | 1 | 1.0005 | 0.99978 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230614 | 1.704 | 0.2943% | -0.6993% | 0.53 | 0.63 | -0.27 | 0.1 | 1.1 | -0.26 | 0.04 | 1.704 | 8.3 | 8.5 | 7.3 | 7.7 | 104.7 | 105.1 | EF | YM | GTR2_20230614_OF_01 | 2023-06-14 | 1.0029 | 1.0007 | 0.9998 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230615 | 1.701 | 0.1177% | -0.8741% | 0.6 | 0.46 | -0.32 | 0.05 | 1.05 | -0.32 | 0.04 | 1.701 | 8.3 | 8.5 | 7.1 | 8 | 104.8 | 105.1 | EF | YM | GTR2_20230615_OF_01 | 2023-06-15 | 1.0012 | 1.0008 | 0.9999 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230616 | 1.694 | -0.2943% | -1.282% | 0.47 | 0.66 | -0.29 | 0.12 | 1.11 | -0.29 | 0.09 | 1.694 | 8.2 | 8.6 | 7.6 | 7.5 | 104.8 | 105.2 | EN | AC | GTR2_20230616_OF_01 | 2023-06-16 | 0.99706 | 1.001 | 0.99996 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230619 | 1.721 | 1.295% | 0.2914% | 0.47 | 0.6 | -0.25 | 0.05 | 1.07 | -0.24 | 0.02 | 1.721 | 8.4 | 8.4 | 7.7 | 7.4 | 104.8 | 105.1 | EN | ML | GTR2_20230619_OF_01 | 2023-06-19 | 1.0129 | 1.001 | 0.99994 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230620 | 1.697 | -0.1177% | -1.107% | 0.43 | 0.61 | -0.18 | -0.02 | 0.96 | -0.17 | 0.01 | 1.697 | 8.9 | 7.6 | 7.7 | 7.4 | 104.7 | 105.1 | GK | GB | GTR2_20230620_OF_01 | 2023-06-20 | 0.99882 | 1.0012 | 0.99994 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230621 | 1.699 | 0% | -0.9907% | 0.56 | 0.63 | -0.32 | 0.03 | 1.05 | -0.26 | 0.04 | 1.699 | 8.3 | 8.5 | 7.5 | 7.7 | 104.8 | 105.1 | OJ | CVD | GTR2_20230621_OF_01 | 2023-06-21 | 1 | 1.0013 | 1 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230622 | 1.703 | 0.2354% | -0.7576% | 0.43 | 0.55 | -0.24 | 0.02 | 1.05 | -0.26 | 0 | 1.703 | 8.3 | 8.4 | 7.6 | 7.4 | 104.8 | 105.1 | ON | ML | GTR2_20230622_OF_01 | 2023-06-22 | 1.0024 | 1.0013 | 1.0001 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230626 | 1.704 | 0.2943% | -0.6993% | 0.43 | 0.49 | -0.26 | 0.05 | 0.91 | -0.2 | 0.02 | 1.704 | 7.8 | 8.9 | 7.3 | 7.7 | 104.7 | 105.1 | AR | ML | GTR2_20230626_OF_01 | 2023-06-26 | 1.0029 | 1.0012 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230627 | 1.707 | 0.4709% | -0.5245% | 0.56 | 0.54 | -0.35 | 0.06 | 1.07 | -0.29 | 0.02 | 1.707 | 8 | 8.8 | 7.4 | 7.7 | 104.7 | 105.1 | EN | AC | GTR2_20230627_OF_01 | 2023-06-27 | 1.0047 | 1.0011 | 1.0002 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230628 | 1.704 | 0.2943% | -0.6993% | 0.45 | 0.67 | -0.17 | 0.03 | 1.02 | -0.22 | 0.03 | 1.704 | 8.6 | 8.1 | 7.3 | 7.7 | 104.8 | 105.2 | AR | AD | GTR2_20230628_OF_01 | 2023-06-28 | 1.0029 | 1.0011 | 1.0004 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230629 | 1.706 | 0.412% | -0.5828% | 0.58 | 0.46 | -0.37 | 0.02 | 1 | -0.32 | 0.03 | 1.706 | 7.8 | 8.8 | 7.2 | 7.9 | 104.7 | 105.1 | ON | ML | GTR2_20230629_OF_01 | 2023-06-29 | 1.0041 | 1.0011 | 1.0005 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230630 | 1.702 | 0.1766% | -0.8159% | 0.57 | 0.59 | -0.35 | 0.05 | 1.12 | -0.29 | 0.01 | 1.702 | 8 | 8.7 | 7.5 | 7.5 | 104.7 | 105.1 | EN | ML | GTR2_20230630_OF_01 | 2023-06-30 | 1.0018 | 1.001 | 1.0007 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230703 | 1.698 | -0.05886% | -1.049% | 0.59 | 0.5 | -0.35 | -0.01 | 1.04 | -0.3 | -0.02 | 1.698 | 7.9 | 8.8 | 7.4 | 7.6 | 104.7 | 105.2 | AR | AC | GTR2_20230703_OF_01 | 2023-07-03 | 0.99941 | 1.001 | 1.0008 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230704 | 1.7 | 0.05886% | -0.9324% | 0.53 | 0.57 | -0.32 | 0.06 | 1.03 | -0.29 | 0.03 | 1.7 | 8.2 | 8.6 | 7.3 | 7.7 | 104.8 | 105.2 | EN | AC | GTR2_20230704_OF_01 | 2023-07-04 | 1.0006 | 1.0012 | 1.001 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230705 | 1.697 | -0.1177% | -1.107% | 0.42 | 0.57 | -0.27 | 0.04 | 0.97 | -0.25 | 0.03 | 1.697 | 7.7 | 9 | 7.3 | 7.6 | 104.7 | 105.2 | AR | AD | GTR2_20230705_OF_01 | 2023-07-05 | 0.99882 | 1.0008 | 1.001 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230706 | 1.694 | -0.2943% | -1.282% | 0.54 | 0.57 | -0.3 | 0.1 | 1.05 | -0.28 | 0.05 | 1.694 | 8.4 | 8.4 | 7.4 | 7.6 | 104.8 | 105.1 | OJ | AD | GTR2_20230706_OF_01 | 2023-07-06 | 0.99706 | 1.0008 | 1.0012 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230707 | 1.696 | -0.1766% | -1.166% | 0.52 | 0.59 | -0.31 | 0.11 | 1.08 | -0.32 | 0.07 | 1.696 | 8 | 8.8 | 7.4 | 7.6 | 104.8 | 105.2 | AJ, OJ | ML | GTR2_20230707_OF_01 | 2023-07-07 | 0.99823 | 1.0009 | 1.0013 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230710 | 1.695 | -0.2354% | -1.224% | 0.58 | 0.59 | -0.34 | 0.13 | 1.02 | -0.31 | 0.09 | 1.695 | 8.2 | 8.7 | 7.6 | 7.5 | 104.8 | 105.2 | EN | AC | GTR2_20230710_OF_01 | 2023-07-10 | 0.99765 | 1.0008 | 1.0013 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230711 | 1.695 | -0.2354% | -1.224% | 0.49 | 0.55 | -0.31 | 0.09 | 1.04 | -0.3 | 0.05 | 1.695 | 8 | 8.7 | 7.3 | 7.6 | 104.8 | 105.2 | OJ | YM | GTR2_20230711_OF_01 | 2023-07-11 | 0.99765 | 1.0007 | 1.0012 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230712 | 1.701 | 0.1177% | -0.8741% | 0.47 | 0.54 | -0.3 | 0.1 | 1.06 | -0.28 | 0.05 | 1.701 | 8 | 8.7 | 7.5 | 7.4 | 104.8 | 105.2 | ON | YM | GTR2_20230712_OF_01 | 2023-07-12 | 1.0012 | 1.0005 | 1.0011 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230713 | 1.702 | 0.1766% | -0.8159% | 0.47 | 0.58 | -0.28 | 0.09 | 1 | -0.25 | 0.06 | 1.702 | 7.8 | 8.9 | 7.5 | 7.5 | 104.7 | 105.2 | ON | ME | GTR2_20230713_OF_01 | 2023-07-13 | 1.0018 | 1.0004 | 1.0011 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230714 | 1.711 | 0.7063% | -0.2914% | 0.46 | 0.6 | -0.26 | 0.07 | 0.95 | -0.25 | 0.04 | 1.711 | 8 | 8.7 | 7.1 | 7.6 | 104.7 | 105.2 | ON | YM | GTR2_20230714_OF_01 | 2023-07-14 | 1.0071 | 1.0003 | 1.0011 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230717 | 1.705 | 0.3531% | -0.641% | 0.48 | 0.53 | -0.29 | 0.03 | 1.09 | -0.28 | 0.05 | 1.705 | 7.7 | 8.9 | 7.6 | 7.2 | 104.8 | 105.2 | ON | YM | GTR2_20230717_OF_01 | 2023-07-17 | 1.0035 | 1.0002 | 1.001 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230718 | 1.695 | -0.2354% | -1.224% | 0.53 | 0.69 | -0.3 | 0.04 | 1.13 | -0.29 | 0.04 | 1.695 | 7.9 | 8.8 | 7.2 | 7.8 | 104.7 | 105.1 | AC | AC | GTR2_20230718_OF_01 | 2023-07-18 | 0.99765 | 1.0002 | 1.001 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230719 | 1.707 | 0.4709% | -0.5245% | 0.44 | 0.52 | -0.25 | 0.07 | 1 | -0.23 | 0.06 | 1.707 | 8.2 | 8.5 | 7.2 | 7.7 | 104.7 | 105.2 | ON | AC,YM | GTR2_20230719_OF_01 | 2023-07-19 | 1.0047 | 1.0002 | 1.0012 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230720 | 1.7 | 0.05886% | -0.9324% | 0.37 | 0.64 | -0.14 | 0.02 | 1.03 | -0.18 | 0.06 | 1.7 | 8.8 | 7.8 | 7.5 | 7.5 | 104.7 | 105.2 | EN | YM | GTR2_20230720_OF_01 | 2023-07-20 | 1.0006 | 1.0003 | 1.0008 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230721 | 1.704 | 0.2943% | -0.6993% | 0.45 | 0.5 | -0.28 | 0.1 | 1 | -0.28 | 0.09 | 1.704 | 8 | 8.7 | 7.4 | 7.6 | 104.7 | 105.1 | AR | YM, NIC | GTR2_20230721_OF_01 | 2023-07-21 | 1.0029 | 1.0005 | 1.0008 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230724 | 1.696 | -0.1766% | -1.166% | 0.39 | 0.7 | -0.22 | 0.05 | 1.16 | -0.25 | 0.06 | 1.696 | 7.9 | 8.9 | 7.6 | 7.4 | 104.7 | 105.2 | EN | YM KW | GTR2_20230724_OF_01 | 2023-07-24 | 0.99823 | 1.0006 | 1.0009 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230725 | 1.693 | -0.3531% | -1.34% | 0.58 | 0.56 | -0.31 | 0.1 | 1.1 | -0.32 | 0.1 | 1.693 | 7.9 | 8.9 | 7.6 | 7.5 | 104.8 | 105.2 | EN | YM KW | GTR2_20230725_OF_01 | 2023-07-25 | 0.99647 | 1.0008 | 1.0008 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230726 | 1.701 | 0.1177% | -0.8741% | 0.48 | 0.58 | -0.32 | 0.06 | 1.04 | -0.29 | 0.07 | 1.701 | 8.1 | 8.7 | 7.1 | 7.7 | 104.8 | 105.2 | AR | YM | GTR2_20230726_OF_01 | 2023-07-26 | 1.0012 | 1.0011 | 1.0007 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230727 | 1.695 | -0.2354% | -1.224% | 0.46 | 0.65 | -0.29 | 0.07 | 1.08 | -0.28 | 0.04 | 1.695 | 8.2 | 8.7 | 7.2 | 7.7 | 104.8 | 105.2 | ON | YM KW | GTR2_20230727_OF_01 | 2023-07-27 | 0.99765 | 1.0011 | 1.0005 | 1 | 1.01 | 0.99 | 1.02 | 0.98
20230728 | 1.701 | 0.1177% | -0.8741% | 0.54 | 0.61 | -0.32 | 0.11 | 1.1 | -0.3 | 0.08 | 1.701 | 8 | 8.8 | 7.4 | 7.5 | 104.7 | 105.2 | ON | YM MCN | GTR2_20230728_OF_01 | 2023-07-28 | 1.0012 | 1.001 | 1.0004 | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
[21 empty] | #VALUE! |  |  |  | 1 | 1.01 | 0.99 | 1.02 | 0.98
Formulas:
Date = DATE(LEFT([Parameter],4), MID([Parameter],5,2), RIGHT([Parameter],2))

## GTR2_Cubes
ID.nr:
Titel: | [52 empty] | Dokumenttyp:
Protocol for Daily QA: Consistency check of single layer and cubic plans dosimetric parameters | [52 empty] | Strålskyddsinstruktion
Godkänt av: | [51 empty] | Godkänt den:
Kategori: | Strålskydd
Skapat av: | [51 empty] | Skapat den:
Reviderat av: | [51 empty] | Reviderat den:
GTR2
Daily cubes parameters
MatriXX PT_3
Set-up: |  | MatriXX (MatrixPT_3 - position adjusted using X-rays (chambers at iso))
Detector: |  | IBA MatrixPT_3
Phantom: |  | SW
MyQA Project: |  | GTR2_DailyQA_Matrixx
Patient in AdaptDelivery: |  | G2_Daily_MatriXX Skandion
Plan ID: |  | Cube_R15_SOBP10; Cube_R31_SOBP10
Parameters and its criterias
Plan ID | Parameter | Value (Gy) | Good: within 2% |  | Warning: over 2% |  | Action level: over 3%
R15M10 | Ref. dose (Roos IC SW)* | 1.987 | 2.0267 | 1.9473 | 2.0277 | 1.9463 | 2.0476 | 1.9264
R31M10 | Ref. dose (Roos IC SW)* | 2.007 | 2.0471 | 1.9669 | 2.0481 | 1.9659 | 2.0682 | 1.9458
R15M10 | Dose from TPS | 2.004 | 2.0441 | 1.9639 | 2.0451 | 1.9629 | 2.0651 | 1.9429
R31M10 | Dose from TPS | 2.026 | 2.0441 | 1.9639 | 2.0451 | 1.9629 | 2.0651 | 1.9429
Plan ID | Parameter | Value (mm) | Good |  | +/- 1.001 mm |  | +/- 2.001 mm
R15M10 | Penumbra (average) | 8.6125 | 9.6125 | 7.6125 | 9.6135 | 7.6115 | 10.6135 | 6.6115
R31M10 | Penumbra (average) | 11.6175 | 12.6175 | 10.6175 | 12.6185 | 10.6165 | 13.6185 | 9.6165
Plan ID | Parameter | Value (mm) | Good |  | +/- 1.001 mm |  | +/- 2.001 mm
R15M10 | FWHM (average) | 119.945 | 120.945 | 118.945 | 120.946 | 118.944 | 121.946 | 117.944
R31M10 | FWHM (average) | 119.6 | 120.6 | 118.6 | 120.601 | 118.599 | 121.601 | 117.599
Plan ID | Parameter |  | Good: within +/-0.5% |  | Warning: within +/-1% |  | Action level: over 1%
| Symmetry | x | 0.5 | -0.5 | 0.5 | -0.5 | 1 | 1
Plan ID |  |  | Good: within +/-1% |  | Warning: within +/-2% |  | Action level: over 3%
| Flatness | x | 1 | -1 | 2 | -2 | 3 | -3
* reference measurements, Roos IC #2519, 18.11.2022
| R15M10 |  |  | R31M10 |  |  |  | R15_SOBP10_2Gy, 10 cm WET, 9.2 SW | [13 empty] | R31_SOBP10_2Gy , 26 cm WET, 24.8 SW | [17 empty] | R15_SOBP10_2Gy |  | R31_SOBP10_2Gy |  | Plot
Date | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | Dcentral ROI [1 cm x 1 cm] | Difference from reference (%) | Difference from TPS (%) | R31M10 vs. R15M10 | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Fx (%) | Fy (%) | Sx (%) | Sy (%) | 2D F (%) | 2D S L/R (%) | 2D S U/D (%) | DCen (Gy) | PLx (mm) | PRx (mm) | PLy (mm) | PRy (mm) | FWHM x (mm) | FWHM y (mm) | Sign | Checked | File name | Notes | GI (1%, 1 mm) | GI (2%, 2 mm) | GI (1%, 1 mm) | GI (2%, 2 mm) | Date | Difference from reference (a.u.) R15M10 | Medelvärde +/- 10 dygn R15M10 | Medelvärde +/- 20 dygn R15M10 | Difference from reference (a.u.) R31M10 | Medelvärde +/- 10 dygn R31M10 | Medelvärde +/- 20 dygn R31M10 | Ref | 2% | -2% | 3% | -3%
20221118 | 1.987 | 0% | -0.8483% | 2.007 | 0% | -0.9378% | 0.99003 | 0.51 | 0.58 | -0.3 | 0.18 | 1.1 | -0.32 | 0.18 | 1.987 | 8.7 | 9.2 | 8.1 | 8.3 | 119.8 | 120.2 | 0.36 | 0.39 | -0.17 | 0.03 | 0.71 | -0.16 | 0 | 2.007 | 11.5 | 11.5 | 11.3 | 11.2 | 119.4 | 119.9 | ML |  | GTR2_20221118_R15M10_01, GTR2_20221118_R31M10_01 | REFERENCE IMAGES |  |  |  |  | 2022-11-18 | 1 | 0.99323 | 0.99163 | 1 | 0.99334 | 0.99137 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20221118 | 1.983 | -0.2013% | -1.048% | 2.006 | -0.04983% | -0.9872% | 0.98853 | 0.53 | 0.54 | -0.33 | 0.14 | 1.05 | -0.3 | 0.17 | 1.983 | 8.7 | 9.2 | 8.2 | 8.2 | 119.7 | 120.2 | 0.32 | 0.33 | -0.14 | 0.02 | 0.65 | -0.16 | -0.03 | 2.006 | 11.6 | 11.5 | 11.3 | 11.2 | 119.4 | 119.9 | ML |  | GTR2_20221118_R15M10_02, GTR2_20221118_R31M10_02 |  | 100 |  | 100 |  | 2022-11-18 | 0.99799 | 0.99304 | 0.99175 | 0.9995 | 0.99315 | 0.99153 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20221118 | 1.988 | 0.05033% | -0.7984% | 2.008 | 0.04983% | -0.8885% | 0.99004 | 0.55 | 0.61 | -0.36 | 0.21 | 1.05 | -0.31 | 0.21 | 1.988 | 8.6 | 9.3 | 8.1 | 8.3 | 119.7 | 120.2 | 0.38 | 0.4 | -0.14 | 0.1 | 0.69 | -0.13 | 0 | 2.008 | 11.5 | 11.5 | 11.3 | 11.2 | 119.4 | 119.9 | ML |  | GTR2_20221118_R15M10_03, GTR2_20221118_R31M10_03 |  | 100 |  | 100 |  | 2022-11-18 | 1.0005 | 0.99299 | 0.99173 | 1.0005 | 0.99287 | 0.99148 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230313 | 1.959 | -1.409% | -2.246% | 1.986 | -1.046% | -1.974% | 0.9864 | 0.41 | 0.62 | -0.23 | 0.25 | 1.09 | -0.28 | 0.19 | 1.959 | 9 | 8.8 | 8.1 | 8.5 | 119.8 | 120.1 | 0.34 | 0.35 | -0.17 | 0.03 | 0.72 | -0.15 | -0.06 | 1.986 | 11.6 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | AJ | AC | GTR2_20230313_R15M10_01, GTR2_20230313_R31M10_01 |  | 51.8 | 100 | 69.1 | 100 | 2023-03-13 | 0.98591 | 0.9927 | 0.99179 | 0.98954 | 0.99263 | 0.99155 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230314 | 1.969 | -0.9059% | -1.747% | 1.991 | -0.7972% | -1.728% | 0.98895 | 0.65 | 0.6 | -0.31 | 0.23 | 1.09 | -0.29 | 0.24 | 1.969 | 8.8 | 9.1 | 7.9 | 8.6 | 119.8 | 120.1 | 0.32 | 0.42 | -0.14 | -0.01 | 0.73 | -0.12 | -0.07 | 1.991 | 11.6 | 11.6 | 12 | 11.7 | 119.4 | 119.8 | SM, EH | YM | GTR2_20230314_R15M10_01, GTR2_20230314_R31M10_01 |  | 89.3 | 100 | 87.5 | 100 | 2023-03-14 | 0.99094 | 0.99238 | 0.99166 | 0.99203 | 0.99229 | 0.99151 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230315 | 1.967 | -1.007% | -1.846% | 1.987 | -0.9965% | -1.925% | 0.98993 | 0.65 | 0.71 | -0.32 | 0.27 | 1.07 | -0.3 | 0.23 | 1.967 | 8.9 | 9 | 7.9 | 8.7 | 119.8 | 120.1 | 0.28 | 0.38 | -0.1 | -0.01 | 0.71 | -0.11 | -0.09 | 1.987 | 11.5 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | GF | CVD | GTR2_20230315_R15M10_01, GTR2_20230315_R31M10_01 |  | 69.6 | 100 | 77.2 | 100 | 2023-03-15 | 0.98993 | 0.9921 | 0.99165 | 0.99003 | 0.99181 | 0.99153 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230316 | 1.97 | -0.8556% | -1.697% | 1.992 | -0.7474% | -1.678% | 0.98896 | 0.52 | 0.66 | -0.28 | 0.33 | 1.12 | -0.29 | 0.24 | 1.97 | 8.7 | 9.2 | 8 | 8.5 | 119.8 | 120.1 | 0.31 | 0.4 | -0.15 | 0.05 | 0.7 | -0.16 | -0.02 | 1.992 | 11.7 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | EH | ML | GTR2_20230316_R15M10_01, GTR2_20230316_R31M10_01 |  | 95.3 | 100 | 93.8 | 100 | 2023-03-16 | 0.99144 | 0.99204 | 0.99165 | 0.99253 | 0.99162 | 0.99149 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230320 | 1.973 | -0.7046% | -1.547% | 1.989 | -0.8969% | -1.826% | 0.99196 | 0.68 | 0.61 | -0.4 | 0.23 | 1.09 | -0.32 | 0.23 | 1.973 | 8.5 | 9.3 | 8.3 | 8.6 | 119.8 | 120.1 | 0.39 | 0.37 | -0.15 | 0.03 | 0.72 | -0.15 | -0.04 | 1.989 | 11.8 | 11.7 | 11.9 | 11.6 | 119.3 | 119.8 | AC | AC,ML | GTR2_20230320_R15M10_01, GTR2_20230320_R31M10_01 |  | 100 |  | 85.9 | 100 | 2023-03-20 | 0.99295 | 0.99184 | 0.99168 | 0.99103 | 0.9915 | 0.99151 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230320 | 1.968 | -0.9562% | -1.796% | 1.987 | -0.9965% | -1.925% | 0.99044 | 0.47 | 0.64 | -0.26 | 0.17 | 1.16 | -0.28 | 0.18 | 1.968 | 8.6 | 9.3 | 8.2 | 8.5 | 119.7 | 120.1 | 0.31 | 0.38 | -0.15 | -0.01 | 0.75 | -0.16 | -0.06 | 1.987 | 11.8 | 11.7 | 11.9 | 11.5 | 119.3 | 119.8 | AC | AC,ML | GTR2_20230320_R15M10_02, GTR2_20230320_R31M10_02 |  | 73.4 | 100 | 76.9 | 100 | 2023-03-20 | 0.99044 | 0.99184 | 0.99163 | 0.99003 | 0.9915 | 0.99151 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230320 | 1.971 | -0.8052% | -1.647% | 1.988 | -0.9467% | -1.876% | 0.99145 | 0.54 | 0.59 | -0.3 | 0.15 | 1.06 | -0.3 | 0.17 | 1.971 | 8.6 | 9.3 | 8.2 | 8.5 | 119.7 | 120.1 | 0.32 | 0.37 | -0.14 | 0 | 0.74 | -0.15 | -0.05 | 1.988 | 11.8 | 11.7 | 11.9 | 11.6 | 119.3 | 119.8 | AC | AC,ML | GTR2_20230320_R15M10_03, GTR2_20230320_R31M10_03 |  | 94.3 | 100 | 80.6 | 100 | 2023-03-20 | 0.99195 | 0.99184 | 0.9915 | 0.99053 | 0.9915 | 0.99144 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230321 | 1.974 | -0.6543% | -1.497% | 1.989 | -0.8969% | -1.826% | 0.99246 | 0.68 | 0.65 | -0.34 | 0.27 | 1.05 | -0.29 | 0.23 | 1.974 | 8.4 | 9.4 | 8.1 | 8.6 | 119.7 | 120.1 | 0.33 | 0.37 | -0.11 | 0.05 | 0.73 | -0.12 | -0.04 | 1.989 | 11.7 | 11.6 | 11.8 | 11.5 | 119.4 | 119.9 | GF,AE | GB | GTR2_20230321_R15M10_01, GTR2_20230321_R31M10_01 |  | 99 | 100 | 91.3 | 100 | 2023-03-21 | 0.99346 | 0.99163 | 0.99151 | 0.99103 | 0.99137 | 0.99144 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230322 | 1.969 | -0.9059% | -1.747% | 1.989 | -0.8969% | -1.826% | 0.98994 | 0.51 | 0.63 | -0.29 | 0.21 | 1.12 | -0.29 | 0.18 | 1.969 | 8.4 | 9.3 | 8.2 | 8.4 | 119.8 | 120.2 | 0.28 | 0.34 | -0.1 | 0 | 0.67 | -0.13 | -0.05 | 1.989 | 11.9 | 11.7 | 11.9 | 11.6 | 119.4 | 119.8 | SF | CVD | GTR2_20230322_R15M10_01, GTR2_20230322_R31M10_01 |  | 93 | 100 | 88.8 | 100 | 2023-03-22 | 0.99094 | 0.99131 | 0.99148 | 0.99103 | 0.99108 | 0.99146 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230323 | 1.972 | -0.7549% | -1.597% | 1.986 | -1.046% | -1.974% | 0.99295 | 0.64 | 0.71 | -0.31 | 0.28 | 1.07 | -0.28 | 0.26 | 1.972 | 9.2 | 8.8 | 7.6 | 8.9 | 119.7 | 120.1 | 0.33 | 0.39 | -0.12 | 0 | 0.65 | -0.12 | -0.05 | 1.986 | 11.7 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | SM, OJ | ML | GTR2_20230323_R15M10_01, GTR2_20230323_R31M10_01 |  | 99 | 100 | 66.3 | 100 | 2023-03-23 | 0.99245 | 0.99097 | 0.99141 | 0.98954 | 0.99061 | 0.99145 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230324 | 1.965 | -1.107% | -1.946% | 1.986 | -1.046% | -1.974% | 0.98943 | 0.43 | 0.59 | -0.22 | 0.3 | 0.98 | -0.24 | 0.23 | 1.965 | 8.6 | 9.2 | 7.8 | 8.6 | 119.8 | 120.2 | 0.28 | 0.43 | -0.07 | -0.04 | 0.69 | -0.1 | -0.09 | 1.986 | 11.6 | 11.6 | 12.1 | 11.7 | 119.4 | 119.8 | SM, EH | KW | GTR2_20230324_R15M10_01, GTR2_20230324_R31M10_01 |  | 71.6 | 100 | 68.1 | 100 | 2023-03-24 | 0.98893 | 0.99057 | 0.99224 | 0.98954 | 0.99022 | 0.99225 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230327 | 1.963 | -1.208% | -2.046% | 1.982 | -1.246% | -2.172% | 0.99041 | 0.61 | 0.54 | -0.29 | 0.24 | 1.04 | -0.26 | 0.21 | 1.963 | 8.8 | 9.2 | 7.8 | 8.7 | 119.8 | 120.1 | 0.3 | 0.38 | -0.09 | 0.01 | 0.68 | -0.13 | -0.07 | 1.982 | 11.8 | 11.7 | 12.1 | 11.7 | 119.4 | 119.8 | AJ, GK | AC | GTR2_20230327_R15M10_01, GTR2_20230327_R31M10_01 |  | 57.9 | 100 | 51.6 | 100 | 2023-03-27 | 0.98792 | 0.99073 | 0.99229 | 0.98754 | 0.99027 | 0.99229 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230328 | 1.963 | -1.208% | -2.046% | 1.976 | -1.545% | -2.468% | 0.99342 | 0.47 | 0.64 | -0.26 | 0.33 | 1.07 | -0.27 | 0.25 | 1.963 | 9 | 8.9 | 7.6 | 8.8 | 119.8 | 120.1 | 0.29 | 0.34 | -0.12 | 0.02 | 0.72 | -0.13 | -0.06 | 1.976 | 11.4 | 11.4 | 11.7 | 11.5 | 119.4 | 119.8 | AE, GF | YM | GTR2_20230328_R15M10_01, GTR2_20230328_R31M10_01 |  | 64.9 | 100 | 49.1 | 100 | 2023-03-28 | 0.98792 | 0.99076 | 0.99236 | 0.98455 | 0.99027 | 0.99238 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230329 | 1.969 | -0.9059% | -1.747% | 1.984 | -1.146% | -2.073% | 0.99244 | 0.51 | 0.59 | -0.32 | 0.31 | 1.04 | -0.28 | 0.23 | 1.969 | 8.9 | 9 | 7.9 | 8.6 | 119.8 | 120.1 | 0.29 | 0.35 | -0.12 | -0.02 | 0.72 | -0.14 | -0.08 | 1.984 | 11.7 | 11.6 | 11.7 | 11.5 | 119.5 | 119.9 | GK,ON,EH | YM ON | GTR2_20230329_R15M10_01, GTR2_20230329_R31M10_01 |  | 98 | 100 | 69.4 | 100 | 2023-03-29 | 0.99094 | 0.99084 | 0.99239 | 0.98854 | 0.9903 | 0.99243 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230330 | 1.964 | -1.158% | -1.996% | 1.986 | -1.046% | -1.974% | 0.98892 | 0.4 | 0.66 | -0.23 | 0.18 | 1.01 | -0.25 | 0.18 | 1.964 | 8.8 | 9.1 | 8 | 8.5 | 119.8 | 120.2 | 0.34 | 0.48 | -0.14 | -0.06 | 0.74 | -0.12 | -0.07 | 1.986 | 11.6 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | JR,ON | ON | GTR2_20230330_R15M10_01, GTR2_20230330_R31M10_01 |  | 55.9 | 100 | 63.1 | 100 | 2023-03-30 | 0.98842 | 0.99089 | 0.99241 | 0.98954 | 0.99027 | 0.99247 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230331 | 1.963 | -1.208% | -2.046% | 1.985 | -1.096% | -2.024% | 0.98892 | 0.54 | 0.64 | -0.25 | 0.32 | 1.08 | -0.26 | 0.28 | 1.963 | 8.7 | 9.1 | 8 | 8.6 | 119.8 | 120.1 | 0.34 | 0.34 | -0.12 | 0.02 | 0.69 | -0.13 | -0.03 | 1.985 | 11.6 | 11.7 | 12.1 | 11.6 | 119.4 | 119.8 | AR, SM | GB | GTR2_20230331_R15M10_01, GTR2_20230331_R31M10_01 |  | 57.5 | 100 | 60 | 100 | 2023-03-31 | 0.98792 | 0.99062 | 0.99341 | 0.98904 | 0.99032 | 0.99368 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230403 | 1.975 | -0.6039% | -1.447% | 1.996 | -0.5481% | -1.481% | 0.98948 | 0.5 | 0.63 | -0.27 | 0.18 | 1.07 | -0.26 | 0.15 | 1.975 | 9 | 9 | 8.1 | 8.5 | 119.8 | 120.2 | 0.34 | 0.38 | -0.14 | -0.01 | 0.73 | -0.14 | -0.04 | 1.996 | 11.6 | 11.6 | 12 | 11.6 | 119.3 | 119.8 | SM, AR | YM | GTR2_20230403_R15M10_01, GTR2_20230403_R31M10_01 |  | 100 | 100 | 99.7 | 100 | 2023-04-03 | 0.99396 | 0.99046 | 0.99335 | 0.99452 | 0.99038 | 0.99362 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230404 | 1.97 | -0.8556% | -1.697% | 1.988 | -0.9467% | -1.876% | 0.99095 | 0.75 | 0.59 | -0.37 | 0.23 | 1.04 | -0.35 | 0.2 | 1.97 | 8.4 | 9.5 | 8.1 | 8.5 | 119.7 | 120.1 | 0.29 | 0.36 | -0.1 | 0.01 | 0.74 | -0.15 | -0.09 | 1.988 | 12 | 11.8 | 12 | 11.6 | 119.4 | 119.8 | ON | GB | GTR2_20230404_R15M10_01, GTR2_20230404_R31M10_01 |  | 82.8 | 100 | 82.5 | 100 | 2023-04-04 | 0.99144 | 0.99038 | 0.9933 | 0.99053 | 0.99038 | 0.99356 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230405 | 1.973 | -0.7046% | -1.547% | 1.993 | -0.6976% | -1.629% | 0.98996 | 0.6 | 0.65 | -0.24 | 0.24 | 1.1 | -0.24 | 0.24 | 1.973 | 9.3 | 8.5 | 7.5 | 8.8 | 119.8 | 120.1 | 0.27 | 0.32 | -0.12 | 0.03 | 0.66 | -0.14 | -0.05 | 1.993 | 11.6 | 11.7 | 11.9 | 11.5 | 119.4 | 119.8 | AE, OJ | GB | GTR2_20230405_R15M10_01, GTR2_20230405_R31M10_01 |  | 90.3 | 100 | 86.3 | 100 | 2023-04-05 | 0.99295 | 0.99206 | 0.99322 | 0.99302 | 0.9921 | 0.99349 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230406 | 1.965 | -1.107% | -1.946% | 1.988 | -0.9467% | -1.876% | 0.98843 | 0.62 | 0.76 | -0.26 | 0.32 | 1.17 | -0.28 | 0.29 | 1.965 | 9.5 | 8 | 6.9 | 9.4 | 119.8 | 120.1 | 0.32 | 0.38 | -0.15 | -0.02 | 0.73 | -0.17 | -0.08 | 1.988 | 11.8 | 12.1 | 12.2 | 11.8 | 119.4 | 119.8 | EH | CVD | GTR2_20230406_R15M10_01, GTR2_20230406_R31M10_01 |  | 57.1 | 93.4 | 50.9 | 90.3 | 2023-04-06 | 0.98893 | 0.99218 | 0.9933 | 0.99053 | 0.9922 | 0.99353 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230407 | 1.97 | -0.8556% | -1.697% | 1.991 | -0.7972% | -1.728% | 0.98945 | 0.53 | 0.62 | -0.29 | 0.2 | 1.06 | -0.28 | 0.2 | 1.97 | 8.5 | 9.3 | 8 | 8.5 | 119.8 | 120.2 | 0.26 | 0.4 | -0.09 | -0.05 | 0.73 | -0.12 | -0.07 | 1.991 | 11.7 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | SM, JR | CVD | GTR2_20230407_R15M10_01, GTR2_20230407_R31M10_01 |  | 83.2 | 100 | 97.5 | 100 | 2023-04-07 | 0.99144 | 0.99235 | 0.99333 | 0.99203 | 0.99241 | 0.99354 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230411 | 1.97 | -0.8556% | -1.697% | 1.988 | -0.9467% | -1.876% | 0.99095 | 0.55 | 0.61 | -0.26 | 0.26 | 1.06 | -0.25 | 0.23 | 1.97 | 8.8 | 9.1 | 7.9 | 8.7 | 119.8 | 120.1 | 0.33 | 0.41 | -0.13 | -0.03 | 0.72 | -0.12 | -0.07 | 1.988 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | GF, EH, ON | GB | GTR2_20230411_R15M10_01, GTR2_20230411_R31M10_01 |  | 86.3 | 100 | 81.3 | 100 | 2023-04-11 | 0.99144 | 0.99247 | 0.99336 | 0.99053 | 0.99261 | 0.99358 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230412 | 1.972 | -0.7549% | -1.597% | 1.991 | -0.7972% | -1.728% | 0.99046 | 0.75 | 0.69 | -0.37 | 0.29 | 1.12 | -0.32 | 0.25 | 1.972 | 8.4 | 9.5 | 8 | 8.7 | 119.7 | 120.1 | 0.34 | 0.38 | -0.1 | 0.06 | 0.72 | -0.12 | -0.02 | 1.991 | 11.9 | 11.7 | 11.9 | 11.5 | 119.4 | 119.8 | JR,GF | YM | GTR2_20230412_R15M10_01, GTR2_20230412_R31M10_01 |  | 94.9 | 100 | 90.9 | 100 | 2023-04-12 | 0.99245 | 0.99253 | 0.99338 | 0.99203 | 0.99274 | 0.99359 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230413 | 1.968 | -0.9562% | -1.796% | 1.99 | -0.847% | -1.777% | 0.98894 | 0.57 | 0.64 | -0.29 | 0.28 | 1.2 | -0.27 | 0.25 | 1.968 | 8.7 | 9.3 | 7.9 | 8.7 | 119.8 | 120.1 | 0.3 | 0.39 | -0.08 | 0.02 | 0.74 | -0.11 | -0.04 | 1.99 | 11.8 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | SM, AE | YM | GTR2_20230413_R15M10_01, GTR2_20230413_R31M10_01 |  | 78.5 | 100 | 87.5 | 100 | 2023-04-13 | 0.99044 | 0.99291 | 0.99339 | 0.99153 | 0.9931 | 0.99362 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230414 | 1.963 | -1.208% | -2.046% | 1.986 | -1.046% | -1.974% | 0.98842 | 0.53 | 0.64 | -0.28 | 0.29 | 1.04 | -0.29 | 0.25 | 1.963 | 8.8 | 9.1 | 7.8 | 8.7 | 119.8 | 120.1 | 0.31 | 0.36 | -0.11 | 0.03 | 0.72 | -0.13 | -0.04 | 1.986 | 11.7 | 11.7 | 11.9 | 11.6 | 119.4 | 119.9 | OJ, SM | GB | GTR2_20230414_R15M10_01, GTR2_20230414_R31M10_01 |  | 53.2 | 100 | 75.9 | 100 | 2023-04-14 | 0.98792 | 0.99376 | 0.99342 | 0.98954 | 0.99416 | 0.99366 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230417 | 1.971 | -0.8052% | -1.647% | 1.99 | -0.847% | -1.777% | 0.99045 | 0.52 | 0.62 | -0.31 | 0.13 | 1.02 | -0.26 | 0.14 | 1.971 | 8.6 | 9.3 | 8.1 | 8.5 | 119.7 | 120.1 | 0.29 | 0.39 | -0.07 | 0 | 0.75 | -0.11 | -0.08 | 1.99 | 11.8 | 11.7 | 11.8 | 11.5 | 119.4 | 119.8 | AJ, OJ | GB | GTR2_20230417_R15M10_01, GTR2_20230417_R31M10_01 |  | 90.6 |  | 87.8 | 100 | 2023-04-17 | 0.99195 | 0.99376 | 0.99344 | 0.99153 | 0.99416 | 0.99369 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230418 | 1.968 | -0.9562% | -1.796% | 1.991 | -0.7972% | -1.728% | 0.98845 | 0.6 | 0.63 | -0.3 | 0.28 | 1.02 | -0.3 | 0.23 | 1.968 | 7.9 | 9.6 | 7.6 | 8.8 | 119.7 | 120.1 | 0.29 | 0.38 | -0.08 | -0.01 | 0.72 | -0.12 | -0.08 | 1.991 | 12.2 | 11.8 | 12.3 | 11.7 | 119.4 | 119.8 | GK,GF | GB | GTR2_20230418_R15M10_01, GTR2_20230418_R31M10_01 |  | 78.1 | 100 | 81.9 | 100 | 2023-04-18 | 0.99044 | 0.99383 | 0.99343 | 0.99203 | 0.99422 | 0.99372 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230419 | 1.966 | -1.057% | -1.896% | 1.989 | -0.8969% | -1.826% | 0.98844 | 0.64 | 0.58 | -0.34 | 0.25 | 1.13 | -0.34 | 0.22 | 1.966 | 9.1 | 8.9 | 7.9 | 8.7 | 119.8 | 120.1 | 0.34 | 0.37 | -0.15 | 0 | 0.71 | -0.15 | -0.04 | 1.989 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | OJ, AR | CVD | GTR2_20230419_R15M10_01, GTR2_20230419_R31M10_01 |  | 71.6 | 100 | 92.5 | 100 | 2023-04-19 | 0.98943 | 0.99383 | 0.99346 | 0.99103 | 0.99422 | 0.99375 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230420 | 1.971 | -0.8052% | -1.647% | 1.993 | -0.6976% | -1.629% | 0.98896 | 0.52 | 0.56 | -0.3 | 0.16 | 1.16 | -0.26 | 0.19 | 1.971 | 8.9 | 9.1 | 8.1 | 8.5 | 119.8 | 120.2 | 0.34 | 0.37 | -0.14 | -0.01 | 0.69 | -0.15 | -0.04 | 1.993 | 11.5 | 11.5 | 11.9 | 11.5 | 119.4 | 119.8 | ON, EH | CVD | GTR2_20230420_R15M10_01, GTR2_20230420_R31M10_01 |  | 95.7 | 100 | 96.6 | 100 | 2023-04-20 | 0.99195 | 0.99386 | 0.99347 | 0.99302 | 0.99427 | 0.9938 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230421 | 1.967 | -1.007% | -1.846% | 1.988 | -0.9467% | -1.876% | 0.98944 | 0.45 | 0.6 | -0.26 | 0.21 | 1.03 | -0.28 | 0.15 | 1.967 | 8.8 | 9.2 | 8.1 | 8.5 | 119.8 | 120.2 | 0.33 | 0.4 | -0.13 | -0.01 | 0.71 | -0.11 | -0.05 | 1.988 | 11.7 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | SM, ON | CVD | GTR2_20230421_R15M10_01, GTR2_20230421_R31M10_01 |  | 76.3 | 100 | 81.3 | 100 | 2023-04-21 | 0.98993 | 0.99387 | 0.99353 | 0.99053 | 0.99428 | 0.99385 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230424 | 1.966 | -1.057% | -1.896% | 1.987 | -0.9965% | -1.925% | 0.98943 | 0.48 | 0.62 | -0.28 | 0.25 | 1.14 | -0.3 | 0.18 | 1.966 | 8.9 | 9 | 8.1 | 8.5 | 119.8 | 120.2 | 0.3 | 0.38 | -0.08 | 0.01 | 0.69 | -0.11 | -0.07 | 1.987 | 11.5 | 11.5 | 11.9 | 11.6 | 119.4 | 119.9 | GK,AE | YM | GTR2_20230424_R15M10_01, GTR2_20230424_R31M10_01 |  | 66.9 | 100 | 80.3 | 100 | 2023-04-24 | 0.98943 | 0.99393 | 0.99359 | 0.99003 | 0.99433 | 0.99393 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230425 | 1.964 | -1.158% | -1.996% | 1.988 | -0.9467% | -1.876% | 0.98793 | 0.49 | 0.54 | -0.26 | 0.19 | 0.99 | -0.28 | 0.18 | 1.964 | 8.6 | 9.2 | 8 | 8.4 | 119.8 | 120.2 | 0.27 | 0.4 | -0.1 | -0.03 | 0.69 | -0.1 | -0.04 | 1.988 | 11.7 | 11.6 | 11.7 | 11.4 | 119.4 | 119.9 | AE, GF | CVD | GTR2_20230425_R15M10_01, GTR2_20230425_R31M10_01 |  | 69 | 100 | 86.3 | 100 | 2023-04-25 | 0.98842 | 0.99397 | 0.99366 | 0.99053 | 0.99436 | 0.99404 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230426 | 1.992 | 0.2516% | -0.5988% | 2.001 | -0.299% | -1.234% | 0.9955 | 0.66 | 0.52 | -0.33 | 0.22 | 1.04 | -0.28 | 0.21 | 1.992 | 8.8 | 9.2 | 7.6 | 8.6 | 119.8 | 120.2 | 0.36 | 0.3 | -0.16 | 0.02 | 0.71 | -0.16 | -0.05 | 2.001 | 11.3 | 11.4 | 11.8 | 11.4 | 119.4 | 119.9 | AR, GK | YM | GTR2_20230426_R15M10_01, GTR2_20230426_R31M10_01 |  | 100 | 100 | 93.1 | 100 | 2023-04-26 | 1.0025 | 0.994 | 0.99369 | 0.99701 | 0.99441 | 0.9941 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230427 | 1.98 | -0.3523% | -1.198% | 1.998 | -0.4484% | -1.382% | 0.99099 | 0.51 | 0.57 | -0.26 | 0.27 | 1.12 | -0.27 | 0.24 | 1.98 | 8.8 | 9.2 | 7.9 | 8.5 | 119.8 | 120.2 | 0.33 | 0.33 | -0.17 | -0.01 | 0.72 | -0.15 | -0.04 | 1.998 | 11.5 | 11.6 | 11.8 | 11.5 | 119.4 | 119.9 | GK,OJ | CVD | GTR2_20230427_R15M10_01, GTR2_20230427_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-04-27 | 0.99648 | 0.99402 | 0.99376 | 0.99552 | 0.99445 | 0.99416 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230428 | 1.969 | -0.9059% | -1.747% | 1.992 | -0.7474% | -1.678% | 0.98845 | 0.59 | 0.62 | -0.28 | 0.3 | 1.05 | -0.26 | 0.25 | 1.969 | 9 | 8.9 | 7.8 | 8.6 | 119.8 | 120.2 | 0.32 | 0.4 | -0.11 | -0.01 | 0.72 | -0.13 | -0.09 | 1.992 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | GK,JR | AC | GTR2_20230428_R15M10_01, GTR2_20230428_R31M10_01 |  | 87 | 100 | 94.7 | 100 | 2023-04-28 | 0.99094 | 0.99407 | 0.99383 | 0.99253 | 0.99449 | 0.99422 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230502 | 1.971 | -0.8052% | -1.647% | 1.993 | -0.6976% | -1.629% | 0.98896 | 0.49 | 0.57 | -0.28 | 0.2 | 1.1 | -0.27 | 0.17 | 1.971 | 8.7 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.3 | 0.39 | -0.1 | 0 | 0.71 | -0.13 | -0.07 | 1.993 | 11.5 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | AE, AR, EF | CVD | GTR2_20230502_R15M10_01, GTR2_20230502_R31M10_01 |  | 97.7 | 100 | 98.8 | 100 | 2023-05-02 | 0.99195 | 0.99416 | 0.99383 | 0.99302 | 0.99455 | 0.99422 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230503 | 1.97 | -0.8556% | -1.697% | 1.99 | -0.847% | -1.777% | 0.98995 | 0.59 | 0.63 | -0.32 | 0.16 | 1.08 | -0.28 | 0.14 | 1.97 | 8.8 | 9.2 | 8.2 | 8.3 | 119.8 | 120.2 | 0.3 | 0.33 | -0.14 | 0 | 0.73 | -0.15 | -0.05 | 1.99 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | ON | CVD | GTR2_20230503_R15M10_01, GTR2_20230503_R31M10_01 |  | 95.7 | 100 | 98.1 | 100 | 2023-05-03 | 0.99144 | 0.99419 | 0.99386 | 0.99153 | 0.9946 | 0.99427 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230504 | 1.972 | -0.7549% | -1.597% | 1.995 | -0.5979% | -1.53% | 0.98847 | 0.75 | 0.56 | -0.39 | 0.25 | 1.04 | -0.31 | 0.22 | 1.972 | 8.4 | 9.4 | 7.8 | 8.6 | 119.7 | 120.1 | 0.38 | 0.35 | -0.16 | -0.01 | 0.71 | -0.15 | -0.06 | 1.995 | 11.9 | 11.7 | 11.9 | 11.6 | 119.4 | 119.8 | JR | YM | GTR2_20230504_R15M10_01, GTR2_20230504_R31M10_01 |  | 96.6 | 100 | 100 | 100 | 2023-05-04 | 0.99245 | 0.99424 | 0.99387 | 0.99402 | 0.99463 | 0.99428 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230504 | 1.975 | -0.6039% | -1.447% | 1.995 | -0.5979% | -1.53% | 0.98997 | 0.49 | 0.59 | -0.29 | 0.14 | 1.09 | -0.26 | 0.15 | 1.975 | 8.8 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.36 | 0.36 | -0.13 | -0.01 | 0.73 | -0.14 | -0.06 | 1.995 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | ON, EF | ML | GTR2_20230504_R15M10_02, GTR2_20230504_R31M10_02 |  | 99.7 |  | 100 |  | 2023-05-04 | 0.99396 | 0.99431 | 0.99393 | 0.99402 | 0.99469 | 0.99433 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230505 | 1.971 | -0.8052% | -1.647% | 1.992 | -0.7474% | -1.678% | 0.98946 | 0.43 | 0.6 | -0.24 | 0.21 | 0.97 | -0.24 | 0.17 | 1.971 | 8.7 | 9.2 | 8 | 8.5 | 119.8 | 120.2 | 0.3 | 0.33 | -0.12 | 0.02 | 0.74 | -0.14 | -0.06 | 1.992 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | SM, GK | AC | GTR2_20230505_R15M10_01, GTR2_20230505_R31M10_01 |  | 95.7 | 100 | 98.8 | 100 | 2023-05-05 | 0.99195 | 0.99435 | 0.99397 | 0.99253 | 0.99471 | 0.99436 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230508 | 1.97 | -0.8556% | -1.697% | 1.988 | -0.9467% | -1.876% | 0.99095 | 0.58 | 0.6 | -0.33 | 0.29 | 1.15 | -0.31 | 0.24 | 1.97 | 8.9 | 9.1 | 7.9 | 8.6 | 119.8 | 120.2 | 0.38 | 0.36 | -0.17 | -0.01 | 0.73 | -0.16 | -0.07 | 1.988 | 11.5 | 11.5 | 11.9 | 11.5 | 119.4 | 119.9 | OJ, SF | AC | GTR2_20230508_R15M10_01, GTR2_20230508_R31M10_01 |  | 93.8 | 100 | 79.7 | 100 | 2023-05-08 | 0.99144 | 0.99449 | 0.99402 | 0.99053 | 0.99484 | 0.99445 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230509 | 1.967 | -1.007% | -1.846% | 1.983 | -1.196% | -2.122% | 0.99193 | 0.65 | 0.63 | -0.36 | 0.35 | 1.11 | -0.32 | 0.28 | 1.967 | 8.7 | 9.2 | 7.1 | 8.9 | 119.8 | 120.2 | 0.37 | 0.35 | -0.16 | 0.01 | 0.7 | -0.15 | -0.03 | 1.983 | 11.6 | 11.6 | 11.7 | 11.4 | 119.5 | 120 | GF | GB | GTR2_20230509_R15M10_01, GTR2_20230509_R31M10_01 | [5 empty] | 2023-05-09 | 0.98993 | 0.99458 | 0.99407 | 0.98804 | 0.99491 | 0.99449 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230510 | 1.968 | -0.9562% | -1.796% | 1.987 | -0.9965% | -1.925% | 0.99044 | 0.56 | 0.58 | -0.31 | 0.18 | 1.05 | -0.26 | 0.17 | 1.968 | 8.7 | 9.2 | 7.9 | 8.4 | 119.8 | 120.2 | 0.33 | 0.36 | -0.1 | 0.01 | 0.69 | -0.13 | -0.04 | 1.987 | 11.7 | 11.6 | 11.7 | 11.4 | 119.4 | 119.9 | ON | GB | GTR2_20230510_R15M10_01, GTR2_20230510_R31M10_01 |  | 81.5 | 100 | 79.7 | 100 | 2023-05-10 | 0.99044 | 0.99446 | 0.99416 | 0.99003 | 0.99488 | 0.99455 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230511 | 1.973 | -0.7046% | -1.547% | 1.991 | -0.7972% | -1.728% | 0.99096 | 0.81 | 0.59 | -0.39 | 0.26 | 1.09 | -0.33 | 0.22 | 1.973 | 8.4 | 9.4 | 7.7 | 8.7 | 119.7 | 120.1 | 0.35 | 0.4 | -0.16 | -0.02 | 0.78 | -0.17 | -0.06 | 1.991 | 11.9 | 11.7 | 11.9 | 11.5 | 119.4 | 119.9 | AR | CVD | GTR2_20230511_R15M10_01, GTR2_20230511_R31M10_01 |  | 99.7 | 100 | 97.5 | 100 | 2023-05-11 | 0.99295 | 0.99442 | 0.99419 | 0.99203 | 0.99487 | 0.9946 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230512 | 1.976 | -0.5536% | -1.397% | 1.999 | -0.3986% | -1.333% | 0.98849 | 0.54 | 0.55 | -0.29 | 0.13 | 1.05 | -0.28 | 0.12 | 1.976 | 8.7 | 9.3 | 8.1 | 8.3 | 119.8 | 120.2 | 0.35 | 0.35 | -0.12 | 0.01 | 0.7 | -0.14 | -0.05 | 1.999 | 11.8 | 11.7 | 11.8 | 11.5 | 119.4 | 119.9 | SF, JR | EAL | GTR2_20230512_R15M10_01, GTR2_20230512_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-12 | 0.99446 | 0.99448 | 0.99424 | 0.99601 | 0.9949 | 0.99463 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230515 | 1.966 | -1.057% | -1.896% | 1.985 | -1.096% | -2.024% | 0.99043 | 0.58 | 0.58 | -0.32 | 0.29 | 1.14 | -0.28 | 0.25 | 1.966 | 8.7 | 9.3 | 7.8 | 8.7 | 119.7 | 120.1 | 0.29 | 0.35 | -0.11 | 0.01 | 0.75 | -0.14 | -0.05 | 1.985 | 11.7 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | OJ, ON, EN | GB | GTR2_20230515_R15M10_01, GTR2_20230515_R31M10_01 |  | 71 | 100 | 60 | 100 | 2023-05-15 | 0.98943 | 0.99452 | 0.99431 | 0.98904 | 0.99494 | 0.99469 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230516 | 1.967 | -1.007% | -1.846% | 1.992 | -0.7474% | -1.678% | 0.98745 | 0.67 | 0.62 | -0.36 | 0.13 | 1.18 | -0.34 | 0.13 | 1.967 | 8.6 | 9.4 | 8.1 | 8.5 | 119.8 | 120.2 | 0.4 | 0.36 | -0.16 | 0.01 | 0.76 | -0.17 | -0.04 | 1.992 | 11.8 | 11.7 | 12 | 11.6 | 119.4 | 119.8 | EN,AJ,AR | ON, YM | GTR2_20230516_R15M10_01, GTR2_20230516_R31M10_01 |  | 66.7 | 100 | 97.8 | 100 | 2023-05-16 | 0.98993 | 0.99457 | 0.99435 | 0.99253 | 0.99499 | 0.99471 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230517 | 1.977 | -0.5033% | -1.347% | 1.993 | -0.6976% | -1.629% | 0.99197 | 0.61 | 0.61 | -0.37 | 0.26 | 1.14 | -0.34 | 0.21 | 1.977 | 8.8 | 9.2 | 7.9 | 8.6 | 119.7 | 120.1 | 0.31 | 0.35 | -0.14 | 0.01 | 0.73 | -0.15 | -0.04 | 1.993 | 11.6 | 11.6 | 11.9 | 11.5 | 119.4 | 119.8 | AE, OJ, EF | YM ON | GTR2_20230517_R15M10_01, GTR2_20230517_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-17 | 0.99497 | 0.99461 | 0.99441 | 0.99302 | 0.99501 | 0.99477 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230519 | 1.979 | -0.4026% | -1.248% | 2.001 | -0.299% | -1.234% | 0.98901 | 0.64 | 0.57 | -0.32 | 0.23 | 1.04 | -0.26 | 0.2 | 1.979 | 8.7 | 9.3 | 7.7 | 8.7 | 119.7 | 120.1 | 0.32 | 0.43 | -0.11 | -0.02 | 0.69 | -0.15 | -0.06 | 2.001 | 11.5 | 11.5 | 11.9 | 11.5 | 119.4 | 119.8 | ON | AD | GTR2_20230519_R15M10_01, GTR2_20230519_R31M10_01 |  | 100 |  | 99.1 | 100 | 2023-05-19 | 0.99597 | 0.99462 | 0.99449 | 0.99701 | 0.99503 | 0.99484 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230522 | 1.969 | -0.9059% | -1.747% | 1.989 | -0.8969% | -1.826% | 0.98994 | 0.56 | 0.6 | -0.3 | 0.12 | 1.16 | -0.3 | 0.14 | 1.969 | 8.9 | 9.1 | 8.3 | 8.4 | 119.8 | 120.1 | 0.31 | 0.39 | -0.13 | -0.02 | 0.7 | -0.14 | -0.05 | 1.989 | 11.5 | 11.5 | 11.8 | 11.5 | 119.4 | 119.8 | AE | GB | GTR2_20230522_R15M10_01, GTR2_20230522_R31M10_01 |  | 87.3 | 100 | 93.4 | 100 | 2023-05-22 | 0.99094 | 0.99467 | 0.99458 | 0.99103 | 0.99507 | 0.99491 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230523 | 1.981 | -0.302% | -1.148% | 1.99 | -0.847% | -1.777% | 0.99548 | 0.71 | 0.63 | -0.36 | 0.24 | 1.1 | -0.3 | 0.22 | 1.981 | 9.2 | 9 | 7.6 | 9.1 | 119.7 | 120 | 0.32 | 0.31 | -0.12 | 0.02 | 0.69 | -0.11 | -0.06 | 1.99 | 11.6 | 11.6 | 11.8 | 11.6 | 119.4 | 119.8 | GF, ON | CVD | GTR2_20230523_R15M10_01, GTR2_20230523_R31M10_01 |  | 100 | 100 | 90.9 | 100 | 2023-05-23 | 0.99698 | 0.99467 | 0.99446 | 0.99153 | 0.99507 | 0.99488 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230524 | 1.968 | -0.9562% | -1.796% | 1.988 | -0.9467% | -1.876% | 0.98994 | 0.54 | 0.63 | -0.28 | 0.29 | 1.03 | -0.26 | 0.24 | 1.968 | 8.7 | 9.2 | 7.8 | 8.7 | 119.8 | 120.1 | 0.36 | 0.34 | -0.15 | 0.04 | 0.69 | -0.15 | -0.03 | 1.988 | 11.6 | 11.6 | 12 | 11.6 | 119.4 | 119.8 | AE, SM | AD | GTR2_20230524_R15M10_01, GTR2_20230524_R31M10_01 |  | 79.9 | 100 | 85 | 100 | 2023-05-24 | 0.99044 | 0.99472 | 0.99442 | 0.99053 | 0.99515 | 0.99487 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230525 | 1.969 | -0.9059% | -1.747% | 1.989 | -0.8969% | -1.826% | 0.98994 | 0.6 | 0.58 | -0.32 | 0.12 | 1.07 | -0.3 | 0.12 | 1.969 | 8.6 | 9.4 | 8.1 | 8.5 | 119.7 | 120.1 | 0.32 | 0.37 | -0.17 | -0.04 | 0.72 | -0.17 | -0.07 | 1.989 | 11.8 | 11.7 | 11.9 | 11.6 | 119.4 | 119.8 | AR, SM | GB | GTR2_20230525_R15M10_01, GTR2_20230525_R31M10_01 |  | 83.7 | 100 | 96.6 | 100 | 2023-05-25 | 0.99094 | 0.99481 | 0.99448 | 0.99103 | 0.99528 | 0.9949 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230526 | 1.969 | -0.9059% | -1.747% | 1.989 | -0.8969% | -1.826% | 0.98994 | 0.74 | 0.58 | -0.32 | 0.21 | 1.06 | -0.28 | 0.17 | 1.969 | 9 | 9.1 | 7.9 | 8.7 | 119.7 | 120.1 | 0.29 | 0.33 | -0.12 | 0.01 | 0.78 | -0.16 | -0.05 | 1.989 | 11.5 | 11.5 | 11.8 | 11.5 | 119.4 | 119.8 | OJ, EN | AC | GTR2_20230526_R15M10_01, GTR2_20230526_R31M10_01 |  | 86.3 | 100 | 85.9 | 100 | 2023-05-26 | 0.99094 | 0.99489 | 0.99452 | 0.99103 | 0.99538 | 0.99494 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230529 | 1.968 | -0.9562% | -1.796% | 1.988 | -0.9467% | -1.876% | 0.98994 | 0.52 | 0.63 | -0.26 | 0.28 | 1.04 | -0.26 | 0.25 | 1.968 | 9 | 9.1 | 8 | 8.7 | 119.7 | 120.1 | 0.37 | 0.33 | -0.16 | 0.01 | 0.75 | -0.17 | -0.04 | 1.988 | 11.6 | 11.5 | 11.8 | 11.5 | 119.4 | 119.8 | OJ | CVD | GTR2_20230529_R15M10_01, GTR2_20230529_R31M10_01 |  | 71.9 | 100 | 87.8 | 100 | 2023-05-29 | 0.99044 | 0.99493 | 0.99457 | 0.99053 | 0.99544 | 0.99499 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230530 | 1.975 | -0.6039% | -1.447% | 1.993 | -0.6976% | -1.629% | 0.99097 | 0.74 | 0.58 | -0.36 | 0.19 | 1.12 | -0.28 | 0.16 | 1.975 | 8.7 | 9.3 | 8 | 8.7 | 119.7 | 120.1 | 0.37 | 0.35 | -0.17 | 0.03 | 0.72 | -0.16 | -0.03 | 1.993 | 11.7 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | AE | YM | GTR2_20230530_R15M10_01, GTR2_20230530_R31M10_01 |  | 100 | 100 | 100 | 100 | 2023-05-30 | 0.99396 | 0.99494 | 0.99461 | 0.99302 | 0.99543 | 0.99501 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230531 | 1.97 | -0.8556% | -1.697% | 1.99 | -0.847% | -1.777% | 0.98995 | 0.66 | 0.65 | -0.34 | 0.27 | 1.12 | -0.3 | 0.22 | 1.97 | 8.9 | 9.1 | 7.8 | 8.8 | 119.8 | 120.1 | 0.39 | 0.43 | -0.17 | -0.02 | 0.76 | -0.17 | -0.07 | 1.99 | 11.6 | 11.6 | 12.1 | 11.7 | 119.4 | 119.7 | JR | YM | GTR2_20230531_R15M10_01, GTR2_20230531_R31M10_01 |  | 91 | 100 | 86.3 | 100 | 2023-05-31 | 0.99144 | 0.99505 | 0.99462 | 0.99153 | 0.99555 | 0.99503 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230601 | 1.977 | -0.5033% | -1.347% | 1.996 | -0.5481% | -1.481% | 0.99048 | 0.43 | 0.58 | -0.26 | 0.15 | 1.09 | -0.26 | 0.17 | 1.977 | 8.7 | 9.2 | 8.1 | 8.5 | 119.8 | 120.1 | 0.26 | 0.38 | -0.1 | 0.01 | 0.66 | -0.1 | -0.05 | 1.996 | 11.7 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | EF | AD | GTR2_20230601_R15M10_01, GTR2_20230601_R31M10_01 |  | 100 |  | 100 |  | 2023-06-01 | 0.99497 | 0.99515 | 0.99467 | 0.99452 | 0.99562 | 0.99507 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230602 | 1.966 | -1.057% | -1.896% | 1.987 | -0.9965% | -1.925% | 0.98943 | 0.41 | 0.58 | -0.25 | 0.15 | 1.04 | -0.25 | 0.17 | 1.966 | 9 | 8.9 | 8.2 | 8.4 | 119.8 | 120.1 | 0.28 | 0.41 | -0.09 | -0.01 | 0.74 | -0.12 | -0.06 | 1.987 | 11.6 | 11.6 | 11.9 | 11.6 | 119.4 | 119.8 | AJ | AC | GTR2_20230602_R15M10_01, GTR2_20230602_R31M10_01 |  | 68.2 | 100 | 79.1 | 100 | 2023-06-02 | 0.98943 | 0.99515 | 0.99467 | 0.99003 | 0.99567 | 0.99507 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230605 | 1.977 | -0.5033% | -1.347% | 1.997 | -0.4983% | -1.431% | 0.98998 | 0.68 | 0.61 | -0.35 | 0.24 | 1.12 | -0.31 | 0.22 | 1.977 | 8.7 | 9.3 | 8.1 | 8.6 | 119.7 | 120.1 | 0.41 | 0.41 | -0.14 | 0.02 | 0.77 | -0.16 | -0.02 | 1.997 | 11.6 | 11.6 | 11.8 | 11.5 | 119.3 | 119.7 | EF | AC/KW | GTR2_20230605_R15M10_01, GTR2_20230605_R31M10_01 |  | 100 |  | 100 |  | 2023-06-05 | 0.99497 | 0.99514 | 0.99472 | 0.99502 | 0.99564 | 0.99515 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230607 | 1.981 | -0.302% | -1.148% | 2.006 | -0.04983% | -0.9872% | 0.98754 | 0.5 | 0.57 | -0.27 | 0.14 | 1.07 | -0.26 | 0.14 | 1.981 | 8.8 | 9.1 | 8.2 | 8.4 | 119.8 | 120.1 | 0.37 | 0.36 | -0.14 | -0.01 | 0.71 | -0.14 | -0.07 | 2.006 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | ON | YM | GTR2_20230607_R15M10_01, GTR2_20230607_R31M10_01 |  | 100 |  | 100 |  | 2023-06-07 | 0.99698 | 0.99522 | 0.99481 | 0.9995 | 0.99574 | 0.99528 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230608 | 1.985 | -0.1007% | -0.9481% | 2.007 | 0% | -0.9378% | 0.98904 | 0.74 | 0.56 | -0.33 | 0.21 | 1.15 | -0.3 | 0.18 | 1.985 | 9 | 9 | 8 | 8.6 | 119.7 | 120.1 | 0.35 | 0.38 | -0.17 | -0.03 | 0.72 | -0.17 | -0.06 | 2.007 | 11.5 | 11.6 | 11.8 | 11.5 | 119.4 | 119.7 | EN |  | GTR2_20230608_R15M10_01, GTR2_20230608_R31M10_01 |  | 100 |  | 100 |  | 2023-06-08 | 0.99899 | 0.99519 | 0.99489 | 1 | 0.99583 | 0.99538 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230609 | 1.986 | -0.05033% | -0.8982% | 2.006 | -0.04983% | -0.9872% | 0.99003 | 0.72 | 0.54 | -0.37 | 0.23 | 1.02 | -0.3 | 0.21 | 1.986 | 8.8 | 9.2 | 7.9 | 8.6 | 119.7 | 120.1 | 0.29 | 0.36 | -0.11 | 0.01 | 0.66 | -0.11 | -0.02 | 2.006 | 11.3 | 11.4 | 11.5 | 11.4 | 119.4 | 119.8 | AR | AC | GTR2_20230609_R15M10_01, GTR2_20230609_R31M10_01 |  | 100 |  | 98.4 | 100 | 2023-06-09 | 0.9995 | 0.99529 | 0.99493 | 0.9995 | 0.99595 | 0.99544 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230612 | 1.976 | -0.5536% | -1.397% | 1.996 | -0.5481% | -1.481% | 0.98998 | 0.49 | 0.6 | -0.24 | 0.28 | 1.08 | -0.28 | 0.22 | 1.976 | 8.9 | 9 | 7.9 | 8.6 | 119.8 | 120.1 | 0.28 | 0.33 | -0.09 | 0.05 | 0.68 | -0.13 | -0.05 | 1.996 | 11.5 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | AC | AC | GTR2_20230612_R15M10_01, GTR2_20230612_R31M10_01 | [5 empty] | 2023-06-12 | 0.99446 | 0.99539 | 0.99494 | 0.99452 | 0.99606 | 0.99543 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230613 | 1.978 | -0.4529% | -1.297% | 1.997 | -0.4983% | -1.431% | 0.99049 | 0.59 | 0.63 | -0.31 | 0.28 | 1.13 | -0.32 | 0.24 | 1.978 | 8.7 | 9.3 | 8 | 8.6 | 119.7 | 120.1 | 0.25 | 0.37 | -0.09 | -0.02 | 0.64 | -0.11 | -0.09 | 1.997 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | OJ | GB | GTR2_20230613_R15M10_01, GTR2_20230613_R31M10_01 |  | 100 |  | 100 |  | 2023-06-13 | 0.99547 | 0.99549 | 0.99505 | 0.99502 | 0.99618 | 0.99555 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230614 | 1.981 | -0.302% | -1.148% | 2.004 | -0.1495% | -1.086% | 0.98852 | 0.72 | 0.59 | -0.37 | 0.26 | 1.12 | -0.3 | 0.22 | 1.981 | 9.1 | 9 | 7.7 | 8.8 | 119.7 | 120.1 | 0.39 | 0.38 | -0.18 | 0.04 | 0.77 | -0.19 | -0.02 | 2.004 | 11.5 | 11.5 | 11.4 | 11.3 | 119.4 | 119.8 | EF | YM | GTR2_20230614_R15M10_01, GTR2_20230614_R31M10_01 |  | 100 |  | 100 |  | 2023-06-14 | 0.99698 | 0.99561 | 0.99515 | 0.99851 | 0.99631 | 0.99562 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230615 | 1.976 | -0.5536% | -1.397% | 1.998 | -0.4484% | -1.382% | 0.98899 | 0.68 | 0.63 | -0.39 | 0.31 | 1.15 | -0.34 | 0.24 | 1.976 | 9 | 9 | 7.8 | 8.7 | 119.7 | 120.1 | 0.33 | 0.34 | -0.15 | 0.02 | 0.73 | -0.14 | -0.04 | 1.998 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EF | YM | GTR2_20230615_R15M10_01, GTR2_20230615_R31M10_01 |  | 100 |  | 100 |  | 2023-06-15 | 0.99446 | 0.99565 | 0.99515 | 0.99552 | 0.99639 | 0.99567 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230616 | 1.975 | -0.6039% | -1.447% | 1.998 | -0.4484% | -1.382% | 0.98849 | 0.44 | 0.6 | -0.25 | 0.18 | 1.09 | -0.28 | 0.17 | 1.975 | 8.8 | 9.1 | 8.4 | 8.3 | 119.8 | 120.1 | 0.34 | 0.35 | -0.15 | 0.01 | 0.71 | -0.13 | -0.04 | 1.998 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EN | AC | GTR2_20230616_R15M10_01, GTR2_20230616_R31M10_01 |  | 100 |  | 100 |  | 2023-06-16 | 0.99396 | 0.99576 | 0.99514 | 0.99552 | 0.99651 | 0.99564 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230619 | 2.005 | 0.9059% | 0.0499% | 2.025 | 0.8969% | -0.04936% | 0.99012 | 0.5 | 0.64 | -0.23 | 0.26 | 1 | -0.27 | 0.22 | 2.005 | 8.9 | 9 | 8 | 8.6 | 119.7 | 120.1 | 0.35 | 0.35 | -0.14 | 0.04 | 0.7 | -0.15 | -0.03 | 2.025 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | EN | ML | GTR2_20230619_R15M10_01, GTR2_20230619_R31M10_01 | After LLRF calibration/ML | 90.6 | 100 | 92.8 | 100 | 2023-06-19 | 1.0091 | 0.99578 | 0.99522 | 1.009 | 0.99656 | 0.99574 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230620 | 1.976 | -0.5536% | -1.397% | 1.995 | -0.5979% | -1.53% | 0.99048 | 0.62 | 0.75 | -0.25 | 0.22 | 1.08 | -0.26 | 0.18 | 1.976 | 9.5 | 8.4 | 8 | 8.7 | 119.7 | 120.1 | 0.27 | 0.41 | -0.11 | -0.01 | 0.73 | -0.13 | -0.07 | 1.995 | 11.7 | 12 | 11.7 | 11.5 | 119.4 | 119.8 | GK | GB | GTR2_20230620_R15M10_01, GTR2_20230620_R31M10_01 |  | 90 | 100 | 90.9 | 100 | 2023-06-20 | 0.99446 | 0.99595 | 0.99519 | 0.99402 | 0.99674 | 0.99583 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230621 | 1.98 | -0.3523% | -1.198% | 2 | -0.3488% | -1.283% | 0.99 | 0.49 | 0.58 | -0.22 | 0.27 | 1.02 | -0.26 | 0.2 | 1.98 | 8.7 | 9.2 | 7.9 | 8.6 | 119.8 | 120.1 | 0.26 | 0.37 | -0.09 | 0.01 | 0.67 | -0.1 | -0.04 | 2 | 11.5 | 11.5 | 11.4 | 11.4 | 119.4 | 119.8 | OJ | CVD | GTR2_20230621_R15M10_01, GTR2_20230621_R31M10_01 |  | 100 |  | 100 |  | 2023-06-21 | 0.99648 | 0.99597 | 0.99529 | 0.99651 | 0.99678 | 0.99595 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230622 | 1.981 | -0.302% | -1.148% | 2 | -0.3488% | -1.283% | 0.9905 | 0.47 | 0.63 | -0.28 | 0.24 | 1.08 | -0.29 | 0.14 | 1.981 | 9 | 8.9 | 8 | 8.4 | 119.8 | 120.1 | 0.36 | 0.31 | -0.17 | 0.02 | 0.65 | -0.15 | -0.04 | 2 | 11.5 | 11.5 | 11.3 | 11.3 | 119.4 | 119.8 | ON | ML | GTR2_20230622_R15M10_01, GTR2_20230622_R31M10_01 |  | 100 |  | 100 |  | 2023-06-22 | 0.99698 | 0.99595 | 0.99539 | 0.99651 | 0.99671 | 0.99606 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230626 | 1.985 | -0.1007% | -0.9481% | 2.01 | 0.1495% | -0.7897% | 0.98756 | 0.57 | 0.49 | -0.28 | 0.1 | 0.96 | -0.3 | 0.08 | 1.985 | 8.4 | 9.4 | 8 | 8.5 | 119.7 | 120.1 | 0.27 | 0.38 | -0.09 | -0.02 | 0.68 | -0.11 | -0.04 | 2.01 | 11.7 | 11.7 | 11.7 | 11.5 | 119.3 | 119.8 | AR | ML | GTR2_20230626_R15M10_01, GTR2_20230626_R31M10_01 |  | 100 |  | 100 |  | 2023-06-26 | 0.99899 | 0.99586 | 0.99549 | 1.0015 | 0.99661 | 0.99618 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230627 | 1.983 | -0.2013% | -1.048% | 2.005 | -0.09965% | -1.037% | 0.98903 | 0.73 | 0.61 | -0.38 | 0.25 | 1.01 | -0.33 | 0.21 | 1.983 | 8.5 | 9.3 | 7.8 | 8.6 | 119.7 | 120.1 | 0.29 | 0.42 | -0.1 | -0.04 | 0.68 | -0.11 | -0.08 | 2.005 | 11.6 | 11.6 | 11.6 | 11.5 | 119.4 | 119.8 | EN | AC | GTR2_20230627_R15M10_01, GTR2_20230627_R31M10_01 |  | 100 |  | 100 |  | 2023-06-27 | 0.99799 | 0.99575 | 0.99561 | 0.999 | 0.99653 | 0.99631 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230628 | 1.981 | -0.302% | -1.148% | 2.002 | -0.2491% | -1.185% | 0.98951 | 0.69 | 0.58 | -0.34 | 0.23 | 1.08 | -0.31 | 0.19 | 1.981 | 9.2 | 8.8 | 7.8 | 8.6 | 119.7 | 120.1 | 0.37 | 0.39 | -0.16 | 0.06 | 0.71 | -0.17 | 0 | 2.002 | 11.6 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | AR | AD | GTR2_20230628_R15M10_01, GTR2_20230628_R31M10_01 |  | 99.9 | 100 | 100 |  | 2023-06-28 | 0.99698 | 0.99579 | 0.99565 | 0.99751 | 0.99659 | 0.99639 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230629 | 1.983 | -0.2013% | -1.048% | 2 | -0.3488% | -1.283% | 0.9915 | 0.61 | 0.61 | -0.29 | 0.15 | 0.96 | -0.29 | 0.14 | 1.983 | 8.5 | 9.4 | 8.1 | 8.4 | 119.7 | 120.1 | 0.27 | 0.34 | -0.1 | 0.02 | 0.63 | -0.13 | -0.05 | 2 | 11.7 | 11.6 | 11.4 | 11.3 | 119.4 | 119.9 | ON | ML | GTR2_20230629_R15M10_01, GTR2_20230629_R31M10_01 |  | 100 |  | 100 |  | 2023-06-29 | 0.99799 | 0.9958 | 0.99576 | 0.99651 | 0.99664 | 0.99651 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230630 | 1.978 | -0.4529% | -1.297% | 1.998 | -0.4484% | -1.382% | 0.98999 | 0.58 | 0.58 | -0.32 | 0.26 | 1.08 | -0.3 | 0.22 | 1.978 | 8.7 | 9.2 | 7.9 | 8.6 | 119.7 | 120.1 | 0.31 | 0.37 | -0.12 | 0 | 0.65 | -0.13 | -0.05 | 1.998 | 11.5 | 11.5 | 11.5 | 11.4 | 119.4 | 119.8 | EN | ML | GTR2_20230630_R15M10_01, GTR2_20230630_R31M10_01 |  | 100 |  | 100 |  | 2023-06-30 | 0.99547 | 0.99576 | 0.99578 | 0.99552 | 0.99658 | 0.99656 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230703 | 1.976 | -0.5536% | -1.397% | 1.997 | -0.4983% | -1.431% | 0.98948 | 0.59 | 0.56 | -0.31 | 0.24 | 1.06 | -0.28 | 0.21 | 1.976 | 8.8 | 9.2 | 7.8 | 8.5 | 119.7 | 120.1 | 0.28 | 0.39 | -0.07 | -0.02 | 0.69 | -0.1 | -0.07 | 1.997 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | AR | AC | GTR2_20230703_R15M10_01, GTR2_20230703_R31M10_01 |  | 100 |  | 100 |  | 2023-07-03 | 0.99446 | 0.99581 | 0.99595 | 0.99502 | 0.99661 | 0.99674 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230704 | 1.975 | -0.6039% | -1.447% | 1.998 | -0.4484% | -1.382% | 0.98849 | 0.53 | 0.55 | -0.27 | 0.13 | 1.09 | -0.27 | 0.13 | 1.975 | 8.8 | 9.2 | 8 | 8.5 | 119.7 | 120.1 | 0.34 | 0.38 | -0.12 | -0.01 | 0.75 | -0.12 | -0.05 | 1.998 | 11.5 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | EN | AC | GTR2_20230704_R15M10_01, GTR2_20230704_R31M10_01 |  | 100 |  | 100 |  | 2023-07-04 | 0.99396 | 0.99587 | 0.99597 | 0.99552 | 0.99665 | 0.99678 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230705 | 1.974 | -0.6543% | -1.497% | 1.996 | -0.5481% | -1.481% | 0.98898 | 0.64 | 0.59 | -0.3 | 0.24 | 1.08 | -0.28 | 0.23 | 1.974 | 8.8 | 9.2 | 7.7 | 8.7 | 119.7 | 120.1 | 0.28 | 0.35 | -0.12 | -0.01 | 0.69 | -0.12 | -0.05 | 1.996 | 11.7 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | AR | AD | GTR2_20230705_R15M10_01, GTR2_20230705_R31M10_01 |  | 100 |  | 100 |  | 2023-07-05 | 0.99346 | 0.9954 | 0.99595 | 0.99452 | 0.99621 | 0.99671 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230706 | 1.969 | -0.9059% | -1.747% | 1.996 | -0.5481% | -1.481% | 0.98647 | 0.53 | 0.62 | -0.28 | 0.25 | 1.02 | -0.27 | 0.21 | 1.969 | 9 | 9 | 7.9 | 8.6 | 119.8 | 120.1 | 0.34 | 0.38 | -0.11 | 0.01 | 0.73 | -0.13 | -0.03 | 1.996 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | OJ | AD | GTR2_20230706_R15M10_01, GTR2_20230706_R31M10_01 |  | 95.3 | 100 | 100 |  | 2023-07-06 | 0.99094 | 0.99543 | 0.99586 | 0.99452 | 0.99629 | 0.99661 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230707 | 1.973 | -0.7046% | -1.547% | 1.993 | -0.6976% | -1.629% | 0.98996 | 0.46 | 0.66 | -0.28 | 0.14 | 1.03 | -0.28 | 0.15 | 1.973 | 8.7 | 9.2 | 8.1 | 8.4 | 119.8 | 120.1 | 0.35 | 0.33 | -0.15 | 0.02 | 0.72 | -0.17 | -0.03 | 1.993 | 11.5 | 11.6 | 11.7 | 11.4 | 119.4 | 119.8 | AJ, OJ | ML | GTR2_20230707_R15M10_01, GTR2_20230707_R31M10_01 |  | 99.3 |  | 100 |  | 2023-07-07 | 0.99295 | 0.99539 | 0.99575 | 0.99302 | 0.99628 | 0.99653 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230710 | 1.974 | -0.6543% | -1.497% | 1.994 | -0.6477% | -1.579% | 0.98997 | 0.44 | 0.56 | -0.28 | 0.17 | 1.11 | -0.29 | 0.17 | 1.974 | 8.7 | 9.2 | 8.1 | 8.5 | 119.8 | 120.1 | 0.31 | 0.42 | -0.14 | -0.01 | 0.74 | -0.17 | -0.04 | 1.994 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | EN | AC | GTR2_20230710_R15M10_01, GTR2_20230710_R31M10_01 |  | 100 |  | 100 |  | 2023-07-10 | 0.99346 | 0.99533 | 0.99579 | 0.99352 | 0.99627 | 0.99659 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230711 | 1.973 | -0.7046% | -1.547% | 1.993 | -0.6976% | -1.629% | 0.98996 | 0.48 | 0.59 | -0.29 | 0.13 | 1.05 | -0.28 | 0.13 | 1.973 | 8.6 | 9.2 | 8.4 | 8.1 | 119.8 | 120.2 | 0.26 | 0.41 | -0.11 | -0.06 | 0.71 | -0.12 | -0.07 | 1.993 | 11.6 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | OJ | YM | GTR2_20230711_R15M10_01, GTR2_20230711_R31M10_01 |  | 100 |  | 100 |  | 2023-07-11 | 0.99295 | 0.99518 | 0.9958 | 0.99302 | 0.99606 | 0.99664 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230712 | 1.977 | -0.5033% | -1.347% | 1.999 | -0.3986% | -1.333% | 0.98899 | 0.43 | 0.56 | -0.27 | 0.16 | 1.05 | -0.28 | 0.17 | 1.977 | 8.8 | 9.1 | 8 | 8.3 | 119.8 | 120.2 | 0.28 | 0.36 | -0.07 | -0.02 | 0.68 | -0.11 | -0.07 | 1.999 | 11.5 | 11.5 | 11.6 | 11.4 | 119.4 | 119.9 | ON | YM | GTR2_20230712_R15M10_01, GTR2_20230712_R31M10_01 |  | 100 |  | 100 |  | 2023-07-12 | 0.99497 | 0.99505 | 0.99576 | 0.99601 | 0.99593 | 0.99658 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230713 | 1.975 | -0.6039% | -1.447% | 1.998 | -0.4484% | -1.382% | 0.98849 | 0.53 | 0.62 | -0.3 | 0.2 | 1.05 | -0.31 | 0.18 | 1.975 | 8.4 | 9.4 | 8.1 | 8.3 | 119.7 | 120.2 | 0.28 | 0.38 | -0.06 | 0 | 0.7 | -0.08 | -0.04 | 1.998 | 11.7 | 11.6 | 11.5 | 11.4 | 119.4 | 119.8 | ON | ME | GTR2_20230713_R15M10_01, GTR2_20230713_R31M10_01 |  | 100 |  | 100 |  | 2023-07-13 | 0.99396 | 0.99497 | 0.99581 | 0.99552 | 0.99586 | 0.99661 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230714 | 1.991 | 0.2013% | -0.6487% | 2.015 | 0.3986% | -0.5429% | 0.98809 | 0.51 | 0.53 | -0.3 | 0.13 | 1.05 | -0.28 | 0.15 | 1.991 | 8.7 | 9.2 | 7.9 | 8.4 | 119.7 | 120.2 | 0.36 | 0.39 | -0.16 | -0.06 | 0.73 | -0.13 | -0.06 | 2.015 | 11.3 | 11.4 | 11.4 | 11.3 | 119.4 | 119.9 | ON | YM | GTR2_20230714_R15M10_01, GTR2_20230714_R31M10_01 |  | 100 |  | 100 |  | 2023-07-14 | 1.002 | 0.99482 | 0.99587 | 1.004 | 0.99582 | 0.99665 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230717 | 1.992 | 0.2516% | -0.5988% | 2.01 | 0.1495% | -0.7897% | 0.99104 | 0.6 | 0.61 | -0.32 | 0.12 | 0.99 | -0.29 | 0.13 | 1.992 | 8.4 | 9.5 | 8.2 | 8.2 | 119.8 | 120.2 | 0.38 | 0.39 | -0.14 | -0.02 | 0.71 | -0.13 | -0.01 | 2.01 | 12 | 11.8 | 11.6 | 11.4 | 119.4 | 119.9 | ON | YM | GTR2_20230717_R15M10_01, GTR2_20230717_R31M10_01 |  | 100 |  | 100 |  | 2023-07-17 | 1.0025 | 0.99479 | 0.9954 | 1.0015 | 0.99584 | 0.99621 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230718 | 1.969 | -0.9059% | -1.747% | 1.994 | -0.6477% | -1.579% | 0.98746 | 0.68 | 0.61 | -0.36 | 0.08 | 1.12 | -0.3 | 0.07 | 1.969 | 8.5 | 9.4 | 8.3 | 8.3 | 119.7 | 120.1 | 0.32 | 0.4 | -0.09 | -0.03 | 0.68 | -0.11 | -0.04 | 1.994 | 11.6 | 11.6 | 11.5 | 11.4 | 119.4 | 119.8 | AC | AC | GTR2_20230718_R15M10_01, GTR2_20230718_R31M10_01 |  | 88.3 | 100 | 100 |  | 2023-07-18 | 0.99094 | 0.99481 | 0.99543 | 0.99352 | 0.99588 | 0.99629 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230719 | 1.98 | -0.3523% | -1.198% | 2.003 | -0.1993% | -1.135% | 0.98852 | 0.47 | 0.72 | -0.15 | 0.28 | 1.05 | -0.22 | 0.24 | 1.98 | 9.4 | 8.4 | 7.7 | 8.6 | 119.7 | 120.1 | 0.36 | 0.42 | -0.15 | -0.05 | 0.74 | -0.14 | -0.05 | 2.003 | 11.6 | 11.6 | 11.7 | 11.5 | 119.4 | 119.8 | ON | AC,YM | GTR2_20230719_R15M10_01, GTR2_20230719_R31M10_01 |  | 90 | 100 | 100 |  | 2023-07-19 | 0.99648 | 0.99486 | 0.99539 | 0.99801 | 0.9959 | 0.99628 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230720 | 1.978 | -0.4529% | -1.297% | 2 | -0.3488% | -1.283% | 0.989 | 0.51 | 0.63 | -0.17 | 0.25 | 1.13 | -0.24 | 0.25 | 1.978 | 9.2 | 8.6 | 7.6 | 8.8 | 119.8 | 120.1 | 0.31 | 0.37 | -0.13 | -0.02 | 0.69 | -0.14 | -0.07 | 2 | 11.6 | 11.6 | 11.8 | 11.5 | 119.4 | 119.8 | EN | YM | GTR2_20230720_R15M10_01, GTR2_20230720_R31M10_01 |  | 93.7 |  | 99.1 |  | 2023-07-20 | 0.99547 | 0.99494 | 0.99533 | 0.99651 | 0.99598 | 0.99627 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230721 | 1.986 | -0.05033% | -0.8982% | 2.006 | -0.04983% | -0.9872% | 0.99003 | 0.42 | 0.67 | -0.27 | 0.18 | 1.05 | -0.27 | 0.19 | 1.986 | 8.8 | 9.2 | 7.9 | 8.5 | 119.7 | 120.1 | 0.33 | 0.34 | -0.13 | 0.05 | 0.72 | -0.14 | -0.01 | 2.006 | 11.5 | 11.6 | 11.6 | 11.4 | 119.4 | 119.8 | AR | YM, NIC | GTR2_20230721_R15M10_01, GTR2_20230721_R31M10_01 |  | 100 |  | 100 |  | 2023-07-21 | 0.9995 | 0.99519 | 0.99518 | 0.9995 | 0.99608 | 0.99606 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230724 | 1.974 | -0.6543% | -1.497% | 1.999 | -0.3986% | -1.333% | 0.98749 | 0.41 | 0.51 | -0.26 | 0.11 | 0.98 | -0.27 | 0.13 | 1.974 | 8.7 | 9.2 | 8.2 | 8.2 | 119.7 | 120.2 | 0.3 | 0.35 | -0.13 | 0.01 | 0.72 | -0.15 | -0.04 | 1.999 | 11.4 | 11.5 | 11.6 | 11.4 | 119.4 | 119.8 | EN | YM KW | GTR2_20230724_R15M10_01, GTR2_20230724_R31M10_01 |  | 100 |  | 100 |  | 2023-07-24 | 0.99346 | 0.99534 | 0.99505 | 0.99601 | 0.99628 | 0.99593 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230725 | 1.971 | -0.8052% | -1.647% | 1.986 | -1.046% | -1.974% | 0.99245 | 0.66 | 0.67 | -0.32 | 0.26 | 1.11 | -0.29 | 0.24 | 1.971 | 8.5 | 9.3 | 7.9 | 8.7 | 119.7 | 120.1 | 0.38 | 0.4 | -0.16 | -0.01 | 0.74 | -0.14 | -0.01 | 1.986 | 11.8 | 11.7 | 11.9 | 11.5 | 119.4 | 119.9 | EN | YM KW | GTR2_20230725_R15M10_01, GTR2_20230725_R31M10_01 |  | 94.3 |  | 67.5 | 100 | 2023-07-25 | 0.99195 | 0.99547 | 0.99497 | 0.98954 | 0.99648 | 0.99586 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230726 | 1.976 | -0.5536% | -1.397% | 2.001 | -0.299% | -1.234% | 0.98751 | 0.44 | 0.58 | -0.26 | 0.2 | 1.13 | -0.29 | 0.16 | 1.976 | 9 | 8.9 | 8.2 | 8.2 | 119.8 | 120.2 | 0.28 | 0.4 | -0.07 | -0.03 | 0.73 | -0.11 | -0.06 | 2.001 | 11.5 | 11.5 | 11.7 | 11.4 | 119.4 | 119.9 | AR | YM | GTR2_20230726_R15M10_01, GTR2_20230726_R31M10_01 |  | 100 |  | 100 |  | 2023-07-26 | 0.99446 | 0.99566 | 0.99482 | 0.99701 | 0.99674 | 0.99582 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230727 | 1.973 | -0.7046% | -1.547% | 1.997 | -0.4983% | -1.431% | 0.98798 | 0.48 | 0.62 | -0.28 | 0.17 | 1.14 | -0.25 | 0.18 | 1.973 | 8.8 | 9.2 | 8 | 8.4 | 119.8 | 120.2 | 0.36 | 0.38 | -0.15 | 0.02 | 0.73 | -0.14 | -0.02 | 1.997 | 11.6 | 11.6 | 11.9 | 11.5 | 119.4 | 119.8 | ON | YM KW | GTR2_20230727_R15M10_01, GTR2_20230727_R31M10_01 |  | 99 |  | 100 |  | 2023-07-27 | 0.99295 | 0.99572 | 0.99479 | 0.99502 | 0.9968 | 0.99584 | 1 | 1.02 | 0.98 | 1.03 | 0.97
20230728 | 1.977 | -0.5033% | -1.347% | 1.998 | -0.4484% | -1.382% | 0.98949 | 0.44 | 0.63 | -0.24 | 0.21 | 1.04 | -0.27 | 0.17 | 1.977 | 8.6 | 9.3 | 8 | 8.3 | 119.8 | 120.2 | 0.2 | 0.35 | -0.04 | 0.01 | 0.63 | -0.05 | -0.06 | 1.998 | 11.8 | 11.7 | 11.8 | 11.5 | 119.4 | 119.8 | ON | YM MCN | GTR2_20230728_R15M10_01, GTR2_20230728_R31M10_01 |  | 100 |  | 100 |  | 2023-07-28 | 0.99497 | 0.99588 | 0.99481 | 0.99552 | 0.99692 | 0.99588 | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
[44 empty] | #VALUE! | [6 empty] | 1 | 1.02 | 0.98 | 1.03 | 0.97
Formulas:
Plot = DATE(LEFT([Plan ID],4), MID([Plan ID],5,2), RIGHT([Plan ID],2))

## Safety and lasers
|  | Ljusgardin: | Mån, ons och fredag |  |  |  | Lasers: | distance in mm between the middle of the laser line and the marker on the wall (0.5 mm precision)
|  | Dörrinterlock: | Tis och torsdag
|  | Avsyningsprocedur: | Dagligen
|  | Nödstopp: | Första fredagen i månaden
Datum |  | G2
|  | Safety |  |  |  |  | Lasers [mm] |  |  |  | Dosimetry Manager |  | All
|  | Avsyning | varningslampor | ljusgardin | dörrfunktion | RS | Red1 | Green1 | Red2 | Green2 | T [K] | p [hPa] | Kommentar
2023-02-24 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 993.8
2023-02-27 | Måndag | ok | ok | ok |  | ok | 0 | 0 | 0 | 1 | 296.6 | 1019.9
2023-02-28 | Tisdag | ok | ok |  | ok | ok | 0 | 0 | -0.5 | 1 | 296.8 | 1021.6
2023-03-01 | Onsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1022
2023-03-02 | Torsdag | ok | ok |  | ok | ok | -0.5 | 0 | 0 | 1 | 296.6 | 1017.7
2023-03-03 | Fredag | ok | ok | ok |  | ok | -0.5 | 0 | 0 | 1 | 296.8 | 1009.2
2023-03-06 | Måndag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.5 | 993.1
2023-03-07 | Tisdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.6 | 990.7
2023-03-08 | Onsdag | ok | ok | ok |  | ok | 0 | -0.5 | -0.5 | 1 | 296.6 | 979.6
2023-03-09 | Torsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.9 | 992.3
2023-03-10 | Fredag | ok | ok | ok |  | ok | 0 | 0 | 0 | 1 | 295.8 | 994.5
2023-03-13 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.4 | 993.8
2023-03-14 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 297.3 | 968
2023-03-15 | Onsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 297.3 | 989.5
2023-03-16 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1010 | ok//ML
| Fredag | [11 empty] | Cyclo PM, no treatments
2023-03-20 | Måndag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1001.7 | OK/AD
2023-03-21 | Tisdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1007.3 | OK/GB
2023-03-22 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 297.1 | 996 | ok/CVD
2023-03-23 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.9 | 988.2 | ok/ML
2023-03-24 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 988.6 | KW
2023-03-27 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 997.7 | ok/AC
2023-03-28 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 297.3 | 1005 | ok / YM
2023-03-29 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.25 | 296.8 | 1012.1 | ok / YM ON
2023-03-30 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 1009.1 | ok /ON
2023-03-31 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1001.5 | ok/GB
2023-04-03 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 295.8 | 1032.1 | Ok /YM
2023-04-04 | Tisdag | ok | ok | ok |  | ok | 0 | 0 | -0.5 | 1 | 296.9 | 1024.6 | OK/GB
2023-04-05 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.8 | 1021.3 | OK/GB
2023-04-06 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | -0.5 | 1 | 296.8 | 1022.3 | ok/CVD
2023-04-07 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | -0.5 | 1 | 296.8 | 1023.2 | ok/CVD
Måndag
2023-04-11 | Tisdag | ok | ok |  | ok | ok | 0 | -0.5 | -0.5 | 0.5 | 296.4 | 1010.5 | OK/GB
2023-04-12 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.8 | 1005.5 | ok /YM
2023-04-13 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.9 | 1005.8 | ok/YM
2023-04-14 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1014 | ok/GB
2023-04-17 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.6 | 1026 | OK/GB
2023-04-18 | Tisdag | ok | ok | ok |  | ok | 0.5 | 0 | 0 | 0.5 | 296.8 | 1032.4 | ok/GB
2023-04-19 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1033.2 | ok/CVD
2023-04-20 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0.5 | 1 | 296.4 | 1029.3 | ok/CVD
2023-04-21 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1022.1 | ok/CVD
2023-04-24 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 999.9 | ok/YM
2023-04-25 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 995.9 | ok/CVD
2023-04-26 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 995.4 | ok/YM
2023-04-27 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.9 | 1000.7
2023-04-28 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1007.7 | ok/AC
| Måndag | [11 empty] | Holiday
2023-05-02 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.1 | 1000.2 | ok/CVD
2023-05-03 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1015.2
2023-05-04 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 1024.3 | ok/YM
2023-05-05 | Fredag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.4 | 1027.9 | ok/AC
2023-05-08 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1025.3 | ok/AC
2023-05-09 | Tisdag | ok | ok | ok |  | ok | 0 | 0.5 | 0 | -0.5
2023-05-10 | Onsdag | ok | ok |  | ok | ok | 0 | 0.5 | 0 | -0.5 | 296.6 | 1015.8 | ok/GB
2023-05-11 | Torsdag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.4 | 1017.8 | ok/CVD
2023-05-12 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1021.4 | ok/EAL
2023-05-15 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1013.5 | OK/GB
2023-05-16 | Tisdag | ok | ok | ok |  | ok | 0 | -1 | 0 | 1 | 296.9 | 1003.4 | OK/ON, YM
2023-05-17 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.35 | 997.7 | ok / YM ON
| Torsdag | [11 empty] | Holiday
2023-05-19 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 1025.2 | AD
2023-05-22 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.55 | 1020.8 | GB
2023-05-23 | Tisdag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.9 | 1010.6 | CVD
2023-05-24 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 0.5 | 297.3 | 1012.5 | AD
2023-05-25 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 297.1 | 1009.8 | GB
2023-05-26 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1010.6 | AC
2023-05-29 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.4 | 1016.4 | CVD
2023-05-30 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.35 | 1018.5 | ok/YM
2023-05-31 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 1011.5 | ok/YM
2023-06-01 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1012.2 | AD
2023-06-02 | Fredag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.9 | 1015.9 | AC
2023-06-05 | Måndag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.6 | 1016.5 | AC,KW
2023-06-06 | Tisdag | [11 empty] | Holiday
2023-06-07 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 1 | 296.9 | 1018.4
2023-06-08 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.1 | 1013.7 | GB
2023-06-09 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1021.5 | AC
2023-06-12 | Måndag | ok | ok |  | ok | ok | 0 | 0 | 0 | 0.5 | 296.6 | 1023.4 | AC
2023-06-13 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.9 | 1018 | ok /GB
2023-06-14 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 297.05 | 1017.8 | ok / YM
2023-06-15 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.85 | 1017 | ok / YM
2023-06-16 | Fredag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.4 | 1013.7 | ok/AC
2023-06-19 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 1012.5
2023-06-20 | Tisdag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0 | 296.6 | 1012.8 | GB
2023-06-21 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 0 | 296.8 | 1011.9 | CVD
2023-06-22 | Torsdag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0 | 296.9 | 1006.1 | ML
2023-06-23 | Fredag | [11 empty] | Holiday
2023-06-26 | Måndag | ok | okk |  | ok | ok | 0 | 0 | 0 | 0 | 296.8 | 1011.7 | ML
2023-06-27 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1003.9 | ok/AC
2023-06-28 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 0 | 296.9 | 1007.5 | AD
2023-06-29 | Torsdag | ok | ok | ok |  | ok | 0 | -1 | 0 | 0.5 | 296.9 | 1004.8 | ML
2023-06-30 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 1001.1 | ML
2023-07-03 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 984.8 | AC
2023-07-04 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.1 | 991.5 | AC
2023-07-05 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.4 | 1000.6 | AD
2023-07-06 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 1005.5 | AD
2023-07-07 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.8 | 1007.9 | ML
2023-07-10 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 1 | 296.55 | 1017.8 | AC
2023-07-11 | Tisdag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.3 | 1012.9 | ok / YM
2023-07-12 | Onsdag | ok | ok |  | ok | ok | 0 | 0 | 0 | 0.5 | 296.1 | 1008.1 | ok / YM
2023-07-13 | Torsdag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.3 | 1000.6 | ME
2023-07-14 | Fredag | ok | ok | ok |  | ok | 0 | 0 | 0 | 0.5 | 296.6 | 1003.8 | ok/YM
2023-07-17 | Måndag | ok | ok |  | ok | ok | 0 | 0 | 0 | 0.5 | 296.3 | 1003.5 | ok/YM
2023-07-18 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 1 | 296.1 | 999.6 | ok/AC
2023-07-19 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 999.9 | ok/AC,YM
2023-07-20 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.4 | 997 | ok/YM
2023-07-21 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 997.4 | ok/YM, nic
2023-07-24 | Måndag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 998.4 | ok/YM KW
2023-07-25 | Tisdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 990.7 | ok/YM KW
2023-07-26 | Onsdag | ok | ok |  | ok | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 993.8 | ok/YM
2023-07-27 | Torsdag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.1 | 998.7 | ok/YM KW
2023-07-28 | Fredag | ok | ok | ok |  | ok | 0 | -0.5 | 0 | 0.5 | 296.6 | 997.4 | ok/ YM MCN
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag
Måndag
Tisdag
Onsdag
Torsdag
Fredag

## Signatures
Signature | First name | Last name | Profession | Note
AA | Ali | Alkhiat | Medical Physicist
AC | Athanasia | Christou | Summer physicist
AD | Alexandru | Dasu | Medical physicist
AE | Albin | Eriksson | Uppstuk physicist
AHL | Annika | Hall | Medical physicist
AJ | Anton | Jansson | Uppstuk physicist
AR | Adam | Röjvall | Uppstuk physicist
CVD | Christina | Vallhagen Dahlgren | Medical Physicist
DJ | Dan | Josefsson | Medical physicist
EA | Erik | Almhagen | Medical physicist
EF | Emil | Forsberg | Uppstuk physicist
EL | Elias | Lindbäck | Medical physicist
EN | Erik | Näslund | Uppstuk physicist
EPe | Erik | Pettersson | Medical physicist
GB | Gloria | Bäckström | Medical physicist
GF | Gustav | Furubjelke | Uppstuk physicist
JM | Joakim | Medin | Medical physicist
JR | Johan | Rensfeldt | Uppstuk physicist
GK | Gurpreet | Kaur | Uppstuk physicist
KW | Kenneth | Wikström | Medical physicist
LL | Linnea | Lund | Medical physicist
MB | Mikael | Blomqvist | Medical physicist
ME | Marika | Enmark | Medical physicist
MF | Marcus | Fager | Medical physicist
MJ | Mikael | Johansson | Uppstuk physicist
MK | Marcus | Krantz | Medical physicist
Ml | Malgorzata | Liszka | Medical physicist
NIC | Carl Magnus | Nilsson | Medical physicist
OJ | Olof | Jardfelt | Uppstuk physicist
ON | Olivier | Näslund | Uppstuk physicist
PL | Peter | Larsson | Medical physicist
SF | Stina | Fredriksson | Uppstuk physicist
SM | Sebastian | Mikkelsen | Uppstuk physicist
TH | Thomas | Henry | Medical physicist
UG | Ulf | Granlund | Medical physicist
YM | Younes | Majeddam | Medical physicist

## Updates
No | Date | Responsible | Change Description
1 | 2023-02-23 | ML | The document released for users
2
3
4
5
6