## GTR1_SW_SOBP plans
ID.nr:
Titel: | [11 empty] | Dokumenttyp:
Protocol for measurements of the homogeneous cubic plan in SW with Roos IC | [11 empty] | Strålskyddsinstruktion
Godkänt av: | [10 empty] | Godkänt den:
Kategori: |  | Strålskydd
Skapat av: | [10 empty] | Skapat den:
Reviderat av: | [10 empty] | Reviderat den:
Room ID: | GTR 1
Dosimetry equipment: | Name and model:
phantom | Slab phantom
ionization chamber | PTW Roos TW34001 |  | SN:002519
electrometer | PTW UnidosWebline T10022 |  | SN:000456
cable |  |  | reference
Setup of measurement: | Mode |  | Charge&Current
| Units |  | Electrical
| Range |  | Medium/High
| Nominal voltage [V] |  | -200
calibration factor (proton PBS, Qo=150MeV): | ND,w,Qo [Gy/nC] 2021-03-19 |  | 0.0846
| ND,w,Qo [Gy/nC] 2023-03-27 |  | 0.0849 | 0.3546%
Plan ID | R15M10 | R31M10
Prescribe dose TPS [Gy] | 2.018 | 2.019 |  | -0.9911% | -1.436%
Refernece dose [Gy] | 1.998 | 1.99 |  | -0.3504% | -0.3518%
Prescribe MU's | 2069.67 | 1719.77
zmeas [cm] | 10 | 26
Correction factors
electrometer - kelec | 1 | 1
ion recombination - ks | 1.0057 | 1.004
polarity - kpol | 1 | 1
beam quality - kQ,Qo | 1.0006 | 1.0005
iso at the surface of the IC
Date | Plan ID | zmeas | no.of SW plates | PPS Z-axis | TIC,air | pair | kT,p_air | Qmeas [nC] |  |  |  |  | Average Qmeas | STD | Average Dose | Deviation from reference | Deviation from TPS | Primary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's | Secondary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's
[DD/MM/YYYY] |  | [cm] | [cm] | [cm] | [oC] | [hPa] | [oC] | meas 1 | meas 2 | meas 3 | meas 4 | meas 5 | [nC] | [%] | [Gy] | [%] | [%] | MU_1 | MU_2 | MU_3 | MU_4 | MU_5 |  |  | MU_1 | MU_2 | MU_3 | MU_4 | MU_5
2021-07-14 | R15M10 | 10 | 9.6 |  | 22.59 | 1009 | 1.0131 | -23.12 | -23.12 | -23.12 |  |  | -23.12 | 0% | 1.9939 | -0.2028% | -1.192% | 2068.3 | 2068.52 | 2068.53 |  |  | 2068.45 | -0.05895% | 2071.3 | 2071.44 | 2071.44 |  |  | 2071.3933 | 0.08327% | GTR1_Roos IC+electrometer #470 from GTR2 | REFERENCE MEASUREMENTS
| R31M10 | 26 | 25.3 |  | 22.59 | 1009 | 1.0131 | -23.08 | -23.1 | -23.08 |  |  | -23.0867 | -0.05002% | 1.9876 | -0.1199% | -1.555% |  | 1720.47 | 1720.46 |  |  | 1720.465 | 0.04041% |  | 1722.15 | 1722.18 |  |  | 1722.165 | 0.1393%
2021-07-19 | R15M10 | 10 | 9.6 |  | 22.68 | 1012 | 1.0104 | -23.21 | -23.24 | -23.22 |  |  | -23.2233 | -0.06578% | 1.9975 | -0.02356% | -1.014% |  |  | 2069.15 |  |  | 2069.15 | -0.02512% |  |  | 2070.34 |  |  | 2070.34 | 0.03237%
| R31M10 | 26 | 25.3 |  | 22.68 | 1012 | 1.0104 | -23.16 | -23.17 | -23.17 |  |  | -23.1667 | -0.02492% | 1.9892 | -0.0405% | -1.476% | 1721.1 |  |  |  |  | 1721.1 | 0.07734% | 1721.3 |  |  |  |  | 1721.3 | 0.08897%
2021-09-14 | R15M10 | 10 | 9.6 | -2.12 | 22.08 | 1014 | 1.0064 | -23.32 | -23.34 | -23.34 |  |  | -23.3333 | -0.04949% | 1.999 | 0.04854% | -0.943% | 2067.36 | 2069.8 | 2069.83 |  |  | 2068.9967 | -0.03253% | 2068.68 | 2071.02 | 2071.13 |  |  | 2070.2767 | 0.02931% | GTR1_Roos IC+electrometer #470 from GTR2
| R31M10 | 26 | 25.3 | -2.08 | 22.08 | 1014 | 1.0064 | -23.27 | -23.27 | -23.23 | -23.25 | -23.24 | -23.252 | -0.07693% | 1.9885 | -0.07328% | -1.509% | 1721.61 | 1721.61 | 1719.86 | 1719.88 | 1719.71 | 1720.534 | 0.04442% | 1721.82 | 1721.88 | 1720.18 | 1720.12 | 1720 | 1720.8 | 0.05989%
2021-09-17 | R15M10 | 10 | 9.6 | -2.11 | 22.304 | 1017 | 1.0041 | -23.41 | -23.38 | -23.38 |  |  | -23.39 | -0.07405% | 1.9994 | 0.07153% | -0.9203% | 2075.68 | 2070.52 | 2070.68 |  |  | 2072.2933 | 0.1268% | 2077.25 | 2072.22 | 2072.39 |  |  | 2073.9533 | 0.207% | GTR1_Roos IC+electrometer #470 from GTR2
| R31M10 | 26 | 25.3 | -2.05 | 22.304 | 1017 | 1.0041 | -23.38 | -23.38 | -23.37 |  |  | -23.3767 | -0.0247% | 1.9948 | 0.2421% | -1.198% | 1725.02 | 1725.31 | 1725.32 |  |  | 1725.2167 | 0.3167% | 1725.43 | 1725.83 | 1725.84 |  |  | 1725.7 | 0.3448%
2021-10-25 | R15M10 | 10 | 9.6 | -2.05 | 22.03 | 1008 | 1.0122 | -23.27 | -23.27 | -23.28 |  |  | -23.2733 | -0.02481% | 2.0054 | 0.3683% | -0.6265% | 2070.56 | 2070.54 | 2070.66 |  |  | 2070.5867 | 0.04429% | 2072.91 | 2072.9 | 2072.92 |  |  | 2072.91 | 0.1565%
| R31M10 | 26 | 25.3 | -2.05 | 22.03 | 1008 | 1.0122 | -23.21 | -23.21 | -23.21 |  |  | -23.21 | -1.875e-14% | 1.9964 | 0.323% | -1.118% | 1722.1 | 1721.98 | 1721.83 |  |  | 1721.97 | 0.1279% | 1723.29 | 1723.16 | 1722.95 |  |  | 1723.1333 | 0.1956%
2022-10-14 | R15M10 | 10 | 9.6 | -1.95 | 22.41 | 1005.8 | 1.0157 | -23.08 | -23.08 | -23.08 |  |  | -23.08 | 0% | 1.9956 | -0.1194% | -1.109% | 2068.86 | 2069.24 | 2069.07 |  |  | 2069.0567 | -0.02963% | 2080.13 | 2080.52 | 2080.32 |  |  | 2080.3233 | 0.5147%
| R31M10 | 26 | 25.3 | -1.95 | 22.41 | 1005.8 | 1.0157 | -23.01 | -23.01 | -23.01 |  |  | -23.01 | 0% | 1.9861 | -0.1957% | -1.629% | 1720.23 | 1720.32 | 1720.38 |  |  | 1720.31 | 0.0314% | 1728.86 | 1728.89 | 1728.99 |  |  | 1728.9133 | 0.5317%
2022-10-18 | R15M10 | 10 | 9.6 | -1.95 | 22.2 | 1010.3 | 1.0104 | -23.13 | -23.13 |  |  |  | -23.13 | 0% | 1.9896 | -0.4196% | -1.407% | 2069.47 | 2068.78 |  |  |  | 2069.125 | -0.02633% | 2080.3 | 2079.65 |  |  |  | 2079.975 | 0.4979%
| R31M10 | 26 | 25.3 | -1.95 | 22.2 | 1010.3 | 1.0104 | -23.06 | -23.07 |  |  |  | -23.065 | -0.03066% | 1.9806 | -0.4735% | -1.903% | 1720.18 | 1720.23 |  |  |  | 1720.205 | 0.02529% | 1728.51 | 1728.58 |  |  |  | 1728.545 | 0.5102%
2022-11-04 | R15M10 | 10 | 9.6 | 6.84 | 22.76 | 1008.3 | 1.0144 | -23.23 | -23.22 | -23.23 |  |  | -23.2267 | -0.02486% | 2.0057 | 0.3849% | -0.61% | 2070.26 | 2070.81 | 2071 |  |  | 2070.69 | 0.04928% | 2106.51 | 2107.14 | 2107.34 |  |  | 2106.9967 | 1.804%
| R31M10 | 26 | 25.3 | 6.84 | 22.76 | 1008.3 | 1.0144 | -23.17 | -23.16 | -23.16 |  |  | -23.1633 | -0.02493% | 1.9967 | 0.339% | -1.102% | 1721.78 | 1722.08 | 1722.28 |  |  | 1722.0467 | 0.1324% | 1752.78 | 1753.07 | 1753.29 |  |  | 1753.0467 | 1.935%
2023-03-14 | R15M10 | 10 | 9.6 | -2.05 | 23.6 | 980 | 1.0466 | -22.37 | -22.37 | -22.38 |  |  | -22.3733 | -0.02581% | 1.9934 | -0.2284% | -1.217% | 2067.47 | 2067.3 | 2068.23 |  |  | 2067.6667 | -0.09679% | 2081.94 | 2081.65 | 2082.54 |  |  | 2082.0433 | 0.5978%
| R31M10 | 26 | 25.3 | -2.05 | 23.6 | 980 | 1.0466 | -22.3 | -22.28 | -22.29 |  |  | -22.29 | -0.04486% | 1.9826 | -0.3738% | -1.805% | 1720.34 | 1720.15 | 1720.57 |  |  | 1720.3533 | 0.03392% | 1731.56 | 1731.35 | 1731.75 |  |  | 1731.5533 | 0.6852%
2023-04-21 | R15M10 | 10 | 9.7 (Blue SW) | -2.07 | 22.6 | 1016 | 1.0061 | -23.19 | -23.2 | -23.19 |  |  | -23.1933 | -0.02489% | 1.9936 | -0.2201% | -1.209% | 2067.46 | 2067.82 | 2067.78 |  |  | 2067.6867 | -0.09583% | 2079.4 | 2079.67 | 2079.68 |  |  | 2079.5833 | 0.479% | Blue SW phantom + new calibration coefficient
| R31M10 | 26 | 25.5 (Blue SW) | -2.02 | 22.6 | 1016 | 1.0061 | -23.09 | -23.08 | -23.09 |  |  | -23.0867 | -0.02501% | 1.981 | -0.4529% | -1.883% | 1719.71 | 1719.35 | 1719.88 |  |  | 1719.6467 | -0.007172% | 1729.14 | 1728.7 | 1729.17 |  |  | 1729.0033 | 0.5369%
2023-05-30 | R15M10 | 10 | 9.6 | -2.1 | 23.1 | 1017 | 1.0068 | -23.24 | -23.24 | -23.24 |  |  | -23.24 | 0% | 1.999 | 0.05119% | -0.9404% | 2069.24 | 2069.27 | 2069.13 |  |  | 2069.2133 | -0.02206% | 2082.42 | 2082.41 | 2082.36 |  |  | 2082.3967 | 0.6149%
| R31M10 | 26 | 25.3 | -2.1 | 23.1 | 1017 | 1.0068 | -23.14 | -23.14 | -23.15 |  |  | -23.1433 | -0.02495% | 1.9873 | -0.1382% | -1.573% | 1720.42 | 1720.56 | 1720.63 |  |  | 1720.5367 | 0.04458% | 1730.75 | 1730.83 | 1730.91 |  |  | 1730.83 | 0.6431%
2023-06-13 | R15M10 | 10 | 9.6 | -2.1 | 23.64 | 1015 | 1.0107 | -23.16 | -23.15 | -23.16 |  |  | -23.1567 | -0.02493% | 1.9994 | 0.07095% | -0.9208% | 2067.69 | 2067.75 | 2067.46 |  |  | 2067.6333 | -0.09841% | 2080.85 | 2080.83 | 2080.59 |  |  | 2080.7567 | 0.5357%
| R31M10 | 26 | 25.3 | -2.05 | 23.64 | 1015 | 1.0107 | -23.1 | -23.1 | -23.11 |  |  | -23.1033 | -0.02499% | 1.9913 | 0.06772% | -1.37% | 1719.75 | 1719.68 | 1719.78 |  |  | 1719.7367 | -0.001938% | 1729.86 | 1729.77 | 1729.86 |  |  | 1729.83 | 0.585%
2023-06-19 | R15M10 | 10 | 9.6 | -2.18 | 22.9 | 1012 | 1.0111 | -23.22 | -23.21 | -23.24 |  |  | -23.2233 | -0.06578% | 2.0061 | 0.4056% | -0.5895% | 2069.57 | 2069.42 | 2069.55 |  |  | 2069.5133 | -0.00757% | 2080.59 | 2080.38 | 2080.5 |  |  | 2080.49 | 0.5228%
| R31M10 | 26 | 25.3 | -2.18 | 22.9 | 1012 | 1.0111 | -23.1 | -23.1 | -23.11 |  |  | -23.1033 | -0.02499% | 1.9923 | 0.1141% | -1.324% | 1721.44 | 1721.68 | 1721.37 |  |  | 1721.4967 | 0.1004% | 1729.79 | 1729.96 | 1729.7 |  |  | 1729.8167 | 0.5842%
2023-07-18 | R15M10 | 10 | 9.6 | -2.01 | 23.4 | 1001.6 | 1.0234 | -22.9 | -22.91 | -22.89 |  |  | -22.9 | -0.04367% | 2.0021 | 0.2046% | -0.7885% | 2071.55 | 2072.33 | 2071.65 |  |  | 2071.8433 | 0.105% | 2081.24 | 2081.94 | 2081.27 |  |  | 2081.4833 | 0.5708%
| R31M10 | 26 | 25.3 | -1.82 | 23.4 | 1001.6 | 1.0234 | -22.86 | -22.86 | -22.85 |  |  | -22.8567 | -0.02526% | 1.9948 | 0.2427% | -1.197% | 1724.21 | 1724.5 | 1724.17 |  |  | 1724.2933 | 0.263% | 1731.56 | 1731.75 | 1731.41 |  |  | 1731.5733 | 0.6863%
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!

## GTR2_SW_SOBP plans
ID.nr:
Titel: | [11 empty] | Dokumenttyp:
Protocol for measurements of the homogeneous cubic plan in SW with Roos IC | [11 empty] | Strålskyddsinstruktion
Godkänt av: | [10 empty] | Godkänt den:
Kategori: |  | Strålskydd
Skapat av: | [10 empty] | Skapat den:
Reviderat av: | [10 empty] | Reviderat den:
Room ID: | GTR 2
Dosimetry equipment: | Name and model:
phantom | Slab phantom
ionization chamber | PTW Roos TW34001 |  | SN:002519
electrometer | PTW UnidosWebline T10022 |  | SN:000470
cable |  |  | reference
Setup of measurement: | Mode |  | Charge&Current
| Units |  | Electrical
| Range |  | Medium/High
| Nominal voltage [V] |  | -200
calibration factor (proton PBS, Qo=150MeV): | ND,w,Qo [Gy/nC] 2021-03-19 |  | 0.0846
| ND,w,Qo [Gy/nC] 2023-03-27 |  | 0.0849 | 0.3546%
Plan ID | R15M10 | R31M10
Prescribe dose TPS [Gy] | 2.004 | 2.026
Refernece dose [Gy] | 1.985 | 2.006 |  | -0.3023% | -0.349%
Prescribe MU's | 2069.67 | 1719.77
zmeas [cm] | 10 | 26
Correction factors
electrometer - kelec | 1 | 1
ion recombination - ks | 1.0057 | 1.004
polarity - kpol | 1 | 1
beam quality - kQ,Qo | 1.0006 | 1.0005
iso at the surface of the IC
Date | Plan ID | zmeas | no.of SW plates | PPS Z-axis | TIC,air | pair | kT,p_air | Qmeas [nC] |  |  |  |  | Average Qmeas | STD | Average Dose | Deviation from reference | Deviation from TPS | Primary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's | Secondary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's
[DD/MM/YYYY] |  | [cm] | [cm] | [cm] | [oC] | [hPa] | [oC] | meas 1 | meas 2 | meas 3 | meas 4 | meas 5 | [nC] | [%] | [Gy] | [%] | [%] | MU_1 | MU_2 | MU_3 | MU_4 | MU_5 |  |  | MU_1 | MU_2 | MU_3 | MU_4 | MU_5
2021-07-20 | R15M10 | 10 | 9.6 |  | 22.23 | 1006 | 1.0149 | -22.93 | -22.91 | -22.95 |  |  | -22.93 | -0.08722% | 1.981 | -0.1993% | -1.146% | 2074.28 | 2074.4 | 2044.53 |  |  | 2064.4033 | -0.2545% | 2075.79 | 2075.79 | 2075.99 |  |  | 2075.8567 | 0.2989% | REFERENCE MEASUREMENTS
| R31M10 | 26 | 25.3 |  | 22.23 | 1006 | 1.0149 | -23.24 | -23.24 | -23.25 |  |  | -23.2433 | -0.02484% | 2.0046 | -0.06849% | -1.055% | 1725.06 | 1725.12 | 1725.34 |  |  | 1725.1733 | 0.3142% | 1725.39 | 1725.46 | 1725.7 |  |  | 1725.5167 | 0.3342%
2021-10-04 | R15M10 | 10 | 9.6 |  | 23.2 | 1005.5 | 1.0187 | -22.93 | -22.91 |  |  |  | -22.92 | -0.0617% | 1.9877 | 0.1345% | -0.8148% | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  | 23.2 | 1005.5 | 1.0187 | -23.17 | -23.26 |  |  |  | -23.215 | -0.2741% | 2.0098 | 0.1873% | -0.8018% | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
2022-10-14 | R15M10 | 10 | 9.6 | -2.19 | 22.64 | 1005.3 | 1.017 | -22.98 | -22.99 | -22.98 |  |  | -22.9833 | -0.02512% | 1.9898 | 0.2414% | -0.709% | 2072.42 | 2072.49 | 2072.46 |  |  | 2072.4567 | 0.1346% | 2073.37 | 2073.35 | 2073.3 |  |  | 2073.34 | 0.1773%
| R31M10 | 26 | 25.3 | -2.19 | 22.64 | 1005.3 | 1.017 | -23.28 | -23.26 | -23.28 |  |  | -23.2733 | -0.04961% | 2.0114 | 0.2692% | -0.7207% | 1722.66 | 1722.82 | 1722.91 |  |  | 1722.7967 | 0.176% | 1722.65 | 1722.76 | 1722.84 |  |  | 1722.75 | 0.1733%
2022-11-15 | R15M10 | 10 | 9.6 | -2.29 | 22.055 | 1018 | 1.0023 | -23.18 | -23.17 | -23.18 |  |  | -23.1767 | -0.02491% | 1.9776 | -0.3739% | -1.318% | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 | -2.29 | 22.055 | 1018 | 1.0023 | -23.47 | -23.47 | -23.49 |  |  | -23.4767 | -0.04919% | 1.9997 | -0.3142% | -1.298% | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
2022-11-18 | R15M10 | 10 | 9.6 | -2.28 | 22.392 | 1017 | 1.0044 | -23.24 | -23.23 | -23.26 |  |  | -23.2433 | -0.06572% | 1.9875 | 0.1251% | -0.8242% | 2072.5 | 2072.77 | 2072.94 |  |  | 2072.7367 | 0.1482% | 2072.96 | 2073.28 | 2073.43 |  |  | 2073.2233 | 0.1717%
| R31M10 | 26 | 25.3 | -2.28 | 22.392 | 1017 | 1.0044 | -23.5 | -23.51 | -23.51 |  |  | -23.5067 | -0.02456% | 2.0065 | 0.02539% | -0.962% | 1723.04 | 1723.33 | 1723.32 |  |  | 1723.23 | 0.2012% | 1722.61 | 1722.92 | 1722.89 |  |  | 1722.8067 | 0.1766%
2023-03-14 | R15M10 | 10 | 9.6 | -2.27 | 23.3 | 980 | 1.0456 | -22.19 | -22.18 | -22.2 |  |  | -22.19 | -0.04507% | 1.9751 | -0.4986% | -1.442% | 2070.59 | 2072.02 | 2072 |  |  | 2071.5367 | 0.09019% | 2068.2 | 2069.67 | 2069.54 |  |  | 2069.1367 | -0.02577%
| R31M10 | 26 | 25.3 | -2.27 | 23.3 | 980 | 1.0456 | -22.42 | -22.42 | -22.43 |  |  | -22.4233 | -0.02575% | 1.9924 | -0.6778% | -1.658% | 1723.03 | 1722.65 | 1722.88 |  |  | 1722.8533 | 0.1793% | 1720.4 | 1720.01 | 1720.22 |  |  | 1720.21 | 0.02558%
2023-05-30 | R15M10 | 10 | 9.6 | -2.38 | 23.39 | 1018 | 1.0068 | -23.04 | -23.05 | -23.04 |  |  | -23.0433 | -0.02505% | 1.9821 | -0.1462% | -1.093% | 2072.22 | 2072.07 | 2072.13 |  |  | 2072.14 | 0.1193% | 2071.66 | 2071.48 | 2071.47 |  |  | 2071.5367 | 0.09019% | New calibration coefficient
| R31M10 | 26 | 25.3 | -2.38 | 23.39 | 1018 | 1.0068 | -23.36 | -23.37 | -23.36 |  |  | -23.3633 | -0.02471% | 2.0061 | 0.006584% | -0.9806% | 1722.65 | 1722.98 | 1722.66 |  |  | 1722.7633 | 0.1741% | 1721.32 | 1721.68 | 1721.38 |  |  | 1721.46 | 0.09827%
2023-06-13 | R15M10 | 10 | 9.6 | -2.22 | 23.54 | 1016 | 1.0093 | -23.07 | -23.08 | -23.05 |  |  | -23.0667 | -0.06622% | 1.989 | 0.2023% | -0.7477% | 2071.91 | 2072.01 | 2071.63 |  |  | 2071.85 | 0.1053% | 2071.32 | 2071.58 | 2071.1 |  |  | 2071.3333 | 0.08037%
| R31M10 | 26 | 25.3 | -2.19 | 23.54 | 1016 | 1.0093 | -23.39 | -23.39 | -23.38 |  |  | -23.3867 | -0.02469% | 2.0131 | 0.3543% | -0.6364% | 1722.79 | 1722.83 | 1722.78 |  |  | 1722.8 | 0.1762% | 1721.53 | 1721.52 | 1721.46 |  |  | 1721.5033 | 0.1008%
2023-06-19 | R15M10 | 10 | 9.6 | -2.33 | 24.1 | 1012 | 1.0152 | -23.07 | -23.07 | -23.06 |  |  | -23.0667 | -0.02503% | 2.0006 | 0.7883% | -0.1673% |  | 2073.12 | 2073.13 |  |  | 2073.125 | 0.1669% |  | 2072.01 | 2072.11 |  |  | 2072.06 | 0.1155%
| R31M10 | 26 | 25.3 | -2.33 | 24.1 | 1012 | 1.0152 | -23.31 | -23.32 | -23.32 |  |  | -23.3167 | -0.02476% | 2.0188 | 0.6389% | -0.3545% | 1723.81 | 1723.89 | 1723.93 |  |  | 1723.8767 | 0.2388% | 1722.25 | 1722.28 | 1722.34 |  |  | 1722.29 | 0.1465%
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!