## GTR1_SW_SOBP plans
ID.nr:
Titel: | [11 empty] | Dokumenttyp:
Protocol for measurements of the homogeneous cubic plan in SW with Roos IC | [11 empty] | Strålskyddsinstruktion
Godkänt av: | [10 empty] | Godkänt den:
Kategori: |  | Strålskydd
Skapat av: | [10 empty] | Skapat den:
Reviderat av: | [10 empty] | Reviderat den:
Room ID: | GTR 1
Dosimetry equipment: | Name and model:
phantom | Slab phantom
ionization chamber | PTW Roos TW34001 |  | SN:003389
electrometer | PTW UnidosWebline T10022 |  | SN:000456
cable |  |  | reference
Setup of measurement: | Mode |  | Charge&Current
| Units |  | Electrical
| Range |  | Medium/High
| Nominal voltage [V] |  | -200
calibration factor (proton PBS, Qo=150MeV): | ND,w,Qo [Gy/nC] 2021-06-08 |  | 0.08352
| ND,w,Qo [Gy/nC] 2023-03-27 |  | 0.0839 | 0.455%
Plan ID | R15M10 | R31M10
Prescribe dose TPS [Gy] | 2.018 | 2.019
Reference dose [Gy] | 1.991 | 1.983
Prescribe MU's | 2069.67 | 1719.77
zmeas [cm] | 10 | 26
Correction factors
electrometer - kelec | 1 | 1
ion recombination - ks | 1.0057 | 1.004
polarity - kpol | 1 | 1
beam quality - kQ,Qo | 1.0006 | 1.0005
iso at the surface of the IC
Date | Plan ID | zmeas | no.of SW plates | PPS Z-axis | TIC,air | pair | kT,p_air | Qmeas [nC] |  |  |  |  | Average Qmeas | STD | Average Dose | Deviation from reference | Deviation from TPS | Primary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's | Secondary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's
[DD/MM/YYYY] |  | [cm] | [cm] | [cm] | [oC] | [hPa] | [oC] | meas 1 | meas 2 | meas 3 | meas 4 | meas 5 | [nC] | [%] | [Gy] | [%] | [%] | MU_1 | MU_2 | MU_3 | MU_4 | MU_5 |  |  | MU_1 | MU_2 | MU_3 | MU_4 | MU_5
2021-07-27 | R15M10 | 10 | 9.6 |  | 22.87 | 1007 | 1.0161 | -23.26 | -23.27 | -23.28 |  |  | -23.27 | -0.04297% | 1.9871 | -0.197% | -1.532% | 2071.43 | 2071.37 | 2071.22 |  |  | 2071.34 | 0.08069% | 2071.6 | 2071.4 | 2071.27 |  |  | 2071.4233 | 0.08472% | REFERENCE MEASUREMENTS
| R31M10 | 26 | 25.3 |  | 22.87 | 1007 | 1.0161 | -23.22 | -23.24 | -23.23 |  |  | -23.23 | -0.04305% | 1.9802 | -0.1403% | -1.921% | 1722.51 | 1722.62 | 1722.4 |  |  | 1722.51 | 0.1593% | 1721.9 | 1721.96 | 1721.67 |  |  | 1721.8433 | 0.1206%
2021-07-28 | R15M10 | 10 | 9.6 |  | 23.1 | 1001 | 1.0229 | -23.17 | -23.16 | -23.16 |  |  | -23.1633 | -0.02493% | 1.9914 | 0.01868% | -1.32% | 2067.22 | 2067.82 | 2067.68 |  |  | 2067.5733 | -0.1013% | 2066.83 | 2067.33 | 2067.2 |  |  | 2067.12 | -0.1232%
| R31M10 | 26 | 25.3 |  | 23.1 | 1001 | 1.0229 | -23.09 | -23.08 | -23.08 |  |  | -23.0833 | -0.02501% | 1.981 | -0.09848% | -1.88% | 1719.42 | 1719.66 | 1719.66 |  |  | 1719.58 | -0.01105% | 1718.4 | 1718.47 | 1718.62 |  |  | 1718.4967 | -0.07404%
2021-07-29 | R15M10 | 10 | 9.6 |  | 23.6 | 994 | 1.0319 | -23.01 | -23.02 | -23 |  |  | -23.01 | -0.04346% | 1.9955 | 0.2252% | -1.116% | 2068.35 | 2068.46 | 2067.99 |  |  | 2068.2667 | -0.0678% | 2068.2 | 2068.2 | 2067.76 |  |  | 2068.0533 | -0.07811%
| R31M10 | 26 | 25.3 |  | 23.6 | 994 | 1.0319 | -22.96 | -22.97 | -22.98 |  |  | -22.97 | -0.04354% | 1.9886 | 0.2801% | -1.508% | 1720.68 | 1720.58 | 1720.71 |  |  | 1720.6567 | 0.05156% | 1719.79 | 1719.71 | 1719.81 |  |  | 1719.77 | 0%
2021-08-03 | R15M10 | 10 | 9.6 |  | 22.55 | 1006 | 1.016 | -23.31 | -23.33 | -23.33 |  |  | -23.3233 | -0.04951% | 1.9915 | 0.02297% | -1.315% | 2070.73 | 2072.09 | 2072.58 |  |  | 2071.8 | 0.1029% | 2070.95 | 2072.32 | 2072.68 |  |  | 2071.9833 | 0.1118%
| R31M10 | 26 | 25.3 |  | 22.55 | 1006 | 1.016 | -23.31 | -23.31 | -23.3 |  |  | -23.3067 | -0.02477% | 1.9866 | 0.1804% | -1.606% | 1723.6 | 1723.78 | 1722.59 |  |  | 1723.3233 | 0.2066% | 1723.01 | 1723.17 | 1721.98 |  |  | 1722.72 | 0.1715%
2021-08-04 | R15M10 | 10 | 9.6 |  | 22.21 | 1010 | 1.0108 | -23.39 | -23.43 | -23.4 |  |  | -23.4067 | -0.08893% | 1.9884 | -0.1322% | -1.468% | 2069.26 | 2069.17 | 2069.49 |  |  | 2069.3067 | -0.01756% | 2071.28 | 2071.16 | 2071.45 |  |  | 2071.2967 | 0.0786%
| R31M10 | 26 | 25.3 |  | 22.21 | 1010 | 1.0108 | -23.33 | -23.34 | -23.33 |  |  | -23.3333 | -0.02474% | 1.9787 | -0.217% | -1.996% | 1719.27 | 1719.46 | 1719.22 |  |  | 1719.3167 | -0.02636% | 1720.11 | 1720.35 | 1720.07 |  |  | 1720.1767 | 0.02365%
2021-09-21 | R15M10 | 10 | 9.6 | -2.11 | 22.3 | 1017.5 | 1.0036 | -23.54 | -23.55 | -23.56 |  |  | -23.55 | -0.04246% | 1.9864 | -0.2309% | -1.566% | 2068.86 | 2068.82 | 2068.66 |  |  | 2068.78 | -0.043% | 2070.77 | 2070.68 | 2070.66 |  |  | 2070.7033 | 0.04993%
| R31M10 | 26 | 25.3 | -2.04 | 22.3 | 1017.5 | 1.0036 | -23.53 | -23.56 | -23.52 |  |  | -23.5367 | -0.08844% | 1.9818 | -0.05897% | -1.841% | 1722.31 | 1722.39 | 1722.46 |  |  | 1722.3867 | 0.1522% | 1723.16 | 1723.18 | 1723.25 |  |  | 1723.1967 | 0.1993%
2021-10-25 | R15M10 | 10 | 9.6 | -2.06 | 22.41 | 1007 | 1.0145 | -23.53 | -23.53 | -23.52 |  |  | -23.5267 | -0.02454% | 2.0059 | 0.7471% | -0.6009% | 2070.54 | 2070.48 | 2070.67 |  |  | 2070.5633 | 0.04316% | 2072.93 | 2072.88 | 2073.02 |  |  | 2072.9433 | 0.1582%
| R31M10 | 26 | 25.3 | -2.06 | 22.41 | 1007 | 1.0145 | -23.46 | -23.46 | -23.45 |  |  | -23.4567 | -0.02461% | 1.9964 | 0.6773% | -1.118% | 1722.05 | 1722.15 | 1722.19 |  |  | 1722.13 | 0.1372% | 1723.31 | 1723.37 | 1723.42 |  |  | 1723.3667 | 0.2091%
2021-11-22 | R15M10 | 10 | 9.6 |  | 22.3 | 1006.5 | 1.0146 | -23.45 | -23.45 | -23.46 |  |  | -23.4533 | -0.02462% | 1.9999 | 0.4455% | -0.8984% | 2069.19 | 2069.39 | 2069.4 |  |  | 2069.3267 | -0.01659% | 2072.05 | 2072.14 | 2072.22 |  |  | 2072.1367 | 0.1192%
| R31M10 | 26 | 25.3 |  | 22.3 | 1006.5 | 1.0146 | -23.47 | -23.48 | -23.48 |  |  | -23.4767 | -0.02459% | 1.9984 | 0.7757% | -1.021% | 1720.61 | 1720.81 | 1720.77 |  |  | 1720.73 | 0.05582% | 1722.3 | 1722.42 | 1722.37 |  |  | 1722.3633 | 0.1508%
2021-12-10 | R15M10 | 10 | 9.6 | -2.17 | 22.48 | 1002 | 1.0198 | -23.34 | -23.34 | -23.33 |  |  | -23.3367 | -0.02474% | 2.0001 | 0.4559% | -0.8882% | 2068.8 | 2068.8 | 2069.01 |  |  | 2068.87 | -0.03865% | 2074.24 | 2074.27 | 2074.4 |  |  | 2074.3033 | 0.2239%
| R31M10 | 26 | 25.3 | -2.08 | 22.48 | 1002 | 1.0198 | -23.24 | -23.25 | -23.24 |  |  | -23.2433 | -0.02484% | 1.9886 | 0.2833% | -1.505% | 1719.95 | 1719.91 | 1719.87 |  |  | 1719.91 | 0.008141% | 1723.83 | 1723.68 | 1723.67 |  |  | 1723.7267 | 0.2301%
2022-01-24 | R15M10 | 10 | 9.6 | -2.01 | 22.32 | 1013 | 1.0082 | -23.59 | -23.58 | -23.59 |  |  | -23.5867 | -0.02448% | 1.9985 | 0.3752% | -0.9678% | 2069.13 | 2069.14 | 2069.19 |  |  | 2069.1533 | -0.02496% | 2077.7 | 2077.77 | 2077.82 |  |  | 2077.7633 | 0.391%
| R31M10 | 26 | 25.3 | -1.92 | 22.32 | 1013 | 1.0082 | -23.48 | -23.49 | -23.48 |  |  | -23.4833 | -0.02459% | 1.9863 | 0.1643% | -1.622% | 1720.46 | 1720.68 | 1720.57 |  |  | 1720.57 | 0.04652% | 1726.95 | 1727.16 | 1727.12 |  |  | 1727.0767 | 0.4249%
2022-02-21 | R15M10 | 10 | 9.6 | -1.93 | 22.4 | 972 | 1.051 | -22.6 | -22.61 |  |  |  | -22.605 | -0.03128% | 1.9966 | 0.2825% | -1.059% | 2072.56 | 2072.51 |  |  |  | 2072.535 | 0.1384% | 2079.04 | 2079.04 |  |  |  | 2079.04 | 0.4527%
| R31M10 | 26 | 25.3 | -1.93 | 22.4 | 972 | 1.051 | -22.53 | -22.51 |  |  |  | -22.52 | -0.0628% | 1.9857 | 0.1342% | -1.651% | 1722.93 | 1722.93 |  |  |  | 1722.93 | 0.1837% | 1727.66 | 1727.66 |  |  |  | 1727.66 | 0.4588%
2022-03-13 | R15M10 | 10 | 9.6 | -2 | 22.1 | 1024.3 | 0.9963 | -23.85 | -23.84 | -23.82 |  |  | -23.8367 | -0.06408% | 1.9959 | 0.2453% | -1.096% | 2075.7 | 2075.08 | 2075.11 |  |  | 2075.2967 | 0.2719% | 2082.52 | 2081.87 | 2081.88 |  |  | 2082.09 | 0.6001%
| R31M10 | 26 | 25.3 | -2 | 22.1 | 1024.3 | 0.9963 | -23.75 | -23.74 | -23.76 |  |  | -23.75 | -0.04211% | 1.9852 | 0.1096% | -1.675% | 1725.83 | 1725.83 | 1725.84 |  |  | 1725.8333 | 0.3526% | 1730.56 | 1730.59 | 1730.68 |  |  | 1730.61 | 0.6303%
2022-04-05 | R15M10 | 10 | 9.6 | -1.95 | 22.27 | 984 | 1.0377 | -22.92 | -22.93 | -22.91 |  |  | -22.92 | -0.04363% | 1.9989 | 0.3957% | -0.9475% | 2074.16 | 2074.14 | 2074.2 |  |  | 2074.1667 | 0.2173% | 2081.02 | 2080.89 | 2080.92 |  |  | 2080.9433 | 0.5447%
| R31M10 | 26 | 25.3 | -1.92 | 22.27 | 984 | 1.0377 | -22.85 | -22.85 | -22.85 |  |  | -22.85 | -1.904e-14% | 1.9893 | 0.3183% | -1.47% | 1724.96 | 1724.9 | 1725.02 |  |  | 1724.96 | 0.3018% | 1730.14 | 1729.97 | 1730 |  |  | 1730.0367 | 0.597%
2022-05-13 | R15M10 | 10 | 9.6 | -1.92 | 22.4 | 1000 | 1.0215 | -23.23 | -23.26 | -23.22 |  |  | -23.2367 | -0.08959% | 1.9949 | 0.1983% | -1.142% | 2072.15 | 2072.9 | 2072.19 |  |  | 2072.4133 | 0.1325% | 2080.6 | 2080.54 | 2080.51 |  |  | 2080.55 | 0.5257%
| R31M10 | 26 | 25.3 | -2.01 | 22.4 | 1000 | 1.0215 | -23.17 | -23.19 | -23.18 |  |  | -23.18 | -0.04314% | 1.9866 | 0.1829% | -1.603% | 1722.98 | 1723.21 | 1723.08 |  |  | 1723.09 | 0.193% | 1729.23 | 1729.45 | 1729.32 |  |  | 1729.3333 | 0.5561%
2022-05-13 | R15M10 | 10 | 9.6 | -1.94 | 22.4 | 999 | 1.0226 | -23.22 | -23.23 | -23.2 |  |  | -23.2167 | -0.06579% | 1.9952 | 0.2123% | -1.128% | 2072.12 | 2072.21 | 2072.16 |  |  | 2072.1633 | 0.1205% | 2080.66 | 2080.71 | 2080.6 |  |  | 2080.6567 | 0.5308% | new electrometer UnidosTango
| R31M10 | 26 | 25.3 | -1.94 | 22.4 | 999 | 1.0226 | -23.16 | -23.16 | -23.16 |  |  | -23.16 | 0% | 1.9869 | 0.1967% | -1.59% | 1722.88 | 1723.01 | 1723.18 |  |  | 1723.0233 | 0.1892% | 1729.6 | 1729.31 | 1729.45 |  |  | 1729.4533 | 0.5631%
2022-06-14 | R15M10 | 10 | 9.6 | -1.98 | 22.7 | 1008.4 | 1.0141 | -23.53 | -23.55 |  |  |  | -23.54 | -0.06008% | 2.0062 | 0.763% | -0.5852% | 2076.44 | 2076.26 |  |  |  | 2076.35 | 0.3228% | 2085.49 | 2085.36 |  |  |  | 2085.425 | 0.7612%
| R31M10 | 26 | 25.3 | -1.98 | 22.7 | 1008.4 | 1.0141 | -23.48 | -23.48 |  |  |  | -23.48 | 0% | 1.9976 | 0.7363% | -1.06% | 1725.78 | 1725.72 |  |  |  | 1725.75 | 0.3477% | 1732.72 | 1732.61 |  |  |  | 1732.665 | 0.7498%
2022-07-21 | R15M10 | 10 | 9.6 | -2.04 | 22.7 | 1008 | 1.0145 | -23.32 | -23.33 | -23.32 |  |  | -23.3233 | -0.02475% | 1.9885 | -0.1249% | -1.461% | 2069.33 | 2069.37 | 2069.36 |  |  | 2069.3533 | -0.0153% | 2076.34 | 2076.41 | 2076.42 |  |  | 2076.39 | 0.3247%
| R31M10 | 26 | 25.3 | -2.04 | 22.7 | 1008 | 1.0145 | -23.2 | -23.22 | -23.21 |  |  | -23.21 | -0.04308% | 1.9754 | -0.3825% | -2.159% | 1720.57 | 1720.73 | 1720.61 |  |  | 1720.6367 | 0.05039% | 1725.49 | 1725.69 | 1725.59 |  |  | 1725.59 | 0.3384%
2022-08-09 | R15M10 | 10 | 9.6 | -1.93 | 22.3 | 1022 | 0.99922 | -23.67 | -23.65 | -23.67 |  |  | -23.6633 | -0.0488% | 1.9872 | -0.1921% | -1.528% | 2068.46 | 2068.33 | 2068.29 |  |  | 2068.36 | -0.0633% | 2076.07 | 2076.01 | 2076.99 |  |  | 2076.3567 | 0.3231% | Ny config efter IC cyclo byte.
| R31M10 | 26 | 25.3 | -1.93 | 22.3 | 1022 | 0.99922 | -23.58 | -23.58 | -23.59 |  |  | -23.5833 | -0.02448% | 1.977 | -0.3017% | -2.079% | 1719.28 | 1719.85 | 1719.84 |  |  | 1719.6567 | -0.00659% | 1724.71 | 1725.26 | 1725.25 |  |  | 1725.0733 | 0.3084%
2022-09-16 | R15M10 | 10 | 9.6 | -1.95 | 22.1 | 986 | 1.035 | -22.84 | -22.84 | -22.82 |  |  | -22.8333 | -0.05057% | 1.9861 | -0.2442% | -1.579% | 2067.13 | 2067.02 | 2067.1 |  |  | 2067.0833 | -0.125% | 2076.28 | 2076.16 | 2076.18 |  |  | 2076.2067 | 0.3158% | new electrometer UnidosTango
| R31M10 | 26 | 25.3 | -1.95 | 22.1 | 986 | 1.035 | -22.74 | -22.73 | -22.75 |  |  | -22.74 | -0.04398% | 1.9746 | -0.4244% | -2.2% | 1719.02 | 1719.11 | 1719.24 |  |  | 1719.1233 | -0.0376% | 1725.83 | 1725.85 | 1725.98 |  |  | 1725.8867 | 0.3557%
2022-10-09 | R15M10 | 10 | 9.6 | -2.11 | 22.5 | 1013 | 1.0088 | -23.5 | -23.52 | -23.51 |  |  | -23.51 | -0.04254% | 1.9932 | 0.1099% | -1.23% | 2070.56 |  | 2070.7 |  |  | 2070.63 | 0.04638% | 2082.75 |  | 2082.88 |  |  | 2082.815 | 0.6351%
| R31M10 | 26 | 25.3 | -2 | 22.5 | 1013 | 1.0088 | -23.43 | -23.43 | -23.43 |  |  | -23.43 | -1.857e-14% | 1.983 | -0.002298% | -1.785% | 1721.04 | 1721.21 | 1721.38 |  |  | 1721.21 | 0.08373% | 1730.34 | 1730.48 | 1730.68 |  |  | 1730.5 | 0.6239%
2023-01-31 | R15M10 | 10 | 9.6 | 2.97 | 22.28 | 987 | 1.0346 | -22.92 | -22.93 | -22.93 |  |  | -22.9267 | -0.02518% | 1.9935 | 0.1231% | -1.217% | 2067.76 | 2068.26 | 2068.21 |  |  | 2068.0767 | -0.07698% | 2081.66 | 2081.54 | 2081.43 |  |  | 2081.5433 | 0.5737%
| R31M10 | 26 | 25.3 | 3 | 22.28 | 987 | 1.0346 | -22.84 | -22.85 | -22.83 |  |  | -22.84 | -0.04378% | 1.9825 | -0.02698% | -1.81% | 1719.63 | 1719.76 | 1719.4 |  |  | 1719.5967 | -0.01008% | 1730.06 | 1730.14 | 1729.75 |  |  | 1729.9833 | 0.5939%
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | New calibration coefficient
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!

## GTR2_SW_SOBP plans
ID.nr:
Titel: | [11 empty] | Dokumenttyp:
Protocol for measurements of the homogeneous cubic plan in SW with Roos IC | [11 empty] | Strålskyddsinstruktion
Godkänt av: | [10 empty] | Godkänt den:
Kategori: |  | Strålskydd
Skapat av: | [10 empty] | Skapat den:
Reviderat av: | [10 empty] | Reviderat den:
Room ID: | GTR 2
Dosimetry equipment: | Name and model:
phantom | Slab phantom
ionization chamber | PTW Roos TW34001 |  | SN:003389
electrometer | PTW UnidosWebline T10022 |  | SN:000456
cable |  |  | reference
Setup of measurement: | Mode |  | Charge&Current
| Units |  | Electrical
| Range |  | Medium/High
| Nominal voltage [V] |  | -200
calibration factor (proton PBS, Qo=150MeV): | ND,w,Qo [Gy/nC] 2021-06-08 |  | 0.08352
| ND,w,Qo [Gy/nC] 2023-03-27 |  | 0.0839 | 0.455%
Plan ID | R15M10 | R31M10
Prescribe dose TPS [Gy] | 2.004 | 2.026
Refernece dose [Gy] | 1.979 | 1.999
Prescribe MU's | 2069.67 | 1719.77
zmeas [cm] | 10 | 26
Correction factors
electrometer - kelec | 1 | 1
ion recombination - ks | 1.0057 | 1.004
polarity - kpol | 1 | 1
beam quality - kQ,Qo | 1.0006 | 1.0005
iso at the surface of the IC
Date | Plan ID | zmeas | no.of SW plates | PPS Z-axis | TIC,air | pair | kT,p_air | Qmeas [nC] |  |  |  |  | Average Qmeas | STD | Average Dose | Deviation from reference | Deviation from TPS | Primary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's | Secondary Chamber |  |  |  |  | Average MU's | Deviation from Prescribe MU's
[DD/MM/YYYY] |  | [cm] | [cm] | [cm] | [oC] | [hPa] | [oC] | meas 1 | meas 2 | meas 3 | meas 4 | meas 5 | [nC] | [%] | [Gy] | [%] | [%] | MU_1 | MU_2 | MU_3 | MU_4 | MU_5 |  |  | MU_1 | MU_2 | MU_3 | MU_4 | MU_5
2021-07-27 | R15M10 | 10 | 9.6 |  | 22.63 | 1007 | 1.0152 | -23.18 | -23.2 | -23.18 |  |  | -23.1867 | -0.0498% | 1.9784 | -0.03249% | -1.28% | 2074.56 | 2074.61 | 2074.75 |  |  | 2074.64 | 0.2401% | 2076.26 | 2076.26 | 2076.42 |  |  | 2076.3133 | 0.321% | REFERENCE MEASUREMENTS
| R31M10 | 26 | 25.3 |  | 22.63 | 1007 | 1.0152 | -23.49 | -23.46 | -23.46 |  |  | -23.47 | -0.0738% | 1.9991 | 0.002679% | -1.33% | 1724.77 | 1724.56 | 1724.69 |  |  | 1724.6733 | 0.2851% | 1725.27 | 1724.97 | 1725.11 |  |  | 1725.1167 | 0.3109%
2021-07-29 | R15M10 | 10 | 9.6 |  | 22.93 | 994 | 1.0296 | -22.78 | -22.88 | -22.81 |  |  | -22.8233 | -0.2248% | 1.9748 | -0.2109% | -1.456% | 2071.77 | 2071.99 | 2071.9 |  |  | 2071.8867 | 0.1071% | 2073.51 | 2073.69 | 2073.63 |  |  | 2073.61 | 0.1904%
| R31M10 | 26 | 25.3 |  | 22.93 | 994 | 1.0296 | -23.13 | -23.12 | -23.14 |  |  | -23.13 | -0.04323% | 1.9979 | -0.05581% | -1.388% | 1722.66 | 1722.67 | 1722.68 |  |  | 1722.67 | 0.1686% | 1723.21 | 1723.21 | 1723.26 |  |  | 1723.2267 | 0.201%
2021-08-02 | R15M10 | 10 | 9.6 |  | 22.62 | 1000 | 1.0223 | -23.07 | -23.03 | -23.09 |  |  | -23.0633 | -0.1325% | 1.9815 | 0.1284% | -1.121% | 2075.9 | 2075.43 | 2075.18 |  |  | 2075.5033 | 0.2818% | 2077.35 | 2076.94 | 2076.73 |  |  | 2077.0067 | 0.3545%
| R31M10 | 26 | 25.3 |  | 22.62 | 1000 | 1.0223 | -23.38 | -23.37 | -23.38 |  |  | -23.3767 | -0.0247% | 2.005 | 0.2988% | -1.038% | 1727.32 | 1727.24 | 1728.89 |  |  | 1727.8167 | 0.4679% | 1727.68 | 1727.58 | 1729.25 |  |  | 1728.17 | 0.4884%
2021-08-03 | R15M10 | 10 | 9.6 |  | 22.26 | 1006 | 1.015 | -23.19 | -23.27 | -23.25 |  |  | -23.2367 | -0.1792% | 1.9821 | 0.1572% | -1.092% | 2075.75 | 2076.02 | 2075.15 |  |  | 2075.64 | 0.2885% | 2077.05 | 2077.12 | 2076.59 |  |  | 2076.92 | 0.3503%
| R31M10 | 26 | 25.3 |  | 22.26 | 1006 | 1.015 | -23.44 | -23.46 | -23.44 |  |  | -23.4467 | -0.04925% | 1.9966 | -0.1225% | -1.454% | 1722.99 | 1722.9 | 1722.78 |  |  | 1722.89 | 0.1814% | 1723.02 | 1722.89 | 1722.71 |  |  | 1722.8733 | 0.1805%
2021-08-05 | R15M10 | 10 | 9.6 |  | 22.316 | 1011 | 1.0101 | -23.3 | -23.31 | -23.34 |  |  | -23.3167 | -0.08928% | 1.9795 | 0.02396% | -1.224% | 2077.74 | 2077.77 | 2078.64 |  |  | 2078.05 | 0.4049% | 2078.85 | 2078.92 | 2079.73 |  |  | 2079.1667 | 0.4588%
| R31M10 | 26 | 25.3 |  | 22.316 | 1011 | 1.0101 | -23.55 | -23.55 | -23.62 |  |  | -23.5733 | -0.1714% | 1.9978 | -0.06064% | -1.393% | 1723.95 | 1723.77 | 1727.12 |  |  | 1724.9467 | 0.301% | 1724.06 | 1723.77 | 1727.15 |  |  | 1724.9933 | 0.3037%
2021-09-02 | R15M10 | 10 | 9.6 | -2.36 | 22.93 | 1012 | 1.0112 | -23.24 | -23.24 | -23.27 |  |  | -23.25 | -0.0745% | 1.976 | -0.1535% | -1.399% | 2071.57 | 2070.23 | 2069.96 |  |  | 2070.5867 | 0.04429% | 2072.65 | 2071.28 | 2070.98 |  |  | 2071.6367 | 0.09502%
| R31M10 | 26 | 25.3 | -2.28 | 22.93 | 1012 | 1.0112 | -23.59 | -23.61 | -23.59 |  |  | -23.5967 | -0.04893% | 2.0019 | 0.1471% | -1.188% | 1722.55 | 1722.66 | 1722.78 |  |  | 1722.6633 | 0.1682% | 1722.72 | 1722.77 | 1722.9 |  |  | 1722.7967 | 0.176%
2021-10-17 | R15M10 | 10 | 9.6 | -2.6 | 22.2 | 1006 | 1.0148 | -23.35 | -23.4 | -23.35 |  |  | -23.3667 | -0.1235% | 1.9928 | 0.6971% | -0.5591% | 2072.25 | 2072.17 | 2072.25 |  |  | 2072.2233 | 0.1234% | 2073.63 | 2073.62 | 2073.61 |  |  | 2073.62 | 0.1909%
| R31M10 | 26 | 25.3 | -2.41 | 22.2 | 1006 | 1.0148 | -23.7 | -23.65 | -23.65 |  |  | -23.6667 | -0.122% | 2.0149 | 0.7941% | -0.5491% | 1722.75 | 1723.08 | 1723.38 |  |  | 1723.07 | 0.1919% | 1723.11 | 1723.35 | 1723.66 |  |  | 1723.3733 | 0.2095%
2021-11-01 | R15M10 | 10 | 9.6 | -2.23 | 22.618 | 996 | 1.0264 | -23.06 | -23.08 | -23.07 |  |  | -23.07 | -0.04335% | 1.9901 | 0.5589% | -0.6955% | 2076.15 | 2076.21 | 2076.28 |  |  | 2076.2133 | 0.3162% | 2077.65 | 2077.67 | 2077.75 |  |  | 2077.69 | 0.3875%
| R31M10 | 26 | 25.3 | -2.23 | 22.618 | 996 | 1.0264 | -23.37 | -23.36 | -23.37 |  |  | -23.3667 | -0.02471% | 2.0122 | 0.6579% | -0.6836% | 1725.57 | 1725.63 | 1725.68 |  |  | 1725.6267 | 0.3405% | 1725.93 | 1726.01 | 1726.02 |  |  | 1725.9867 | 0.3615%
2021-12-10 | R15M10 | 10 | 9.6 | -2.19 | 22.8 | 1004 | 1.0189 | -23.26 | -23.19 | -23.17 |  |  | -23.2067 | -0.2036% | 1.9871 | 0.4104% | -0.8422% | 2072.56 | 2072.63 | 2072.55 |  |  | 2072.58 | 0.1406% | 2073.54 | 2073.48 | 2073.45 |  |  | 2073.49 | 0.1846%
| R31M10 | 26 | 25.3 | -2.19 | 22.8 | 1004 | 1.0189 | -23.44 | -23.45 | -23.45 |  |  | -23.4467 | -0.02462% | 2.0042 | 0.2594% | -1.077% | 1722.46 | 1722.45 | 1722.61 |  |  | 1722.5067 | 0.1591% | 1722.5 | 1722.45 | 1722.52 |  |  | 1722.49 | 0.1582%
2022-02-15 | R15M10 | 10 | 9.6 | -2.17 | 22.23 | 991.3 | 1.0299 | -22.8 | -22.77 | -22.78 |  |  | -22.7833 | -0.06705% | 1.9721 | -0.3506% | -1.594% | 2068.8 | 2068.83 | 2068.79 |  |  | 2068.8067 | -0.04171% | 2069.55 | 2069.54 | 2069.48 |  |  | 2069.5233 | -0.007086%
| R31M10 | 26 | 25.3 | -2.17 | 22.23 | 991.3 | 1.0299 | -23.03 | -23.02 | -23.03 |  |  | -23.0267 | -0.02507% | 1.9897 | -0.4672% | -1.794% | 1719.76 | 1719.79 | 1719.89 |  |  | 1719.8133 | 0.00252% | 1719.6 | 1719.61 | 1719.68 |  |  | 1719.63 | -0.008141%
2022-04-05 | R15M10 | 10 | 9.6 | -2.18 | 22.26 | 984 | 1.0377 | -22.68 | -22.69 | -22.67 |  |  | -22.68 | -0.04409% | 1.9779 | -0.05654% | -1.303% | 2076.02 | 2076.14 | 2076.31 |  |  | 2076.1567 | 0.3134% | 2077.2 | 2077.24 | 2077.49 |  |  | 2077.31 | 0.3691%
| R31M10 | 26 | 25.3 | -2.18 | 22.26 | 984 | 1.0377 | -22.98 | -22.98 | -22.98 |  |  | -22.98 | 0% | 2.0006 | 0.07816% | -1.256% | 1726.45 | 1726.28 | 1726.61 |  |  | 1726.4467 | 0.3882% | 1726.74 | 1726.5 | 1726.8 |  |  | 1726.68 | 0.4018%
2022-05-17 | R15M10 | 10 | 9.6 | -2.16 | 23.9 | 1022 | 1.0046 | -23.56 | -23.56 | -23.56 |  |  | -23.56 | 0% | 1.9892 | 0.516% | -0.7379% | 2076.98 | 2077.06 | 2076.93 |  |  | 2076.99 | 0.3537% | 2080.08 | 2080.16 | 2079.98 |  |  | 2080.0733 | 0.5027%
| R31M10 | 26 | 25.3 | -2.16 | 23.9 | 1022 | 1.0046 | -23.89 | -23.89 | -23.91 |  |  | -23.8967 | -0.04832% | 2.0141 | 0.757% | -0.5858% | 1727.04 | 1727.11 | 1727.27 |  |  | 1727.14 | 0.4285% | 1728.89 | 1728.9 | 1729.03 |  |  | 1728.94 | 0.5332%
2022-07-01 | R15M10 | 10 | 9.6 | -2.23 | 22.53 | 1010 | 1.0119 | -23.41 | -23.39 | -23.37 | -23.38 |  | -23.3875 | -0.07302% | 1.9889 | 0.4999% | -0.7538% | 2075.8 | 2075.91 | 2076.2 | 2075.99 |  | 2075.975 | 0.3046% | 2079.01 | 20,79,14 | 2079.54 | 2079.23 |  | 2079.12 | 0.4566% | 17:00 start measuring // YM AC
| R31M10 | 26 | 25.3 |  | 22.53 | 1010 | 1.0119 | -23.69 | -23.69 | -23.7 |  |  | -23.6933 | -0.02437% | 2.0114 | 0.6204% | -0.7206% | 1725.56 | 1725.61 | 1725.71 |  |  | 1725.6267 | 0.3405% | 1727.52 | 1727.52 | 1727.6 |  |  | 1727.5467 | 0.4522%
2022-07-15 | R15M10 | 10 | 9.6 | -2.21 | 22.27 | 1002 | 1.0191 | -22.98 | -23 | -22.94 |  |  | -22.9733 | -0.133% | 1.9675 | -0.5792% | -1.819% | 2071.26 | 2071.48 | 2071.02 |  |  | 2071.2533 | 0.0765% | 2072.6 | 2072.35 | 2071.91 |  |  | 2072.2867 | 0.1264%
| R31M10 | 26 | 25.3 | -2.21 | 22.27 | 1002 | 1.0191 | -23.11 | -23.12 | -23.18 |  |  | -23.1367 | -0.1636% | 1.9781 | -1.046% | -2.365% | 1722.12 | 1722.39 | 1722.23 |  |  | 1722.2467 | 0.144% | 1722.04 | 1722.28 | 1722.15 |  |  | 1722.1567 | 0.1388%
2022-08-10 | R15M10 | 10 | 9.6 | -2.22 | 22.2 | 1017 | 1.0038 | -23.38 | -23.24 | -23.37 | -23.38 |  | -23.3425 | -0.2934% | 1.9692 | -0.4951% | -1.736% | 2070.66 | 2070.63 | 2070.69 | 2070.64 |  | 2070.655 | 0.04759% | 2071.57 | 2071.5 | 2071.48 | 2071.4 |  | 2071.4875 | 0.08782% | Ny config efter IC cyclo byte.
| R31M10 | 26 | 25.3 | -2.22 | 22.2 | 1017 | 1.0038 | -23.55 | -23.54 | -23.54 |  |  | -23.5433 | -0.02452% | 1.9827 | -0.8156% | -2.137% | 1721.15 | 1721.28 | 1721.32 |  |  | 1721.25 | 0.08606% | 1720.91 | 1721.01 | 1721.01 |  |  | 1720.9767 | 0.07016%
2022-09-10 | R15M10 | 10 | 9.6 | -2.2 | 22.4 | 1014 | 1.0074 | -23.21 | -23.24 | -23.3 | -23.27 |  | -23.255 | -0.1665% | 1.969 | -0.5074% | -1.749% | 2070.49 | 2070.5 | 2070.57 | 2070.63 |  | 2070.5475 | 0.0424% | 2070.6 | 2070.6 | 2070.69 | 2070.83 |  | 2070.68 | 0.0488%
| R31M10 | 26 | 25.3 | -2.2 | 22.4 | 1014 | 1.0074 | -23.45 | -23.43 | -23.43 |  |  | -23.4367 | -0.04927% | 1.9809 | -0.9058% | -2.226% | 1721.19 | 1721.04 | 1721.21 |  |  | 1721.1467 | 0.08005% | 1720.46 | 1720.31 | 1720.41 |  |  | 1720.3933 | 0.03625%
2023-02-04 | R15M10 | 10 | 9.6 | -2.29 | 22.097 | 1034 | 0.98694 | -23.69 | -23.67 | -23.65 |  |  | -23.67 | -0.0845% | 1.9633 | -0.7925% | -2.03% | 2072.43 | 2072.7 | 2073.06 |  |  | 2072.73 | 0.1478% | 2069.58 | 2069.78 | 2070.12 |  |  | 2069.8267 | 0.00757%
| R31M10 | 26 | 25.3 | -2.29 | 22.097 | 1034 | 0.98694 | -23.99 | -24 | -24 |  |  | -23.9967 | -0.02406% | 1.987 | -0.6026% | -1.927% | 1723.75 | 1723.75 | 1723.17 |  |  | 1723.5567 | 0.2202% | 1720.59 | 1720.6 | 1720.5 |  |  | 1720.5633 | 0.04613%
2023-04-22 | R15M10 | 10 | 9.7 (Blue SW) | -2.23 | 22.9 | 1007.7 | 1.0155 | -23 | -22.97 | -23.02 |  |  | -22.9967 | -0.1094% | 1.9715 | -0.3789% | -1.622% | 2070.29 | 2070.36 | 2070.46 |  |  | 2070.37 | 0.03382% | 2070.4 | 2070.38 | 2070.43 |  |  | 2070.4033 | 0.03543% | Blue SW phantom + new calibration coefficient
| R31M10 | 26 | 25.5 (Blue SW) | -2.18 | 22.9 | 1007.7 | 1.0155 | -23.29 | -23.28 | -23.29 |  |  | -23.2867 | -0.02479% | 1.9929 | -0.3054% | -1.634% | 1720.95 | 1721.26 | 1721.6 |  |  | 1721.27 | 0.08722% | 1720.25 | 1720.56 | 1720.85 |  |  | 1720.5533 | 0.04555%
2023-07-17 | R15M10 | 10 | 9.7 (Blue SW) | -4.15 | 22.41 | 1001 | 1.0206 | -22.79 | -22.67 | -22.69 |  |  | -22.7167 | -0.283% | 1.9573 | -1.097% | -2.331% | 2079.2 | 2076.06 | 2076.25 |  |  | 2077.17 | 0.3624% | 2077.54 | 2074.41 | 2074.5 |  |  | 2075.4833 | 0.2809% | Continued with the Blue SW
| R31M10 | 26 | 25.5 (Blue SW) | -4.15 | 22.41 | 1001 | 1.0206 | -22.99 | -22.99 | -23 |  |  | -22.9933 | -0.02511% | 1.9777 | -1.066% | -2.385% | 1727.51 | 1727.17 | 1727.54 |  |  | 1727.4067 | 0.4441% | 1725.45 | 1725.06 | 1725.43 |  |  | 1725.3133 | 0.3223% | AJ, JR
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R15M10 | 10 | 9.6 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!
| R31M10 | 26 | 25.3 |  |  |  | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0!