## GTR1
Gantry: | GTR1
Inställningar: | Fältstorlek 10x10cm2, Spotstorlek 3mm, Spotavstånd 2.5mm, 1681 spots, 0.4 MU/spot, Vattenytan i isocentrum
Mätdjup: | 2 | cm | SW phantom
Jonkammare: | PTW Roos TW34001 #002519
N_D,w,Q0 = | 0.08465 | Gy/nC (Q0 = 150 MeV p+, kalibrerad på Skandionkliniken 2015-03-18 och 2015-04-10)
Elektrometer: | PTW UNIDOSE T10022 SN00456
k_elec = | 1 | [53 empty] | IC Roos + electrometer from GTR2
|  |  | REFERENCE | [16 empty] | new optical solution with new iso | [7 empty] | new iso position |  | deflector change | site config update | [9 empty] | cyclo opening |  | Acute deflector change |  |  | Before cyclo PM | After cyclo PM |  | fast check | fast check |  | After IC cyclo replacement |  |  | Before IC1 replacement | After IC1 replacement |  |  |  | Data for Roos IC #3389 | [8 empty] | ? |  | Quick test for PROBREAST project |  |  | New calibration coefficient (0.35% higher than the previous one), blue slab phantom | New calibration coefficient | Intervention to LLRF | Validation after LLRF calibration
|  |  | 2018-10-16 | 2018-10-18 | 2018-10-19 | 2018-10-22 | 2018-10-26 | 2018-11-05 | 2018-11-15 | 2018-11-27 | 2018-12-12 | 2019-01-08 | 2019-01-24 | 2019-03-04 | 2019-03-17 | 2019-04-11 | 2019-04-22 | 2019-05-07 | 2019-05-13 | 2019-05-13 | 2019-07-01 | 2019-07-17 | 2019-08-14 | 2019-10-08 | 2019-10-14 | 2019-11-07 | 2020-01-09 | 2020-01-18 | 2020-01-27 | 2020-02-03 | 2020-03-06 | 2020-04-21 | 2020-05-22 | 2020-06-15 | 2020-06-30 | 2020-07-06 | 2020-07-21 | 2020-07-27 | 2020-08-18 | 2020-09-08 | 2020-10-12 | 2020-11-03 | 2020-11-18 | 2021-01-12 | 2021-01-26 | 2021-02-11 | 2021-02-14 | 2021-02-15 | 2021-03-23 | 2021-04-08 | 2021-05-05 | 2021-05-09 | 2021-06-17 | 2021-07-13 | 2021-07-14 | 2021-07-19 | 2021-09-14 | 2021-10-25 | 2022-04-25 | 2022-05-13 | 2022-05-13 | 2022-06-14 | 2022-07-21 | 2022-08-09 | 2022-08-09 | 2022-09-16 | 2022-10-09 | 2022-10-14 | 2022-10-18 | 2022-11-04 | 2022-12-03 | 2023-01-31 (Roos IC #3389) | 2023-03-14 | 2023-04-21 | 2023-05-30 | 2023-06-13 | 2023-06-19 | 2023-07-18 |  | Mean | SD | SD%
|  | Energi | Dos | Dos | Dos | Dos | Dos | Dos | Dos | Dos | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | [8 empty] | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Energi | Dose | Dose | %
|  | (MeV) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | [8 empty] | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (MeV) | (Gy) | (Gy)
|  | 60 | 0.95015 | 0.95145 | 0.94958 | 0.94481 | 0.95454 | 0.95125 | 0.95002 | 0.94824 | 0.94855 | 0.94398 | 0.94256 | 0.9408 | 0.95033 | 0.94519 | 0.95444 | 0.96982 | 0.94181 | 0.94546 | 0.9457 | 0.94537 | 0.94759 | 0.95025 | 0.95258 | 0.94764 | 0.94972 | 0.94991 | 0.94361 | 0.94861 | 0.94515 | 0.9476 | 0.94456 | 0.9481 | 0.95011 | 0.94077 | 0.94905 | 0.94737 | 0.9499 | 0.94772 | 0.94642 | 0.95122 | 0.94816 | 0.94984 | 0.9501 | 0.94566 | 0.95017 | 0.94532 | 0.94509 | 0.95559 | 0.9432 | 0.94435 | 0.94491 | 0.94181 | 0.94177 | 0.94156 | 0.94406 | 0.94608 | 0.9482 | 0.93864 | 0.93849 | 0.93897 | 0.93666 | 0.94152 | 0.93888 | 0.93736 | 0.93956 | 0.94247 | 0.93904 | 0.94706 | 0.94872 | 0.93951 | 0.94274 | 0.94006 | 0.94638 | 0.94545 | 0.94593 | 0.94511 | 60 | 0.95011 | 0.0035301 | 0.3716%
|  | 70 | 0.81474 | 0.81307 | 0.81094 | 0.81184 | 0.81663 | 0.81528 | 0.81271 | 0.8104 | 0.81218 | 0.8099 | 0.81162 | 0.80904 | 0.81198 | 0.81168 | 0.81453 | 0.82848 | 0.81112 | 0.8103 | 0.8095 | 0.80988 | 0.81103 | 0.81534 | 0.81468 | 0.81389 | 0.81442 | 0.81328 |  | 0.81352 | 0.81282 | 0.8125 | 0.81085 | 0.8126 | 0.8121 | 0.80899 | 0.81049 | 0.81029 | 0.8132 | 0.81061 |  | 0.81207 | 0.81184 | 0.81269 | 0.81344 | 0.81145 | 0.81326 | 0.81285 | 0.814 | 0.81595 | 0.80854 | 0.80915 | 0.81071 | 0.80736 | 0.80781 | 0.80992 | 0.80999 | 0.8123 | 0.81362 | 0.80632 | 0.80624 | 0.80857 | 0.80606 | 0.80685 | 0.80704 | 0.80635 | 0.80722 | 0.80838 | 0.80474 | 0.81177 | 0.81296 | 0.80715 | 0.8083 | 0.80761 | 0.81104 | 0.81165 | 0.81054 | 0.81118 | 70 | 0.81344 | 0.0022829 | 0.2806%
|  | 100 | 0.71295 | 0.71074 | 0.70882 | 0.71092 | 0.71359 | 0.71067 | 0.7101 | 0.70777 | 0.71069 | 0.70739 | 0.71113 | 0.70663 | 0.7096 | 0.70847 | 0.71263 | 0.72357 | 0.70838 | 0.70827 | 0.7095 | 0.70845 | 0.70858 | 0.71174 | 0.712 | 0.71031 | 0.7095 | 0.7107 | 0.7088 | 0.71206 | 0.71013 | 0.7108 | 0.70953 | 0.712 | 0.71113 | 0.70968 | 0.70842 | 0.70719 | 0.7121 | 0.70845 | 0.70967 | 0.71174 | 0.71084 | 0.70774 | 0.70968 | 0.70791 | 0.71093 | 0.70993 | 0.71052 | 0.70947 | 0.70688 | 0.70461 | 0.7061 | 0.70331 | 0.70408 | 0.70455 | 0.70556 | 0.70835 | 0.7098 | 0.70294 | 0.70308 | 0.70609 | 0.70209 |  | 0.70284 | 0.7017 | 0.70351 | 0.70376 | 0.70122 | 0.70796 | 0.70907 | 0.70243 | 0.70394 | 0.70368 | 0.70567 | 0.70423 | 0.70514 | 0.70642 | 100 | 0.71141 | 0.0019065 | 0.268%
|  | 120 | 0.68457 | 0.68213 | 0.6802 | 0.68167 | 0.68413 | 0.68167 | 0.6799 | 0.67819 | 0.68135 | 0.67723 | 0.68071 | 0.67721 | 0.67933 | 0.67781 | 0.68515 | 0.69384 | 0.67735 | 0.67876 | 0.6812 | 0.67942 | 0.67976 | 0.68363 | 0.68435 | 0.68182 | 0.68135 | 0.68305 |  | 0.68615 | 0.68265 | 0.6836 | 0.68106 | 0.6836 | 0.68404 | 0.68379 | 0.6815 |  | 0.6852 | 0.68136 |  | 0.68279 | 0.68323 | 0.67907 | 0.68243 | 0.68024 | 0.6842 | 0.68155 | 0.68363 | 0.68003 | 0.67789 | 0.676 | 0.67946 | 0.67662 | 0.67605 | 0.67772 | 0.67924 | 0.68174 | 0.68258 | 0.67534 | 0.67485 | 0.6795 | 0.67291 |  | 0.67315 | 0.67283 | 0.67639 | 0.67756 | 0.6744 | 0.68131 | 0.6829 | 0.67484 | 0.67621 | 0.6764 | 0.67894 | 0.67801 | 0.67847 | 0.68026 | 120 | 0.68254 | 0.0018058 | 0.2646%
|  | 150 | 0.67736 | 0.67293 | 0.67178 | 0.67393 | 0.67654 | 0.67242 | 0.6726 | 0.66838 | 0.67281 | 0.66769 | 0.67407 | 0.66901 | 0.66986 | 0.67088 | 0.67572 | 0.68501 | 0.67026 | 0.66928 | 0.6709 | 0.67094 | 0.67124 | 0.67552 | 0.67548 | 0.6729 | 0.67371 | 0.67376 | 0.67289 |  | 0.67446 | 0.6753 | 0.67343 | 0.6752 | 0.67592 | 0.67464 | 0.67209 | 0.67306 | 0.6771 | 0.67256 | 0.67303 | 0.67471 | 0.67497 | 0.67071 | 0.67312 | 0.67135 | 0.67506 | 0.67285 | 0.67519 | 0.67117 | 0.67025 | 0.66759 | 0.67179 | 0.66814 | 0.66757 | 0.66903 | 0.66927 | 0.67182 | 0.67334 | 0.66704 | 0.66624 | 0.67254 | 0.66416 | 0.66483 | 0.66422 | 0.66359 | 0.66853 | 0.66879 | 0.66546 | 0.67271 | 0.67405 | 0.66596 | 0.66735 | 0.66752 | 0.66985 | 0.66843 | 0.67007 | 0.6711 | 150 | 0.67451 | 0.0023748 | 0.3521%
|  | 170 | 0.6774 | 0.67392 | 0.67193 | 0.67444 | 0.67632 | 0.67358 | 0.67274 | 0.67018 | 0.67433 | 0.66908 | 0.6743 | 0.6699 | 0.67195 | 0.67211 | 0.67565 | 0.68532 | 0.67064 |  | 0.6706 | 0.67042 | 0.67216 | 0.67598 | 0.67513 | 0.67289 | 0.67349 | 0.67402 |  | 0.67739 | 0.67449 | 0.6753 | 0.67317 | 0.6752 | 0.67569 | 0.67539 | 0.67186 |  | 0.6767 | 0.67247 | 0.67317 | 0.6745 | 0.67584 | 0.67027 | 0.67384 | 0.6712 | 0.67553 | 0.67267 | 0.67553 | 0.67212 | 0.66986 | 0.66778 | 0.67201 | 0.66833 | 0.66793 | 0.66888 | 0.66852 | 0.67253 | 0.67308 | 0.66729 | 0.66755 | 0.6722 | 0.66474 |  | 0.66478 | 0.66331 | 0.66848 | 0.66798 | 0.66553 | 0.67206 | 0.67342 | 0.66664 | 0.66713 | 0.66737 | 0.66969 | 0.66865 | 0.66932 | 0.67086 | 170 | 0.6748 | 0.0021325 | 0.316%
|  | 200 | 0.68194 | 0.67745 | 0.67496 | 0.67789 | 0.67997 | 0.67753 | 0.67561 | 0.67307 | 0.67719 | 0.67322 | 0.67684 | 0.67199 | 0.6761 | 0.67526 | 0.67959 | 0.6868 | 0.67366 | 0.67385 | 0.6738 | 0.67429 | 0.67452 | 0.67858 | 0.67829 | 0.67652 | 0.67646 | 0.67649 | 0.67578 | 0.68102 | 0.67531 | 0.6778 | 0.67573 | 0.6774 | 0.67833 | 0.679 | 0.67493 | 0.67635 | 0.6799 | 0.67546 |  | 0.67688 | 0.67749 | 0.67252 | 0.67575 | 0.67376 | 0.67789 | 0.6756 | 0.67724 | 0.67381 | 0.67196 | 0.66963 | 0.67354 | 0.67006 | 0.66963 | 0.67084 | 0.67038 | 0.67426 | 0.67519 | 0.67004 | 0.66898 | 0.6734 | 0.66589 |  | 0.66528 | 0.665 | 0.66946 | 0.6702 | 0.66763 | 0.67423 | 0.67565 | 0.66798 | 0.6698 | 0.66873 | 0.67102 | 0.67061 | 0.67116 | 0.67261 | 200 | 0.67844 | 0.0026465 | 0.3901%
|  | 226 | 0.68151 | 0.67811 | 0.67587 | 0.67918 | 0.68047 | 0.67727 | 0.67642 | 0.67417 | 0.67794 | 0.67399 | 0.67755 | 0.67303 | 0.67687 | 0.67687 | 0.68164 | 0.6869 | 0.67593 | 0.6743 | 0.6765 | 0.67339 | 0.67535 | 0.68085 | 0.67833 | 0.6774 | 0.67777 | 0.67841 |  | 0.68137 | 0.67677 | 0.679 | 0.67702 | 0.6782 | 0.67979 | 0.67963 | 0.67645 | 0.67543 | 0.6817 | 0.67732 | 0.67669 | 0.67771 | 0.67811 | 0.67373 | 0.67677 | 0.67477 | 0.67783 | 0.67542 | 0.67711 | 0.67438 | 0.67226 | 0.66932 | 0.67351 | 0.67049 | 0.6705 | 0.67141 | 0.67164 | 0.67515 | 0.67571 | 0.6705 | 0.66995 | 0.67436 | 0.66505 | 0.66543 | 0.66547 | 0.66549 | 0.67054 | 0.67124 | 0.66786 | 0.6757 | 0.67708 | 0.66865 | 0.67153 | 0.66995 | 0.67174 | 0.67369 | 0.67247 | 0.67358 | 226 | 0.67903 | 0.0021843 | 0.3217%
[36 empty] | 0.94489 |  |  | 0.9512 | 0.94772
% Diff from mean
|  |  | 2018-10-16 | 2018-10-18 | 2018-10-19 | 2018-10-22 | 2018-10-26 | 2018-11-05 | 2018-11-15 | 2018-11-27 | 2018-12-12 | 2019-01-08 | 2019-01-24 | 2019-03-04 | 2019-03-17 | 2019-04-11 | 2019-04-22 | 2019-05-07 | 2019-05-13 | 2019-05-13 | 2019-07-01 | 2019-07-17 | 2019-08-14 | 2019-10-08 | 2019-10-14 | 2019-11-07 | 2020-01-09 | 2020-01-18 | 2020-01-27 | 2020-02-03 | 2020-03-06 | 2020-04-21 | 2020-05-22 | 2020-06-15 | 2020-06-30 | 2020-07-06 | 2020-07-21 | 2020-07-27 | 2020-08-18 | 2020-09-08 | 2020-10-12 | 2020-11-03 | 2020-11-18 | 2021-01-12 | 2021-01-26 | 2021-02-11 | 2021-02-14 | 2021-01-26 | 2021-03-23 | 2021-04-08 | 2021-05-05 | 2021-05-09 | 2021-06-17 | 2021-07-13 | 2021-07-14 | 2021-07-19 | 2021-09-14 | 2021-10-25 | 2021-10-25 | [19 empty] | Energi
|  | Energi | [5 empty] | afer cyclo PM | Dos | after ICcyclo calib change | ICcyclo repairment | [67 empty] | (MeV)
|  | (MeV) | % | % | % | % | % | % | %
|  | 60 | 0.004674% | 0.1416% | -0.05543% | -0.5577% | 0.4669% | 0.1208% | -0.009585% | -0.196% | -0.164% | -0.6452% | -0.7941% | -0.98% | 0.02298% | -0.5181% | 0.4565% | 2.075% | -0.8737% | -0.4892% | -0.4639% | -0.4991% | -0.2651% | 0.01513% | 0.2606% | -0.2593% | -0.04091% | -0.0208% | -0.6838% | -0.158% | -0.5218% | -0.2639% | -0.5837% | -0.2112% | 0.0007412% | -0.9827% | -0.1116% | -0.2877% | -0.0218% | -0.2513% | -0.3881% | 0.1174% | -0.205% | -0.02759% | -0.001077% | -0.4676% | 0.006916% | -0.5036% | -0.5279% | 0.5771% | -0.727% | -0.606% | -0.5475% | -0.8735% | -0.877% | -0.9% | -0.6366% | -0.4239% | -0.2011% | -1.207% | -1.222% | -1.172% | -1.415% | -0.9035% | -1.182% | -1.341% | -1.11% | -0.8037% | -1.165% | -0.3206% | -0.1458% | -1.116% | -0.7749% | -1.057% | -0.3918% | -0.4905% | -0.4399% | -0.5255% | 60
|  | 70 | 0.1597% | -0.04575% | -0.3081% | -0.1976% | 0.3917% | 0.2257% | -0.09059% | -0.3739% | -0.1559% | -0.4355% | -0.2242% | -0.5416% | -0.1798% | -0.2172% | 0.1332% | 1.848% | -0.286% | -0.3862% | -0.4849% | -0.4378% | -0.2967% | 0.233% | 0.1518% | 0.05493% | 0.12% | -0.02075% |  | 0.009676% | -0.07655% | -0.1161% | -0.3191% | -0.1038% | -0.1658% | -0.548% | -0.3629% | -0.3881% | -0.03008% | -0.3483% |  | -0.1688% | -0.1979% | -0.09334% | -0.0009431% | -0.2454% | -0.02319% | -0.07346% | 0.06883% | 0.3085% | -0.6025% | -0.5276% | -0.3362% | -0.7482% | -0.6921% | -0.4329% | -0.4248% | -0.1412% | 0.02131% | -0.8754% | -0.8859% | -0.5997% | -0.9084% | -0.8111% | -0.7871% | -0.8727% | -0.7651% | -0.6231% | -1.07% | -0.2063% | -0.05933% | -0.7744% | -0.6325% | -0.7178% | -0.2952% | -0.2211% | -0.3573% | -0.2787% | 70
|  | 100 | 0.2177% | -0.09327% | -0.3637% | -0.06789% | 0.3071% | -0.1028% | -0.1841% | -0.5115% | -0.09992% | -0.5641% | -0.03877% | -0.6719% | -0.2537% | -0.4129% | 0.172% | 1.71% | -0.4251% | -0.4409% | -0.2678% | -0.4155% | -0.3968% | 0.04688% | 0.08407% | -0.1544% | -0.2681% | -0.09904% | -0.3667% | 0.09226% | -0.1798% | -0.0851% | -0.2641% | 0.08358% | -0.03916% | -0.2429% | -0.4189% | -0.5932% | 0.09763% | -0.4151% | -0.2445% | 0.04687% | -0.07988% | -0.5149% | -0.2428% | -0.491% | -0.06693% | -0.208% | -0.124% | -0.2727% | -0.6361% | -0.9552% | -0.7461% | -1.137% | -1.03% | -0.9632% | -0.822% | -0.4299% | -0.2262% | -1.19% | -1.17% | -0.7469% | -1.309% |  | -1.203% | -1.364% | -1.11% | -1.075% | -1.432% | -0.484% | -0.3279% | -1.262% | -1.05% | -1.086% | -0.8061% | -1.008% | -0.8805% | -0.7014% | 100
|  | 120 | 0.2976% | -0.06027% | -0.3426% | -0.1275% | 0.2327% | -0.1277% | -0.3868% | -0.6368% | -0.1735% | -0.777% | -0.2677% | -0.7804% | -0.4694% | -0.6928% | 0.383% | 1.656% | -0.7595% | -0.553% | -0.1959% | -0.4561% | -0.4069% | 0.1606% | 0.265% | -0.1056% | -0.1741% | 0.07582% |  | 0.529% | 0.01641% | 0.1557% | -0.2169% | 0.1557% | 0.2196% | 0.183% | -0.1519% |  | 0.3901% | -0.1719% |  | 0.03657% | 0.1013% | -0.5082% | -0.01536% | -0.3371% | 0.2439% | -0.1451% | 0.1603% | -0.3673% | -0.6813% | -0.9572% | -0.4504% | -0.8672% | -0.9508% | -0.7061% | -0.4836% | -0.1165% | 0.006108% | -1.055% | -1.126% | -0.4451% | -1.411% |  | -1.375% | -1.422% | -0.9009% | -0.7299% | -1.192% | -0.1798% | 0.05282% | -1.127% | -0.9265% | -0.899% | -0.5275% | -0.6634% | -0.5957% | -0.3339% | 120
|  | 150 | 0.4231% | -0.2342% | -0.4044% | -0.08621% | 0.3017% | -0.3102% | -0.283% | -0.908% | -0.2511% | -1.011% | -0.06551% | -0.8147% | -0.6889% | -0.5385% | 0.1796% | 1.558% | -0.6294% | -0.7758% | -0.535% | -0.5286% | -0.484% | 0.1501% | 0.1448% | -0.2392% | -0.1191% | -0.1105% | -0.2397% |  | -0.007376% | 0.1173% | -0.1595% | 0.1025% | 0.2089% | 0.02006% | -0.3588% | -0.2149% | 0.3842% | -0.2883% | -0.219% | 0.03053% | 0.06876% | -0.5639% | -0.2059% | -0.468% | 0.0818% | -0.2461% | 0.1016% | -0.4946% | -0.6312% | -1.026% | -0.4028% | -0.9439% | -1.029% | -0.8117% | -0.776% | -0.3983% | -0.1726% | -1.108% | -1.227% | -0.2926% | -1.534% | -1.435% | -1.526% | -1.619% | -0.8861% | -0.8473% | -1.342% | -0.2671% | -0.06864% | -1.268% | -1.062% | -1.036% | -0.6914% | -0.9019% | -0.6574% | -0.5048% | 150
|  | 170 | 0.3847% | -0.1302% | -0.4255% | -0.05377% | 0.2248% | -0.1809% | -0.3054% | -0.685% | -0.06899% | -0.8475% | -0.07435% | -0.7266% | -0.4223% | -0.3985% | 0.1264% | 1.56% | -0.617% |  | -0.6223% | -0.6497% | -0.3912% | 0.1749% | 0.04883% | -0.2834% | -0.1934% | -0.116% |  | 0.3835% | -0.04527% | 0.07416% | -0.242% | 0.05934% | 0.1318% | 0.08754% | -0.4351% |  | 0.2816% | -0.3454% | -0.2419% | -0.04385% | 0.1544% | -0.6719% | -0.1426% | -0.5333% | 0.1083% | -0.3158% | 0.1077% | -0.3966% | -0.7315% | -1.04% | -0.4131% | -0.9581% | -1.018% | -0.877% | -0.9302% | -0.3361% | -0.2554% | -1.113% | -1.075% | -0.3845% | -1.491% |  | -1.485% | -1.702% | -0.9366% | -1.011% | -1.373% | -0.4053% | -0.2049% | -1.209% | -1.136% | -1.101% | -0.7571% | -0.9117% | -0.8127% | -0.5833% | 170
|  | 200 | 0.5159% | -0.1462% | -0.5134% | -0.08174% | 0.2254% | -0.1336% | -0.4173% | -0.7919% | -0.1836% | -0.7689% | -0.2354% | -0.9509% | -0.3446% | -0.4686% | 0.17% | 1.232% | -0.704% | -0.6769% | -0.6839% | -0.6122% | -0.5771% | 0.0205% | -0.02243% | -0.2826% | -0.2919% | -0.2878% | -0.3914% | 0.3803% | -0.4617% | -0.09432% | -0.3994% | -0.1533% | -0.01675% | 0.08311% | -0.5169% | -0.3084% | 0.2152% | -0.4395% |  | -0.2307% | -0.1396% | -0.8729% | -0.3958% | -0.6904% | -0.08077% | -0.4192% | -0.1775% | -0.6827% | -0.9544% | -1.298% | -0.7227% | -1.236% | -1.298% | -1.121% | -1.187% | -0.616% | -0.4794% | -1.239% | -1.395% | -0.743% | -1.851% |  | -1.94% | -1.981% | -1.323% | -1.215% | -1.593% | -0.6207% | -0.411% | -1.542% | -1.274% | -1.432% | -1.094% | -1.154% | -1.072% | -0.8593% | 200
|  | 226 | 0.3657% | -0.1356% | -0.4651% | 0.02305% | 0.212% | -0.2589% | -0.3844% | -0.7146% | -0.1595% | -0.7421% | -0.2178% | -0.8827% | -0.3169% | -0.3174% | 0.3851% | 1.16% | -0.4555% | -0.6965% | -0.372% | -0.8296% | -0.5407% | 0.2683% | -0.1022% | -0.2396% | -0.1849% | -0.09099% |  | 0.3458% | -0.3325% | -0.003874% | -0.2962% | -0.1217% | 0.1129% | 0.08877% | -0.3787% | -0.5302% | 0.3938% | -0.2515% | -0.3434% | -0.1933% | -0.1349% | -0.7807% | -0.3317% | -0.6265% | -0.1758% | -0.5305% | -0.2828% | -0.6845% | -0.9961% | -1.43% | -0.8125% | -1.257% | -1.256% | -1.121% | -1.087% | -0.5703% | -0.4884% | -1.256% | -1.336% | -0.6865% | -2.058% | -2.003% | -1.996% | -1.993% | -1.25% | -1.147% | -1.644% | -0.4899% | -0.2865% | -1.529% | -1.104% | -1.336% | -1.074% | -0.7855% | -0.966% | -0.8026% | 226
|  | average | 0.2961% | -0.08799% | -0.3598% | -0.1437% | 0.2953% | -0.09594% | -0.2576% | -0.6022% | -0.1571% | -0.7239% | -0.2397% | -0.7936% | -0.3316% | -0.4455% | 0.2507% | 1.6% | -0.5938% | -0.5741% | -0.4532% | -0.5536% | -0.4198% | 0.1337% | 0.1038% | -0.1886% | -0.144% | -0.08376% | -0.4204% | 0.2261% | -0.2011% | -0.02702% | -0.3101% | -0.02362% | 0.05653% | -0.1639% | -0.3419% | -0.3871% | 0.2138% | -0.3139% | -0.2874% | -0.05065% | -0.05409% | -0.5042% | -0.167% | -0.4824% | 0.01178% | -0.3052% | -0.08422% | -0.2516% | -0.745% | -0.98% | -0.5539% | -1.003% | -1.019% | -0.8666% | -0.7935% | -0.379% | -0.2245% | -1.13% | -1.18% | -0.6338% | -1.497% | -1.288% | -1.437% | -1.537% | -1.035% | -0.9314% | -1.352% | -0.3717% | -0.1814% | -1.228% | -0.9949% | -1.083% | -0.7046% | -0.7671% | -0.7227% | -0.5737% |  | #DIV/0!
[48 empty] | -0.5305% | -0.5279% | -0.6845% | -0.9961% | -1.43% | -0.8125% | -1.257% | -1.298% | -1.121% | -1.187% | -0.616% | -0.4884% | -1.256% | -1.395% | -1.172% | -2.058% | -2.003% | -1.996% | -1.993% | -1.323% | -1.215% | -1.644% | -0.6207% | -0.411% | -1.542% | -1.274% | -1.432% | -1.094% | -1.154% | -1.072% | -0.8593%
[48 empty] | -0.07346% | 0.1603% | 0.5771% | -0.6025% | -0.5276% | -0.3362% | -0.7482% | -0.6921% | -0.4329% | -0.4248% | -0.1165% | 0.02131% | -0.8754% | -0.8859% | -0.2926% | -0.9084% | -0.8111% | -0.7871% | -0.8727% | -0.7651% | -0.6231% | -1.07% | -0.1798% | 0.05282% | -0.7744% | -0.6325% | -0.7178% | -0.2952% | -0.2211% | -0.3573% | -0.2787%

## Summary_GTR1
GTR1
Average dose (Gy)
| 60 (MeV) | 70 (MeV) | 100 (MeV) | 120 (MeV) | 150 (MeV) | 170 (MeV) | 200 (MeV) | 226 (MeV)
Reference dose (Gy) | 0.95011 | 0.81344 | 0.71141 | 0.68254 | 0.67451 | 0.6748 | 0.67844 | 0.67903
RSD (%) | 0.3716% | 0.2806% | 0.268% | 0.2646% | 0.3521% | 0.316% | 0.3901% | 0.3217%
| Average dose (Gy) | [7 empty] | Difference from Reference value
Date | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226 | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226
(DD/MM/YYYY) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV)
2018-10-16 | 0.95015 | 0.81474 | 0.71295 | 0.68457 | 0.67736 | 0.6774 | 0.68194 | 0.68151 | 0.004674% | 0.1597% | 0.2177% | 0.2976% | 0.4231% | 0.3847% | 0.5159% | 0.3657% | Reference measurements
2018-10-18 | 0.95145 | 0.81307 | 0.71074 | 0.68213 | 0.67293 | 0.67392 | 0.67745 | 0.67811 | 0.1416% | -0.04575% | -0.09327% | -0.06027% | -0.2342% | -0.1302% | -0.1462% | -0.1356%
2018-10-19 | 0.94958 | 0.81094 | 0.70882 | 0.6802 | 0.67178 | 0.67193 | 0.67496 | 0.67587 | -0.05543% | -0.3081% | -0.3637% | -0.3426% | -0.4044% | -0.4255% | -0.5134% | -0.4651%
2018-10-22 | 0.94481 | 0.81184 | 0.71092 | 0.68167 | 0.67393 | 0.67444 | 0.67789 | 0.67918 | -0.5577% | -0.1976% | -0.06789% | -0.1275% | -0.08621% | -0.05377% | -0.08174% | 0.02305%
2018-10-26 | 0.95454 | 0.81663 | 0.71359 | 0.68413 | 0.67654 | 0.67632 | 0.67997 | 0.68047 | 0.4669% | 0.3917% | 0.3071% | 0.2327% | 0.3017% | 0.2248% | 0.2254% | 0.212%
2018-11-05 | 0.95125 | 0.81528 | 0.71067 | 0.68167 | 0.67242 | 0.67358 | 0.67753 | 0.67727 | 0.1208% | 0.2257% | -0.1028% | -0.1277% | -0.3102% | -0.1809% | -0.1336% | -0.2589% | after cyclo PM
2018-11-15 | 0.95002 | 0.81271 | 0.7101 | 0.6799 | 0.6726 | 0.67274 | 0.67561 | 0.67642 | -0.009585% | -0.09059% | -0.1841% | -0.3868% | -0.283% | -0.3054% | -0.4173% | -0.3844% | after cyclo PM
2018-11-27 | 0.94824 | 0.8104 | 0.70777 | 0.67819 | 0.66838 | 0.67018 | 0.67307 | 0.67417 | -0.196% | -0.3739% | -0.5115% | -0.6368% | -0.908% | -0.685% | -0.7919% | -0.7146% | before IC cyclo calib change
2018-12-12 | 0.94855 | 0.81218 | 0.71069 | 0.68135 | 0.67281 | 0.67433 | 0.67719 | 0.67794 | -0.164% | -0.1559% | -0.09992% | -0.1735% | -0.2511% | -0.06899% | -0.1836% | -0.1595% | after IC cyclo calib change
2019-01-08 | 0.94398 | 0.8099 | 0.70739 | 0.67723 | 0.66769 | 0.66908 | 0.67322 | 0.67399 | -0.6452% | -0.4355% | -0.5641% | -0.777% | -1.011% | -0.8475% | -0.7689% | -0.7421%
2019-01-24 | 0.94256 | 0.81162 | 0.71113 | 0.68071 | 0.67407 | 0.6743 | 0.67684 | 0.67755 | -0.7941% | -0.2242% | -0.03877% | -0.2677% | -0.06551% | -0.07435% | -0.2354% | -0.2178%
2019-03-04 | 0.9408 | 0.80904 | 0.70663 | 0.67721 | 0.66901 | 0.6699 | 0.67199 | 0.67303 | -0.98% | -0.5416% | -0.6719% | -0.7804% | -0.8147% | -0.7266% | -0.9509% | -0.8827%
2019-03-17 | 0.95033 | 0.81198 | 0.7096 | 0.67933 | 0.66986 | 0.67195 | 0.6761 | 0.67687 | 0.02298% | -0.1798% | -0.2537% | -0.4694% | -0.6889% | -0.4223% | -0.3446% | -0.3169% |  | v
2019-04-11 | 0.94519 | 0.81168 | 0.70847 | 0.67781 | 0.67088 | 0.67211 | 0.67526 | 0.67687 | -0.5181% | -0.2172% | -0.4129% | -0.6928% | -0.5385% | -0.3985% | -0.4686% | -0.3174%
2019-04-22 | 0.95444 | 0.81453 | 0.71263 | 0.68515 | 0.67572 | 0.67565 | 0.67959 | 0.68164 | 0.4565% | 0.1332% | 0.172% | 0.383% | 0.1796% | 0.1264% | 0.17% | 0.3851%
2019-05-07 | 0.96982 | 0.82848 | 0.72357 | 0.69384 | 0.68501 | 0.68532 | 0.6868 | 0.6869 | 2.075% | 1.848% | 1.71% | 1.656% | 1.558% | 1.56% | 1.232% | 1.16%
2019-05-13 | 0.94181 | 0.81112 | 0.70838 | 0.67735 | 0.67026 | 0.67064 | 0.67366 | 0.67593 | -0.8737% | -0.286% | -0.4251% | -0.7595% | -0.6294% | -0.617% | -0.704% | -0.4555% | new optical solution with new iso
2019-05-13 | 0.94546 | 0.8103 | 0.70827 | 0.67876 | 0.66928 |  | 0.67385 | 0.6743 | -0.4892% | -0.3862% | -0.4409% | -0.553% | -0.7758% |  | -0.6769% | -0.6965%
2019-07-01 | 0.9457 | 0.8095 | 0.7095 | 0.6812 | 0.6709 | 0.6706 | 0.6738 | 0.6765 | -0.4639% | -0.4849% | -0.2678% | -0.1959% | -0.535% | -0.6223% | -0.6839% | -0.372%
2019-07-17 | 0.94537 | 0.80988 | 0.70845 | 0.67942 | 0.67094 | 0.67042 | 0.67429 | 0.67339 | -0.4991% | -0.4378% | -0.4155% | -0.4561% | -0.5286% | -0.6497% | -0.6122% | -0.8296%
2019-08-14 | 0.94759 | 0.81103 | 0.70858 | 0.67976 | 0.67124 | 0.67216 | 0.67452 | 0.67535 | -0.2651% | -0.2967% | -0.3968% | -0.4069% | -0.484% | -0.3912% | -0.5771% | -0.5407%
2019-10-08 | 0.95025 | 0.81534 | 0.71174 | 0.68363 | 0.67552 | 0.67598 | 0.67858 | 0.68085 | 0.01513% | 0.233% | 0.04688% | 0.1606% | 0.1501% | 0.1749% | 0.0205% | 0.2683%
2019-10-14 | 0.95258 | 0.81468 | 0.712 | 0.68435 | 0.67548 | 0.67513 | 0.67829 | 0.67833 | 0.2606% | 0.1518% | 0.08407% | 0.265% | 0.1448% | 0.04883% | -0.02243% | -0.1022%
2019-11-07 | 0.94764 | 0.81389 | 0.71031 | 0.68182 | 0.6729 | 0.67289 | 0.67652 | 0.6774 | -0.2593% | 0.05493% | -0.1544% | -0.1056% | -0.2392% | -0.2834% | -0.2826% | -0.2396%
2020-01-09 | 0.94972 | 0.81442 | 0.7095 | 0.68135 | 0.67371 | 0.67349 | 0.67646 | 0.67777 | -0.04091% | 0.12% | -0.2681% | -0.1741% | -0.1191% | -0.1934% | -0.2919% | -0.1849%
2020-01-18 | 0.94991 | 0.81328 | 0.7107 | 0.68305 | 0.67376 | 0.67402 | 0.67649 | 0.67841 | -0.0208% | -0.02075% | -0.09904% | 0.07582% | -0.1105% | -0.116% | -0.2878% | -0.09099% | new iso position
2020-01-27 | 0.94361 |  | 0.7088 |  | 0.67289 |  | 0.67578 |  | -0.6838% |  | -0.3667% |  | -0.2397% |  | -0.3914%
2020-02-03 | 0.94861 | 0.81352 | 0.71206 | 0.68615 |  | 0.67739 | 0.68102 | 0.68137 | -0.158% | 0.009676% | 0.09226% | 0.529% |  | 0.3835% | 0.3803% | 0.3458% | deflector change
2020-03-06 | 0.94515 | 0.81282 | 0.71013 | 0.68265 | 0.67446 | 0.67449 | 0.67531 | 0.67677 | -0.5218% | -0.07655% | -0.1798% | 0.01641% | -0.007376% | -0.04527% | -0.4617% | -0.3325% | site config update
2020-04-21 | 0.9476 | 0.8125 | 0.7108 | 0.6836 | 0.6753 | 0.6753 | 0.6778 | 0.679 | -0.2639% | -0.1161% | -0.0851% | 0.1557% | 0.1173% | 0.07416% | -0.09432% | -0.003874%
2020-05-22 | 0.94456 | 0.81085 | 0.70953 | 0.68106 | 0.67343 | 0.67317 | 0.67573 | 0.67702 | -0.5837% | -0.3191% | -0.2641% | -0.2169% | -0.1595% | -0.242% | -0.3994% | -0.2962%
2020-06-15 | 0.9481 | 0.8126 | 0.712 | 0.6836 | 0.6752 | 0.6752 | 0.6774 | 0.6782 | -0.2112% | -0.1038% | 0.08358% | 0.1557% | 0.1025% | 0.05934% | -0.1533% | -0.1217%
2020-06-30 | 0.95011 | 0.8121 | 0.71113 | 0.68404 | 0.67592 | 0.67569 | 0.67833 | 0.67979 | 0.0007412% | -0.1658% | -0.03916% | 0.2196% | 0.2089% | 0.1318% | -0.01675% | 0.1129%
2020-07-06 | 0.94077 | 0.80899 | 0.70968 | 0.68379 | 0.67464 | 0.67539 | 0.679 | 0.67963 | -0.9827% | -0.548% | -0.2429% | 0.183% | 0.02006% | 0.08754% | 0.08311% | 0.08877%
2020-07-21 | 0.94905 | 0.81049 | 0.70842 | 0.6815 | 0.67209 | 0.67186 | 0.67493 | 0.67645 | -0.1116% | -0.3629% | -0.4189% | -0.1519% | -0.3588% | -0.4351% | -0.5169% | -0.3787%
2020-07-27 | 0.94737 | 0.81029 | 0.70719 |  | 0.67306 |  | 0.67635 | 0.67543 | -0.2877% | -0.3881% | -0.5932% |  | -0.2149% |  | -0.3084% | -0.5302%
2020-08-18 | 0.9499 | 0.8132 | 0.7121 | 0.6852 | 0.6771 | 0.6767 | 0.6799 | 0.6817 | -0.0218% | -0.03008% | 0.09763% | 0.3901% | 0.3842% | 0.2816% | 0.2152% | 0.3938%
2020-09-08 | 0.94772 | 0.81061 | 0.70845 | 0.68136 | 0.67256 | 0.67247 | 0.67546 | 0.67732 | -0.2513% | -0.3483% | -0.4151% | -0.1719% | -0.2883% | -0.3454% | -0.4395% | -0.2515%
2020-10-12 | 0.94642 |  | 0.70967 |  | 0.67303 | 0.67317 |  | 0.67669 | -0.3881% |  | -0.2445% |  | -0.219% | -0.2419% |  | -0.3434% | cyclo opening
2020-11-03 | 0.95122 | 0.81207 | 0.71174 | 0.68279 | 0.67471 | 0.6745 | 0.67688 | 0.67771 | 0.1174% | -0.1688% | 0.04687% | 0.03657% | 0.03053% | -0.04385% | -0.2307% | -0.1933%
2020-11-18 | 0.94816 | 0.81184 | 0.71084 | 0.68323 | 0.67497 | 0.67584 | 0.67749 | 0.67811 | -0.205% | -0.1979% | -0.07988% | 0.1013% | 0.06876% | 0.1544% | -0.1396% | -0.1349% | Acute deflector change
2021-01-12 | 0.94984 | 0.81269 | 0.70774 | 0.67907 | 0.67071 | 0.67027 | 0.67252 | 0.67373 | -0.02759% | -0.09334% | -0.5149% | -0.5082% | -0.5639% | -0.6719% | -0.8729% | -0.7807%
2021-01-26 | 0.9501 | 0.81344 | 0.70968 | 0.68243 | 0.67312 | 0.67384 | 0.67575 | 0.67677 | -0.001077% | -0.0009431% | -0.2428% | -0.01536% | -0.2059% | -0.1426% | -0.3958% | -0.3317%
2021-02-11 | 0.94566 | 0.81145 | 0.70791 | 0.68024 | 0.67135 | 0.6712 | 0.67376 | 0.67477 | -0.4676% | -0.2454% | -0.491% | -0.3371% | -0.468% | -0.5333% | -0.6904% | -0.6265% | Before cyclo PM
2021-02-14 | 0.95017 | 0.81326 | 0.71093 | 0.6842 | 0.67506 | 0.67553 | 0.67789 | 0.67783 | 0.006916% | -0.02319% | -0.06693% | 0.2439% | 0.0818% | 0.1083% | -0.08077% | -0.1758% | After cyclo PM
2021-02-15 | 0.94532 | 0.81285 | 0.70993 | 0.68155 | 0.67285 | 0.67267 | 0.6756 | 0.67542 | -0.5036% | -0.07346% | -0.208% | -0.1451% | -0.2461% | -0.3158% | -0.4192% | -0.5305%
2021-03-23 | 0.94509 | 0.814 | 0.71052 | 0.68363 | 0.67519 | 0.67553 | 0.67724 | 0.67711 | -0.5279% | 0.06883% | -0.124% | 0.1603% | 0.1016% | 0.1077% | -0.1775% | -0.2828%
2021-04-08 | 0.95559 | 0.81595 | 0.70947 | 0.68003 | 0.67117 | 0.67212 | 0.67381 | 0.67438 | 0.5771% | 0.3085% | -0.2727% | -0.3673% | -0.4946% | -0.3966% | -0.6827% | -0.6845%
2021-05-05 | 0.9432 | 0.80854 | 0.70688 | 0.67789 | 0.67025 | 0.66986 | 0.67196 | 0.67226 | -0.727% | -0.6025% | -0.6361% | -0.6813% | -0.6312% | -0.7315% | -0.9544% | -0.9961%
2021-05-09 | 0.94435 | 0.80915 | 0.70461 | 0.676 | 0.66759 | 0.66778 | 0.66963 | 0.66932 | -0.606% | -0.5276% | -0.9552% | -0.9572% | -1.026% | -1.04% | -1.298% | -1.43% | After IC cyclo replacement
2021-06-17 | 0.94491 | 0.81071 | 0.7061 | 0.67946 | 0.67179 | 0.67201 | 0.67354 | 0.67351 | -0.5475% | -0.3362% | -0.7461% | -0.4504% | -0.4028% | -0.4131% | -0.7227% | -0.8125%
2021-07-13 | 0.94181 | 0.80736 | 0.70331 | 0.67662 | 0.66814 | 0.66833 | 0.67006 | 0.67049 | -0.8735% | -0.7482% | -1.137% | -0.8672% | -0.9439% | -0.9581% | -1.236% | -1.257%
2021-07-14 | 0.94177 | 0.80781 | 0.70408 | 0.67605 | 0.66757 | 0.66793 | 0.66963 | 0.6705 | -0.877% | -0.6921% | -1.03% | -0.9508% | -1.029% | -1.018% | -1.298% | -1.256% | Before IC1 replacement
2021-07-19 | 0.94156 | 0.80992 | 0.70455 | 0.67772 | 0.66903 | 0.66888 | 0.67084 | 0.67141 | -0.9% | -0.4329% | -0.9632% | -0.7061% | -0.8117% | -0.877% | -1.121% | -1.121% | After IC1 replacement
2021-09-14 | 0.94406 | 0.80999 | 0.70556 | 0.67924 | 0.66927 | 0.66852 | 0.67038 | 0.67164 | -0.6366% | -0.4248% | -0.822% | -0.4836% | -0.776% | -0.9302% | -1.187% | -1.087%
2021-10-25 | 0.94608 | 0.8123 | 0.70835 | 0.68174 | 0.67182 | 0.67253 | 0.67426 | 0.67515 | -0.4239% | -0.1412% | -0.4299% | -0.1165% | -0.3983% | -0.3361% | -0.616% | -0.5703% | After BCREU replacement
2022-04-25 | 0.9482 | 0.81362 | 0.7098 | 0.68258 | 0.67334 | 0.67308 | 0.67519 | 0.67571 | -0.2011% | 0.02131% | -0.2262% | 0.006108% | -0.1726% | -0.2554% | -0.4794% | -0.4884%
2022-10-14 | 0.94247 | 0.80838 | 0.70376 | 0.67756 | 0.66879 | 0.66798 | 0.6702 | 0.67124 | -0.8037% | -0.6231% | -1.075% | -0.7299% | -0.8473% | -1.011% | -1.215% | -1.147%
2022-10-18 | 0.93904 | 0.80474 | 0.70122 | 0.6744 | 0.66546 | 0.66553 | 0.66763 | 0.66786 | -1.165% | -1.07% | -1.432% | -1.192% | -1.342% | -1.373% | -1.593% | -1.644%
2022-11-04 | 0.94706 | 0.81177 | 0.70796 | 0.68131 | 0.67271 | 0.67206 | 0.67423 | 0.6757 | -0.3206% | -0.2063% | -0.484% | -0.1798% | -0.2671% | -0.4053% | -0.6207% | -0.4899%
2022-12-03 | 0.94872 | 0.81296 | 0.70907 | 0.6829 | 0.67405 | 0.67342 | 0.67565 | 0.67708 | -0.1458% | -0.05933% | -0.3279% | 0.05282% | -0.06864% | -0.2049% | -0.411% | -0.2865%
2023-03-14 | 0.94274 | 0.8083 | 0.70394 | 0.67621 | 0.66735 | 0.66713 | 0.6698 | 0.67153 | -0.7749% | -0.6325% | -1.05% | -0.9265% | -1.062% | -1.136% | -1.274% | -1.104%
2023-04-21 | 0.94006 | 0.80761 | 0.70368 | 0.6764 | 0.66752 | 0.66737 | 0.66873 | 0.66995 | -1.057% | -0.7178% | -1.086% | -0.899% | -1.036% | -1.101% | -1.432% | -1.336% | New calibration coefficient (0.35% higher than the previous one), blue slab phantom
2023-05-30 | 0.94638 | 0.81104 | 0.70567 | 0.67894 | 0.66985 | 0.66969 | 0.67102 | 0.67174 | -0.3918% | -0.2952% | -0.8061% | -0.5275% | -0.6914% | -0.7571% | -1.094% | -1.074% | New calibration coefficient
2023-06-13 | 0.94545 | 0.81165 | 0.70423 | 0.67801 | 0.66843 | 0.66865 | 0.67061 | 0.67369 | -0.4905% | -0.2211% | -1.008% | -0.6634% | -0.9019% | -0.9117% | -1.154% | -0.7855% | Intervention to LLRF
2023-06-19 | 0.94593 | 0.81054 | 0.70514 | 0.67847 | 0.67007 | 0.66932 | 0.67116 | 0.67247 | -0.4399% | -0.3573% | -0.8805% | -0.5957% | -0.6574% | -0.8127% | -1.072% | -0.966% | Validation after LLRF calibration
2023-07-18 | 0.94511 | 0.81118 | 0.70642 | 0.68026 | 0.6711 | 0.67086 | 0.67261 | 0.67358 | -0.5255% | -0.2787% | -0.7014% | -0.3339% | -0.5048% | -0.5833% | -0.8593% | -0.8026%

## GTR2
[15 empty] | 0.9693 | 60
Gantry: | GTR2 | [13 empty] | 0.82515 | 70
Inställningar: | Fältstorlek 10x10cm2, Spotstorlek 3mm, Spotavstånd 2.5mm, 1681 spots, 0.4 MU/spot, Vattenytan i isocentrum
Mätdjup: | 2 | cm | SW phantom
Jonkammare: | PTW Roos TW34001 #002519
N_D,w,Q0 = | 0.08465 | Gy/nC (Q0 = 150 MeV p+, kalibrerad på Skandionkliniken 2015-03-18 och 2015-04-10)
Elektrometer: | PTW UNIDOSE T10022 SN00470
k_elec = | 1
[47 empty] | Data for Roos IC #3389 | [6 empty] | Data for Roos IC #3389 | [7 empty] | Data for Roos IC #3389
|  |  | REFERENCE | [19 empty] | deflector change | [6 empty] | Cyclo opening | Repetition in the afternoon |  |  | New site config |  | Cyclo opening | fast check | Acute deflector change |  | after cyclo PM |  |  |  | After IC cyclo replacement |  |  | REFERENCE | [6 empty] | After cyclo PM |  | New site config | New Deflector | [6 empty] | ATTN!! Data after postprocesseing | After IC Cyclo gains calibration |  | After ECPM |  |  | Data for Roos IC #3389 | ATTN!! Data after postprocesseing | Data for Roos IC #3389 | New calibration coefficient | Intervention to LLRF | Validation after LLRF calibration | Data for Roos IC #3389
|  |  | 2018-10-18 | 2018-10-19 | 2018-10-22 | 2018-10-25 | 2018-10-26 | 2018-11-05 | 2018-11-06 | 2018-11-27 | 2018-11-27 | 2019-03-05 | 2019-03-18 | 2019-04-22 | 2019-04-22 | 2019-05-28 | 2019-07-24 | 2019-10-01 | 2019-10-14 | 2019-11-14 | 2019-12-07 | 2020-01-09 | 2020-02-03 | 2020-03-05 | 2020-04-09 | 2020-04-21 | 2020-05-27 | 2020-06-17 | 2020-06-30 | 2020-07-06 | 2020-07-06 | 2020-07-27 | 2020-08-07 | 2020-08-24 | 2020-10-06 | 2020-10-11 | 2020-11-04 | 2020-11-18 | 2021-01-21 | 2021-02-14 | 2021-03-23 | 2021-04-27 | 2021-05-05 | 2021-05-09 | 2021-06-15 | 2021-07-20 | 2021-07-27 | 2021-07-29 | 2021-08-02 | 2021-08-03 | 2021-08-05 | 2021-09-02 | 2021-10-04 | 2021-10-17 | 2021-11-01 | 2021-12-10 | 2022-01-04 | 2022-01-06 | 2022-02-15 | 2022-04-05 | 2022-04-25 | 2022-05-17 | 2022-07-01 | 2022-07-16 | 2022-08-10 | 2022-09-10 | 2022-10-09 | 2022-10-14 | 2022-11-18 | 2023-02-04 | 2023-03-14 | 2023-04-22 | 2023-05-30 | 2023-06-13 | 2023-06-19 | 2023-07-17 |  | Mean | SD | SD%
|  | Energi | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose (Gy) | Dose | Dose | Dose (Gy) | Dose | Dose (Gy) | Dose | Dose | Dose | Dose (Gy) | Energi | Dose | Dose | %
|  | (MeV) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | (Gy) | [6 empty] | (Gy) | [7 empty] | (Gy) | [6 empty] | (Gy) | (Gy) |  | (Gy) |  | (Gy) | (Gy) | (Gy) |  | (MeV) | (Gy) | (Gy)
|  | 60 | 0.94574 | 0.94284 | 0.94461 | 0.95093 | 0.94347 |  | 0.94178 |  | 0.94117 | 0.9396 | 0.93981 | 0.94157 | 0.94015 | 0.94439 | 0.94542 | 0.9501 | 0.94479 | 0.94396 | 0.94549 | 0.9434 | 0.9458 | 0.947 | 0.94642 | 0.9462 | 0.94658 | 0.94338 | 0.9448 | 0.9453 | 0.94204 | 0.9476 | 0.9468 | 0.94571 | 0.94631 | 0.94852 | 0.94763 | 0.94318 | 0.94363 | 0.94328 | 0.94329 | 0.93803 | 0.93941 | 0.94002 | 0.9398 | 0.94114 | 0.93565 | 0.9355 | 0.9352 | 0.93585 | 0.9355 | 0.93862 | 0.94358 | 0.93886 | 0.9401 | 0.93866 | 0.93949 | 0.93202 | 0.93441 | 0.93681 | 0.94466 | 0.93739 | 0.9409 | 0.93558 | 0.93945 | 0.93565 | 0.93899 | 0.94425 | 0.94255 | 0.9345 | 0.93883 | 0.93316 | 0.94387 | 0.94649 | 0.94367 | 0.9261 | 60 | 0.94552 | 0.0032209 | 0.3407%
|  | 70 | 0.80654 | 0.80634 | 0.80971 | 0.81095 | 0.80496 |  | 0.80321 |  | 0.80337 | 0.8015 | 0.80287 | 0.80273 | 0.80384 | 0.80629 | 0.80693 | 0.81067 | 0.80798 | 0.80757 | 0.80639 | 0.8042 | 0.8063 | 0.807 | 0.80726 | 0.8087 | 0.807 | 0.80608 | 0.8066 | 0.8109 | 0.80526 | 0.80975 | 0.80864 | 0.80713 | 0.81054 | 0.81122 | 0.80868 | 0.80427 | 0.80661 | 0.80593 | 0.80443 | 0.80171 | 0.804 | 0.80167 | 0.8046 | 0.80499 | 0.7998 | 0.8004 | 0.8013 | 0.8011 | 0.8011 | 0.80389 | 0.80848 | 0.80338 | 0.80438 | 0.80391 | 0.80523 | 0.80043 | 0.79799 | 0.80056 | 0.80567 | 0.79972 | 0.80362 | 0.80074 | 0.80211 | 0.79851 | 0.80318 | 0.80302 | 0.80358 | 0.79698 | 0.8013 | 0.79776 | 0.80418 | 0.80625 | 0.80294 | 0.79307 | 70 | 0.8077 | 0.0025141 | 0.3113%
|  | 100 | 0.70348 | 0.69719 | 0.70252 | 0.70291 | 0.70045 | 0.69707 | 0.69695 | 0.69893 | 0.69749 | 0.69708 | 0.69771 | 0.69651 | 0.69885 | 0.70302 | 0.70211 | 0.7041 | 0.70129 | 0.70194 | 0.699 | 0.6982 |  | 0.7021 | 0.70255 | 0.7041 | 0.70298 | 0.70224 | 0.703 | 0.7066 | 0.70081 | 0.70618 | 0.70658 | 0.70255 | 0.70534 | 0.70558 | 0.70282 | 0.70069 | 0.70133 | 0.7015 | 0.69887 | 0.69618 | 0.6977 | 0.69532 | 0.6988 | 0.69992 | 0.6958 | 0.6955 | 0.6962 | 0.6977 | 0.6971 | 0.69842 | 0.70262 | 0.69812 | 0.70129 | 0.69926 | 0.70106 | 0.69394 | 0.69362 | 0.69768 | 0.70287 | 0.69728 | 0.70151 | 0.69275 | 0.69481 | 0.6933 | 0.69999 | 0.69982 | 0.70026 | 0.69302 | 0.69707 | 0.69442 | 0.70111 | 0.70098 | 0.70032 | 0.69389 | 100 | 0.70131 | 0.0025708 | 0.3666%
|  | 120 | 0.67705 | 0.67475 | 0.68037 | 0.68098 | 0.67431 | 0.67179 | 0.67253 | 0.67258 | 0.67084 | 0.67054 | 0.67249 | 0.67093 | 0.67278 | 0.67715 | 0.67732 | 0.67997 | 0.67686 | 0.67547 | 0.67162 | 0.6753 | 0.6818 | 0.6797 | 0.67841 | 0.6793 | 0.67829 | 0.67722 | 0.6783 | 0.6839 | 0.67775 | 0.6847 | 0.68363 | 0.67905 | 0.68415 | 0.68066 | 0.67716 | 0.67755 | 0.6765 | 0.67789 | 0.67505 | 0.67016 | 0.6744 | 0.6691 | 0.6737 | 0.67488 | 0.67275 | 0.6705 | 0.67295 | 0.6737 | 0.6708 | 0.67297 | 0.67846 | 0.67453 | 0.67812 | 0.67518 | 0.67797 | 0.66885 | 0.66812 | 0.67416 | 0.6778 | 0.67261 | 0.67622 | 0.66702 | 0.66584 | 0.66655 |  | 0.67649 | 0.67598 | 0.66736 | 0.67313 | 0.66855 | 0.67619 | 0.67665 | 0.67584 | 0.67097 | 120 | 0.67749 | 0.0030926 | 0.4565%
|  | 150 | 0.66265 | 0.66351 | 0.66824 | 0.66781 | 0.66353 | 0.66151 | 0.66076 | 0.6596 | 0.65955 | 0.65823 | 0.66563 | 0.65978 | 0.66073 | 0.6647 | 0.6649 | 0.66613 | 0.66845 | 0.66563 | 0.66338 | 0.6622 | 0.6699 | 0.6672 | 0.66638 | 0.6676 | 0.66613 | 0.6649 | 0.6681 | 0.6735 | 0.66542 | 0.67256 | 0.67263 | 0.6696 | 0.67365 | 0.66994 | 0.66554 | 0.66499 | 0.66585 | 0.66586 | 0.66247 | 0.65895 | 0.66171 | 0.65709 | 0.661 | 0.66358 | 0.6634 | 0.6615 | 0.6606 | 0.6618 | 0.6638 | 0.66226 | 0.6656 | 0.66287 | 0.66682 | 0.66462 | 0.66896 | 0.65758 | 0.65762 | 0.66286 | 0.66579 | 0.66222 | 0.66605 | 0.65488 | 0.65586 | 0.65495 | 0.66645 | 0.66629 | 0.66701 | 0.65718 | 0.66128 | 0.65993 | 0.66437 | 0.66651 | 0.66554 | 0.6617 | 150 | 0.66515 | 0.002652 | 0.3987%
|  | 170 | 0.66717 | 0.66537 | 0.66864 | 0.66899 | 0.66387 | 0.66296 | 0.66184 | 0.66227 | 0.66176 | 0.66033 | 0.66569 | 0.65975 | 0.66132 | 0.66885 | 0.66777 | 0.66941 | 0.66762 | 0.66739 | 0.6656 | 0.6629 |  | 0.6689 | 0.66719 | 0.6708 | 0.66945 | 0.6673 | 0.6693 | 0.6757 | 0.6678 | 0.67649 | 0.67533 | 0.67143 | 0.67537 | 0.67093 | 0.66778 | 0.66758 | 0.66718 | 0.66892 | 0.66549 | 0.65963 | 0.66229 | 0.65899 | 0.6631 | 0.66407 | 0.6642 | 0.6631 | 0.6633 | 0.6629 | 0.6675 | 0.66351 | 0.66742 | 0.66628 | 0.66966 | 0.66622 | 0.67069 | 0.66014 | 0.65974 | 0.66415 | 0.66746 | 0.66555 | 0.6687 | 0.65757 | 0.65575 | 0.65582 |  | 0.66831 | 0.66791 | 0.66012 | 0.66317 | 0.66141 | 0.66597 | 0.66817 | 0.6686 | 0.66559 | 170 | 0.66681 | 0.0021739 | 0.326%
|  | 200 | 0.67621 | 0.67478 | 0.68 | 0.67924 | 0.67681 | 0.67582 | 0.67397 | 0.6744 | 0.67376 | 0.66984 | 0.67344 | 0.6721 | 0.6727 | 0.67651 | 0.67726 | 0.67779 | 0.67669 | 0.67639 | 0.67397 | 0.6732 | 0.6799 | 0.6775 | 0.67593 | 0.6788 | 0.67773 | 0.67688 | 0.679 | 0.6845 | 0.67682 | 0.68362 | 0.68246 | 0.67981 | 0.68353 | 0.67886 | 0.67566 | 0.67579 | 0.67532 | 0.67794 | 0.67387 | 0.66878 | 0.67169 | 0.66752 | 0.6722 | 0.67297 | 0.6722 | 0.6719 | 0.673 | 0.6711 | 0.6739 | 0.67204 | 0.67418 | 0.67462 | 0.67696 | 0.67358 | 0.67835 | 0.66778 | 0.66742 | 0.67255 | 0.67566 | 0.67297 | 0.67606 | 0.66364 | 0.66339 | 0.66463 | 0.67567 | 0.6763 | 0.67559 | 0.66747 | 0.67043 | 0.67002 | 0.6741 | 0.67646 | 0.67631 | 0.67171 | 200 | 0.67741 | 0.0021669 | 0.3199%
|  | 226 | 0.6805 | 0.68025 | 0.68471 | 0.68525 | 0.68224 | 0.68027 | 0.67819 | 0.6798 | 0.6795 | 0.67633 | 0.67994 | 0.67666 | 0.67723 | 0.67892 | 0.68168 | 0.68291 | 0.67871 | 0.681 | 0.67891 | 0.6777 | 0.6832 | 0.6829 | 0.68068 | 0.6836 | 0.68247 | 0.68051 | 0.683 | 0.6892 | 0.68069 | 0.68773 | 0.68493 | 0.68208 | 0.68881 | 0.6845 | 0.6816 | 0.6806 | 0.68046 | 0.68209 | 0.67908 | 0.67411 | 0.67587 | 0.67297 | 0.6774 | 0.67981 | 0.6767 | 0.6765 | 0.6764 | 0.6753 | 0.6764 | 0.67577 | 0.68061 | 0.67872 | 0.68087 | 0.67841 | 0.68015 | 0.67172 | 0.67195 | 0.6775 | 0.68041 | 0.67744 | 0.6796 | 0.66978 | 0.66926 | 0.66838 | 0.67958 | 0.68162 | 0.67961 | 0.67167 | 0.67646 | 0.67385 | 0.67794 | 0.68086 | 0.6796 | 0.67544 | 226 | 0.68259 | 0.0023196 | 0.3398%
% Diff from mean | [7 empty] | after cyclo PM | after cyclo PM | before ICcyclo calib change | after ICcyclo calib change | [65 empty] | Energi
|  | Energi (MeV) | 2018-10-18 | 2018-10-19 | 2018-10-22 | 2018-10-25 | 2018-10-26 | 2018-11-05 | 2018-11-06 | 2018-11-27 | 2018-11-27 | 2019-03-05 | 2019-03-18 | 2019-04-22 | 2019-04-22 | 2019-05-28 | 2019-07-24 | 2019-10-01 | 2019-10-14 | 2019-11-14 | 2019-12-07 | 2020-01-09 | 2020-02-03 | 2020-03-05 | 2020-04-09 | 2020-04-21 | 2020-05-27 | 2020-06-17 | 2020-06-18 | 2020-07-06 | 2020-07-06 | 2020-07-27 | 2020-08-07 | 2020-08-24 | 2020-10-06 | 2020-10-11 | 2020-11-04 | 2020-11-18 | 2021-01-21 | 2021-02-14 | 2021-03-23 | 2021-04-27 | 2021-05-05 | 2021-05-09 | 2021-06-15 | 2021-07-20 | 2021-07-27 | 2021-07-29 | 2021-08-02 | 2021-08-03 | 2021-08-05 | 2021-09-02 | 2021-10-04 | 2021-10-17 | 2021-11-01 | 2021-12-10 | 2022-01-04 | 2022-01-06 | 2022-02-15 | 2022-04-05 | 2022-04-25 | 2022-05-17 | 2022-07-01 | 2022-07-16 | 2022-08-10 | 2022-09-10 | 2022-10-09 | 2022-10-14 | 2022-11-18 | 2023-02-04 | 2023-03-14 | 2023-04-22 | 2023-05-30 | 2023-06-13 | 2023-06-19 | 2023-07-17 | (MeV)
|  | 60 | 0.02373% | -0.2834% | -0.09611% | 0.5721% | -0.2163% |  | -0.3949% |  | -0.4599% | -0.6256% | -0.6034% | -0.4177% | -0.568% | -0.1198% | -0.01033% | 0.4845% | -0.0769% | -0.1645% | -0.003457% | -0.224% | 0.0298% | 0.1567% | 0.09542% | 0.0721% | 0.112% | -0.2258% | -0.07596% | -0.02308% | -0.3684% | 0.2197% | 0.1358% | 0.01986% | 0.08389% | 0.3175% | 0.2232% | -0.2474% | -0.1994% | -0.2368% | -0.2359% | -0.792% | -0.6464% | -0.5817% | -0.6048% | -0.4635% | -1.044% | -1.06% | -1.091% | -1.023% | -1.06% | -0.7301% | -0.205% | -0.7045% | -0.5734% | -0.7251% | -0.6379% | -1.428% | -1.175% | -0.9209% | -0.09031% | -0.8595% | -0.4889% | -1.051% | -0.6418% | -1.044% | -0.6902% | -0.1342% | -0.3137% | -1.166% | -0.7071% | -1.307% | -0.1746% | 0.1029% | -0.1953% | -2.054% | 60
|  | 70 | -0.1436% | -0.1687% | 0.2491% | 0.4021% | -0.3388% |  | -0.5565% |  | -0.5357% | -0.7679% | -0.5974% | -0.6155% | -0.4784% | -0.1741% | -0.09596% | 0.3683% | 0.03471% | -0.01557% | -0.1623% | -0.4333% | -0.1734% | -0.08669% | -0.05458% | 0.1238% | -0.08679% | -0.2003% | -0.1362% | 0.3962% | -0.3023% | 0.2542% | 0.1167% | -0.07059% | 0.3521% | 0.4362% | 0.1215% | -0.4243% | -0.1346% | -0.2187% | -0.4048% | -0.7421% | -0.4575% | -0.7464% | -0.3838% | -0.3357% | -0.9781% | -0.9038% | -0.7924% | -0.8172% | -0.8172% | -0.4717% | 0.09642% | -0.5343% | -0.4117% | -0.4687% | -0.3062% | -0.9004% | -1.202% | -0.8844% | -0.2519% | -0.9885% | -0.5046% | -0.8622% | -0.6926% | -1.138% | -0.5594% | -0.5793% | -0.5101% | -1.327% | -0.7923% | -1.23% | -0.4361% | -0.1799% | -0.5894% | -1.811% | 70
|  | 100 | 0.3093% | -0.5875% | 0.1718% | 0.2286% | -0.1222% | -0.6051% | -0.622% | -0.3393% | -0.5451% | -0.6031% | -0.5141% | -0.684% | -0.3515% | 0.2431% | 0.1142% | 0.3973% | -0.003027% | 0.09009% | -0.329% | -0.4437% |  | 0.1124% | 0.1767% | 0.3976% | 0.2372% | 0.1317% | 0.2407% | 0.7541% | -0.07208% | 0.694% | 0.7516% | 0.1764% | 0.5744% | 0.6085% | 0.2152% | -0.08837% | 0.003305% | 0.02621% | -0.3481% | -0.7324% | -0.5152% | -0.8543% | -0.3582% | -0.1991% | -0.7859% | -0.8287% | -0.7289% | -0.515% | -0.6006% | -0.4123% | 0.1871% | -0.4556% | -0.002714% | -0.292% | -0.03551% | -1.051% | -1.096% | -0.5179% | 0.2215% | -0.5754% | 0.02897% | -1.221% | -0.9268% | -1.143% | -0.1887% | -0.2128% | -0.1493% | -1.182% | -0.6053% | -0.9827% | -0.02904% | -0.04752% | -0.1408% | -1.058% | 100
|  | 120 | -0.06582% | -0.4044% | 0.4242% | 0.5153% | -0.4693% | -0.8423% | -0.7318% | -0.7246% | -0.982% | -1.027% | -0.7389% | -0.9693% | -0.6959% | -0.05056% | -0.02544% | 0.3662% | -0.09319% | -0.2987% | -0.8669% | -0.3236% | 0.6359% | 0.3259% | 0.1357% | 0.2669% | 0.1179% | -0.03968% | 0.1193% | 0.9458% | 0.03769% | 1.064% | 0.9066% | 0.2303% | 0.9833% | 0.4676% | -0.04935% | 0.00807% | -0.1464% | 0.05896% | -0.3612% | -1.082% | -0.4558% | -1.239% | -0.5597% | -0.386% | -0.6999% | -1.032% | -0.6704% | -0.5597% | -0.9878% | -0.668% | 0.143% | -0.4366% | 0.0927% | -0.341% | 0.07076% | -1.276% | -1.384% | -0.4923% | 0.04549% | -0.7208% | -0.1882% | -1.546% | -1.72% | -1.615% |  | -0.1485% | -0.2229% | -1.495% | -0.6443% | -1.32% | -0.1926% | -0.1237% | -0.2441% | -0.9628% | 120
|  | 150 | -0.3749% | -0.2466% | 0.4643% | 0.4% | -0.2429% | -0.5462% | -0.6598% | -0.8344% | -0.8408% | -1.04% | 0.07317% | -0.8063% | -0.6648% | -0.06761% | -0.03795% | 0.1473% | 0.497% | 0.07318% | -0.2661% | -0.4431% | 0.7145% | 0.3086% | 0.1852% | 0.3687% | 0.1471% | -0.03756% | 0.4439% | 1.256% | 0.04129% | 1.115% | 1.125% | 0.6698% | 1.278% | 0.72% | 0.0585% | -0.02388% | 0.1061% | 0.1075% | -0.402% | -0.9314% | -0.5171% | -1.211% | -0.6236% | -0.235% | -0.2627% | -0.5484% | -0.6837% | -0.5033% | -0.2026% | -0.4336% | 0.06817% | -0.3423% | 0.251% | -0.07901% | 0.5735% | -1.138% | -1.132% | -0.3445% | 0.09612% | -0.4409% | 0.1358% | -1.544% | -1.396% | -1.532% | 0.1957% | 0.1717% | 0.2796% | -1.197% | -0.5819% | -0.7838% | -0.1169% | 0.2044% | 0.05948% | -0.519% | 150
|  | 170 | 0.0538% | -0.2157% | 0.2747% | 0.3269% | -0.4397% | -0.5775% | -0.7447% | -0.6803% | -0.7568% | -0.9716% | -0.1678% | -1.058% | -0.8233% | 0.3067% | 0.1439% | 0.3907% | 0.1214% | 0.08802% | -0.1812% | -0.5858% |  | 0.314% | 0.05757% | 0.5989% | 0.397% | 0.07473% | 0.374% | 1.334% | 0.1496% | 1.452% | 1.279% | 0.6935% | 1.285% | 0.6179% | 0.1455% | 0.1153% | 0.05629% | 0.3171% | -0.1977% | -1.076% | -0.6775% | -1.173% | -0.5558% | -0.4108% | -0.3909% | -0.5558% | -0.5258% | -0.5858% | 0.104% | -0.4946% | 0.09158% | -0.07844% | 0.4278% | -0.0884% | 0.5826% | -1% | -1.06% | -0.3986% | 0.09854% | -0.188% | 0.2845% | -1.385% | -1.658% | -1.648% |  | 0.2256% | 0.1657% | -1.002% | -0.5461% | -0.809% | -0.1251% | 0.2041% | 0.2684% | -0.1821% | 170
|  | 200 | -0.1763% | -0.3883% | 0.3828% | 0.2704% | -0.0885% | -0.2344% | -0.5071% | -0.4443% | -0.5386% | -1.116% | -0.5858% | -0.7833% | -0.6949% | -0.132% | -0.02203% | 0.0569% | -0.1051% | -0.1497% | -0.5078% | -0.6208% | 0.3683% | 0.01398% | -0.2176% | 0.2059% | 0.04837% | -0.07709% | 0.2354% | 1.047% | -0.08676% | 0.9171% | 0.7457% | 0.3545% | 0.9046% | 0.2141% | -0.2574% | -0.2383% | -0.3075% | 0.07884% | -0.5225% | -1.274% | -0.8436% | -1.459% | -0.7684% | -0.6542% | -0.7684% | -0.8127% | -0.6503% | -0.9308% | -0.5175% | -0.7927% | -0.4767% | -0.4111% | -0.06529% | -0.5646% | 0.1388% | -1.421% | -1.474% | -0.7161% | -0.2579% | -0.6553% | -0.199% | -2.032% | -2.069% | -1.887% | -0.2555% | -0.1634% | -0.2677% | -1.466% | -1.03% | -1.091% | -0.4884% | -0.1401% | -0.1621% | -0.8406% | 200
|  | 226 | -0.3062% | -0.3427% | 0.3103% | 0.3896% | -0.05108% | -0.34% | -0.6456% | -0.4089% | -0.4526% | -0.9168% | -0.3894% | -0.8688% | -0.7852% | -0.5388% | -0.1341% | 0.04652% | -0.5691% | -0.2327% | -0.5391% | -0.7168% | 0.08896% | 0.04501% | -0.2809% | 0.1476% | -0.01841% | -0.3056% | 0.05966% | 0.968% | -0.2795% | 0.7521% | 0.343% | -0.07492% | 0.9102% | 0.2791% | -0.1453% | -0.2913% | -0.3128% | -0.07304% | -0.5149% | -1.242% | -0.9842% | -1.409% | -0.7607% | -0.4071% | -0.8633% | -0.8926% | -0.9072% | -1.068% | -0.9072% | -0.9989% | -0.2911% | -0.5677% | -0.2518% | -0.6135% | -0.3576% | -1.594% | -1.559% | -0.7461% | -0.3195% | -0.7546% | -0.4377% | -1.878% | -1.953% | -2.082% | -0.441% | -0.1423% | -0.4363% | -1.6% | -0.8981% | -1.281% | -0.6811% | -0.254% | -0.4383% | -1.048% | 226
|  |  | -0.085% | -0.3297% | 0.2726% | 0.3881% | -0.2461% | -0.5243% | -0.6078% | -0.572% | -0.639% | -0.8835% | -0.4404% | -0.7754% | -0.6328% | -0.06663% | -0.008464% | 0.2822% | -0.02428% | -0.07623% | -0.357% | -0.4739% | 0.2773% | 0.1487% | 0.01219% | 0.2727% | 0.1193% | -0.08495% | 0.1576% | 0.8347% | -0.1101% | 0.8085% | 0.6754% | 0.2499% | 0.7964% | 0.4576% | 0.03897% | -0.1488% | -0.1169% | 0.007513% | -0.3734% | -0.9839% | -0.6372% | -1.084% | -0.5769% | -0.3864% | -0.7241% | -0.8292% | -0.7563% | -0.7503% | -0.6235% | -0.6252% | -0.04833% | -0.4413% | -0.06666% | -0.3965% | 0.003559% | -1.226% | -1.26% | -0.6276% | -0.05724% | -0.6479% | -0.1711% | -1.44% | -1.382% | -1.511% | -0.3232% | -0.1229% | -0.1818% | -1.304% | -0.7257% | -1.101% | -0.2805% | -0.02923% | -0.1803% | -1.06% | mean
[41 empty] | -0.5225% | -1.274% | -0.9842% | -1.459% | -0.7684% | -0.6542% | -1.044% | -1.06% | -1.091% | -1.068% | -1.06% | -0.9989% | -0.4767% | -0.7045% | -0.5734% | -0.7251% | -0.6379% | -1.594% | -1.559% | -0.9209% | -0.3195% | -0.9885% | -0.5046% | -2.032% | -2.069% | -2.082% | -0.6902% | -0.5793% | -0.5101% | -1.6% | -1.03% | -1.32% | -0.6811% | -0.254% | -0.5894% | -2.054% | min
[41 empty] | -0.1977% | -0.7324% | -0.4558% | -0.5817% | -0.3582% | -0.1991% | -0.2627% | -0.5484% | -0.5258% | -0.5033% | 0.104% | -0.4123% | 0.1871% | -0.07844% | 0.4278% | -0.07901% | 0.5826% | -0.9004% | -1.06% | -0.3445% | 0.2215% | -0.188% | 0.2845% | -0.8622% | -0.6418% | -1.044% | 0.1957% | 0.2256% | 0.2796% | -1.002% | -0.5461% | -0.7838% | -0.02904% | 0.2044% | 0.2684% | -0.1821% | max

## Summary_GTR2
GTR2
Average dose (Gy)
| 60 (MeV) | 70 (MeV) | 100 (MeV) | 120 (MeV) | 150 (MeV) | 170 (MeV) | 200 (MeV) | 226 (MeV)
Reference dose (Gy) | 0.94552 | 0.8077 | 0.70131 | 0.67749 | 0.66515 | 0.66681 | 0.67741 | 0.68259
RSD (%) | 0.3407% | 0.3113% | 0.3666% | 0.4565% | 0.3987% | 0.326% | 0.3199% | 0.3398%
| Average dose (Gy) | [7 empty] | Difference from Reference value
Date | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226 | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226
(DD/MM/YYYY) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV)
2018-10-18 | 0.94574 | 0.80654 | 0.70348 | 0.67705 | 0.66265 | 0.66717 | 0.67621 | 0.6805 | 0.02373% | -0.1436% | 0.3093% | -0.06582% | -0.3749% | 0.0538% | -0.1763% | -0.3062% | Reference measurements
2018-10-19 | 0.94284 | 0.80634 | 0.69719 | 0.67475 | 0.66351 | 0.66537 | 0.67478 | 0.68025 | -0.2834% | -0.1687% | -0.5875% | -0.4044% | -0.2466% | -0.2157% | -0.3883% | -0.3427%
2018-10-22 | 0.94461 | 0.80971 | 0.70252 | 0.68037 | 0.66824 | 0.66864 | 0.68 | 0.68471 | -0.09611% | 0.2491% | 0.1718% | 0.4242% | 0.4643% | 0.2747% | 0.3828% | 0.3103%
2018-10-25 | 0.95093 | 0.81095 | 0.70291 | 0.68098 | 0.66781 | 0.66899 | 0.67924 | 0.68525 | 0.5721% | 0.4021% | 0.2286% | 0.5153% | 0.4% | 0.3269% | 0.2704% | 0.3896%
2018-10-26 | 0.94347 | 0.80496 | 0.70045 | 0.67431 | 0.66353 | 0.66387 | 0.67681 | 0.68224 | -0.2163% | -0.3388% | -0.1222% | -0.4693% | -0.2429% | -0.4397% | -0.0885% | -0.05108%
2018-11-05 |  |  | 0.69707 | 0.67179 | 0.66151 | 0.66296 | 0.67582 | 0.68027 |  |  | -0.6051% | -0.8423% | -0.5462% | -0.5775% | -0.2344% | -0.34% | after cyclo PM
2018-11-06 | 0.94178 | 0.80321 | 0.69695 | 0.67253 | 0.66076 | 0.66184 | 0.67397 | 0.67819 | -0.3949% | -0.5565% | -0.622% | -0.7318% | -0.6598% | -0.7447% | -0.5071% | -0.6456% | after cyclo PM
2018-11-27 |  |  | 0.69893 | 0.67258 | 0.6596 | 0.66227 | 0.6744 | 0.6798 |  |  | -0.3393% | -0.7246% | -0.8344% | -0.6803% | -0.4443% | -0.4089% | before IC cyclo calib change
2018-11-27 | 0.94117 | 0.80337 | 0.69749 | 0.67084 | 0.65955 | 0.66176 | 0.67376 | 0.6795 | -0.4599% | -0.5357% | -0.5451% | -0.982% | -0.8408% | -0.7568% | -0.5386% | -0.4526% | after IC cyclo calib change
2019-03-05 | 0.9396 | 0.8015 | 0.69708 | 0.67054 | 0.65823 | 0.66033 | 0.66984 | 0.67633 | -0.6256% | -0.7679% | -0.6031% | -1.027% | -1.04% | -0.9716% | -1.116% | -0.9168%
2019-03-18 | 0.93981 | 0.80287 | 0.69771 | 0.67249 | 0.66563 | 0.66569 | 0.67344 | 0.67994 | -0.6034% | -0.5974% | -0.5141% | -0.7389% | 0.07317% | -0.1678% | -0.5858% | -0.3894%
2019-04-22 | 0.94157 | 0.80273 | 0.69651 | 0.67093 | 0.65978 | 0.65975 | 0.6721 | 0.67666 | -0.4177% | -0.6155% | -0.684% | -0.9693% | -0.8063% | -1.058% | -0.7833% | -0.8688%
2019-04-22 | 0.94015 | 0.80384 | 0.69885 | 0.67278 | 0.66073 | 0.66132 | 0.6727 | 0.67723 | -0.568% | -0.4784% | -0.3515% | -0.6959% | -0.6648% | -0.8233% | -0.6949% | -0.7852%
2019-05-28 | 0.94439 | 0.80629 | 0.70302 | 0.67715 | 0.6647 | 0.66885 | 0.67651 | 0.67892 | -0.1198% | -0.1741% | 0.2431% | -0.05056% | -0.06761% | 0.3067% | -0.132% | -0.5388%
2019-07-24 | 0.94542 | 0.80693 | 0.70211 | 0.67732 | 0.6649 | 0.66777 | 0.67726 | 0.68168 | -0.01033% | -0.09596% | 0.1142% | -0.02544% | -0.03795% | 0.1439% | -0.02203% | -0.1341%
2019-10-01 | 0.9501 | 0.81067 | 0.7041 | 0.67997 | 0.66613 | 0.66941 | 0.67779 | 0.68291 | 0.4845% | 0.3683% | 0.3973% | 0.3662% | 0.1473% | 0.3907% | 0.0569% | 0.04652%
2019-10-14 | 0.94479 | 0.80798 | 0.70129 | 0.67686 | 0.66845 | 0.66762 | 0.67669 | 0.67871 | -0.0769% | 0.03471% | -0.003027% | -0.09319% | 0.497% | 0.1214% | -0.1051% | -0.5691%
2019-11-14 | 0.94396 | 0.80757 | 0.70194 | 0.67547 | 0.66563 | 0.66739 | 0.67639 | 0.681 | -0.1645% | -0.01557% | 0.09009% | -0.2987% | 0.07318% | 0.08802% | -0.1497% | -0.2327%
2019-12-07 | 0.94549 | 0.80639 | 0.699 | 0.67162 | 0.66338 | 0.6656 | 0.67397 | 0.67891 | -0.003457% | -0.1623% | -0.329% | -0.8669% | -0.2661% | -0.1812% | -0.5078% | -0.5391%
2020-01-09 | 0.9434 | 0.8042 | 0.6982 | 0.6753 | 0.6622 | 0.6629 | 0.6732 | 0.6777 | -0.224% | -0.4333% | -0.4437% | -0.3236% | -0.4431% | -0.5858% | -0.6208% | -0.7168%
2020-02-03 | 0.9458 | 0.8063 |  | 0.6818 | 0.6699 |  | 0.6799 | 0.6832 | 0.0298% | -0.1734% |  | 0.6359% | 0.7145% |  | 0.3683% | 0.08896% | deflector change
2020-03-05 | 0.947 | 0.807 | 0.7021 | 0.6797 | 0.6672 | 0.6689 | 0.6775 | 0.6829 | 0.1567% | -0.08669% | 0.1124% | 0.3259% | 0.3086% | 0.314% | 0.01398% | 0.04501%
2020-04-09 | 0.94642 | 0.80726 | 0.70255 | 0.67841 | 0.66638 | 0.66719 | 0.67593 | 0.68068 | 0.09542% | -0.05458% | 0.1767% | 0.1357% | 0.1852% | 0.05757% | -0.2176% | -0.2809%
2020-04-21 | 0.9462 | 0.8087 | 0.7041 | 0.6793 | 0.6676 | 0.6708 | 0.6788 | 0.6836 | 0.0721% | 0.1238% | 0.3976% | 0.2669% | 0.3687% | 0.5989% | 0.2059% | 0.1476%
2020-05-27 | 0.94658 | 0.807 | 0.70298 | 0.67829 | 0.66613 | 0.66945 | 0.67773 | 0.68247 | 0.112% | -0.08679% | 0.2372% | 0.1179% | 0.1471% | 0.397% | 0.04837% | -0.01841%
2020-06-17 | 0.94338 | 0.80608 | 0.70224 | 0.67722 | 0.6649 | 0.6673 | 0.67688 | 0.68051 | -0.2258% | -0.2003% | 0.1317% | -0.03968% | -0.03756% | 0.07473% | -0.07709% | -0.3056%
2020-06-30 | 0.9448 | 0.8066 | 0.703 | 0.6783 | 0.6681 | 0.6693 | 0.679 | 0.683 | -0.07596% | -0.1362% | 0.2407% | 0.1193% | 0.4439% | 0.374% | 0.2354% | 0.05966%
2020-07-06 | 0.9453 | 0.8109 | 0.7066 | 0.6839 | 0.6735 | 0.6757 | 0.6845 | 0.6892 | -0.02308% | 0.3962% | 0.7541% | 0.9458% | 1.256% | 1.334% | 1.047% | 0.968% | Cyclo opening
2020-07-06 | 0.94204 | 0.80526 | 0.70081 | 0.67775 | 0.66542 | 0.6678 | 0.67682 | 0.68069 | -0.3684% | -0.3023% | -0.07208% | 0.03769% | 0.04129% | 0.1496% | -0.08676% | -0.2795% | Repetition in the afternoon
2020-07-27 | 0.9476 | 0.80975 | 0.70618 | 0.6847 | 0.67256 | 0.67649 | 0.68362 | 0.68773 | 0.2197% | 0.2542% | 0.694% | 1.064% | 1.115% | 1.452% | 0.9171% | 0.7521%
2020-08-07 | 0.9468 | 0.80864 | 0.70658 | 0.68363 | 0.67263 | 0.67533 | 0.68246 | 0.68493 | 0.1358% | 0.1167% | 0.7516% | 0.9066% | 1.125% | 1.279% | 0.7457% | 0.343%
2020-08-24 | 0.94571 | 0.80713 | 0.70255 | 0.67905 | 0.6696 | 0.67143 | 0.67981 | 0.68208 | 0.01986% | -0.07059% | 0.1764% | 0.2303% | 0.6698% | 0.6935% | 0.3545% | -0.07492% | New site config
2020-10-06 | 0.94631 | 0.81054 | 0.70534 | 0.68415 | 0.67365 | 0.67537 | 0.68353 | 0.68881 | 0.08389% | 0.3521% | 0.5744% | 0.9833% | 1.278% | 1.285% | 0.9046% | 0.9102%
2020-10-11 | 0.94852 | 0.81122 | 0.70558 | 0.68066 | 0.66994 | 0.67093 | 0.67886 | 0.6845 | 0.3175% | 0.4362% | 0.6085% | 0.4676% | 0.72% | 0.6179% | 0.2141% | 0.2791% | Cyclo opening
2020-11-04 | 0.94763 | 0.80868 | 0.70282 | 0.67716 | 0.66554 | 0.66778 | 0.67566 | 0.6816 | 0.2232% | 0.1215% | 0.2152% | -0.04935% | 0.0585% | 0.1455% | -0.2574% | -0.1453%
2020-11-18 | 0.94318 | 0.80427 | 0.70069 | 0.67755 | 0.66499 | 0.66758 | 0.67579 | 0.6806 | -0.2474% | -0.4243% | -0.08837% | 0.00807% | -0.02388% | 0.1153% | -0.2383% | -0.2913% | Acute deflector change
2021-01-21 | 0.94363 | 0.80661 | 0.70133 | 0.6765 | 0.66585 | 0.66718 | 0.67532 | 0.68046 | -0.1994% | -0.1346% | 0.003305% | -0.1464% | 0.1061% | 0.05629% | -0.3075% | -0.3128%
2021-02-14 | 0.94328 | 0.80593 | 0.7015 | 0.67789 | 0.66586 | 0.66892 | 0.67794 | 0.68209 | -0.2368% | -0.2187% | 0.02621% | 0.05896% | 0.1075% | 0.3171% | 0.07884% | -0.07304% | after cyclo PM
2021-03-23 | 0.94329 | 0.80443 | 0.69887 | 0.67505 | 0.66247 | 0.66549 | 0.67387 | 0.67908 | -0.2359% | -0.4048% | -0.3481% | -0.3612% | -0.402% | -0.1977% | -0.5225% | -0.5149%
2021-04-27 | 0.93803 | 0.80171 | 0.69618 | 0.67016 | 0.65895 | 0.65963 | 0.66878 | 0.67411 | -0.792% | -0.7421% | -0.7324% | -1.082% | -0.9314% | -1.076% | -1.274% | -1.242%
2021-05-05 | 0.93941 | 0.804 | 0.6977 | 0.6744 | 0.66171 | 0.66229 | 0.67169 | 0.67587 | -0.6464% | -0.4575% | -0.5152% | -0.4558% | -0.5171% | -0.6775% | -0.8436% | -0.9842%
2021-05-09 | 0.94002 | 0.80167 | 0.69532 | 0.6691 | 0.65709 | 0.65899 | 0.66752 | 0.67297 | -0.5817% | -0.7464% | -0.8543% | -1.239% | -1.211% | -1.173% | -1.459% | -1.409% | After IC cyclo replacement
2021-06-15 | 0.9398 | 0.8046 | 0.6988 | 0.6737 | 0.661 | 0.6631 | 0.6722 | 0.6774 | -0.6048% | -0.3838% | -0.3582% | -0.5597% | -0.6236% | -0.5558% | -0.7684% | -0.7607%
2021-07-20 | 0.94114 | 0.80499 | 0.69992 | 0.67488 | 0.66358 | 0.66407 | 0.67297 | 0.67981 | -0.4635% | -0.3357% | -0.1991% | -0.386% | -0.235% | -0.4108% | -0.6542% | -0.4071%
2021-10-04 | 0.94358 | 0.80848 | 0.70262 | 0.67846 | 0.6656 | 0.66742 | 0.67418 | 0.68061 | -0.205% | 0.09642% | 0.1871% | 0.143% | 0.06817% | 0.09158% | -0.4767% | -0.2911% | After BCREU replacement
2022-04-25 | 0.94466 | 0.80567 | 0.70287 | 0.6778 | 0.66579 | 0.66746 | 0.67566 | 0.68041 | -0.09031% | -0.2519% | 0.2215% | 0.04549% | 0.09612% | 0.09854% | -0.2579% | -0.3195%
2022-10-14 | 0.94425 | 0.80302 | 0.69982 | 0.67649 | 0.66629 | 0.66831 | 0.6763 | 0.68162 | -0.1342% | -0.5793% | -0.2128% | -0.1485% | 0.1717% | 0.2256% | -0.1634% | -0.1423%
2022-11-18 | 0.94255 | 0.80358 | 0.70026 | 0.67598 | 0.66701 | 0.66791 | 0.67559 | 0.67961 | -0.3137% | -0.5101% | -0.1493% | -0.2229% | 0.2796% | 0.1657% | -0.2677% | -0.4363%
2023-03-14 | 0.93883 | 0.8013 | 0.69707 | 0.67313 | 0.66128 | 0.66317 | 0.67043 | 0.67646 | -0.7071% | -0.7923% | -0.6053% | -0.6443% | -0.5819% | -0.5461% | -1.03% | -0.8981%
2023-05-30 | 0.94387 | 0.80418 | 0.70111 | 0.67619 | 0.66437 | 0.66597 | 0.6741 | 0.67794 | -0.1746% | -0.4361% | -0.02904% | -0.1926% | -0.1169% | -0.1251% | -0.4884% | -0.6811% | New calibration coefficient
2023-06-13 | 0.94649 | 0.80625 | 0.70098 | 0.67665 | 0.66651 | 0.66817 | 0.67646 | 0.68086 | 0.1029% | -0.1799% | -0.04752% | -0.1237% | 0.2044% | 0.2041% | -0.1401% | -0.254% | Intervention to LLRF
2023-06-19 | 0.94367 | 0.80294 | 0.70032 | 0.67584 | 0.66554 | 0.6686 | 0.67631 | 0.6796 | -0.1953% | -0.5894% | -0.1408% | -0.2441% | 0.05948% | 0.2684% | -0.1621% | -0.4383% | Validation after LLRF calibration

## Summary_Average GTR1 GTR2
Date | Average Difference from reference | Date | Average Difference from reference
(DD/MM/YYYY) | GTR1 | (DD/MM/YYYY) | GTR2
2018-10-16 | 0.2961% | 2018-10-18 | -0.085%
2018-10-18 | -0.08799% | 2018-10-19 | -0.3297%
2018-10-19 | -0.3598% | 2018-10-22 | 0.2726%
2018-10-22 | -0.1437% | 2018-10-25 | 0.3881%
2018-10-26 | 0.2953% | 2018-10-26 | -0.2461%
2018-11-05 | -0.09594% | 2018-11-05 | -0.5243%
2018-11-15 | -0.2576% | 2018-11-06 | -0.6078%
2018-11-27 | -0.6022% | 2018-11-27 | -0.572%
2018-12-12 | -0.1571% | 2018-11-27 | -0.639%
2019-01-08 | -0.7239% | 2019-03-05 | -0.8835%
2019-01-24 | -0.2397% | 2019-03-18 | -0.4404%
2019-03-04 | -0.7936% | 2019-04-22 | -0.7754%
2019-03-17 | -0.3316% | 2019-04-22 | -0.6328%
2019-04-11 | -0.4455% | 2019-05-28 | -0.06663%
2019-04-22 | 0.2507% | 2019-07-24 | -0.008464%
2019-05-07 | 1.6% | 2019-10-01 | 0.2822%
2019-05-13 | -0.5938% | 2019-10-14 | -0.02428%
2019-05-13 | -0.5741% | 2019-11-14 | -0.07623%
2019-07-01 | -0.4532% | 2019-12-07 | -0.357%
2019-07-17 | -0.5536% | 2020-01-09 | -0.4739%
2019-08-14 | -0.4198% | 2020-02-03 | 0.2773%
2019-10-08 | 0.1337% | 2020-03-05 | 0.1487%
2019-10-14 | 0.1038% | 2020-04-09 | 0.01219%
2019-11-07 | -0.1886% | 2020-04-21 | 0.2727%
2020-01-09 | -0.144% | 2020-05-27 | 0.1193%
2020-01-18 | -0.08376% | 2020-06-17 | -0.08495%
2020-01-27 | -0.4204% | 2020-06-30 | 0.1576%
2020-02-03 | 0.2261% | 2020-07-06 | 0.8347%
2020-03-06 | -0.2011% | 2020-07-06 | -0.1101%
2020-04-21 | -0.02702% | 2020-07-27 | 0.8085%
2020-05-22 | -0.3101% | 2020-08-07 | 0.6754%
2020-06-15 | -0.02362% | 2020-08-24 | 0.2499%
2020-06-30 | 0.05653% | 2020-10-06 | 0.7964%
2020-07-06 | -0.1639% | 2020-10-11 | 0.4576%
2020-07-21 | -0.3419% | 2020-11-04 | 0.03897%
2020-07-27 | -0.3871% | 2020-11-18 | -0.1488%
2020-08-18 | 0.2138% | 2021-01-21 | -0.1169%
2020-09-08 | -0.3139% | 2021-02-14 | 0.007513%
2020-10-12 | -0.2874% | 2021-03-23 | -0.3734%
2020-11-03 | -0.05065% | 2021-04-27 | -0.9839%
2020-11-18 | -0.05409% | 2021-05-05 | -0.6372%
2021-01-12 | -0.5042% | 2021-05-09 | -1.084%
2021-01-26 | -0.167% | 2021-06-15 | -0.5769%
2021-02-11 | -0.4824% | 2021-07-20 | -0.3864%
2021-02-14 | 0.01178% | 2021-10-04 | -0.04833%
2021-02-15 | -0.3052% | 2022-04-25 | -0.05724%
2021-03-23 | -0.08422% | 2022-10-14 | -0.1229%
2021-04-08 | -0.2516% | 2022-11-18 | -0.1818%
2021-05-05 | -0.745% | 2023-03-14 | -0.7257%
2021-05-09 | -0.98% | 2023-05-30 | -0.2805%
2021-06-17 | -0.5539% | 2023-06-13 | -0.02923%
2021-07-13 | -1.003% | 2023-06-19 | -0.1803%
2021-07-14 | -1.019%
2021-07-19 | -0.8666%
2021-09-14 | -0.7935%
2021-10-25 | -0.379%
2022-04-25 | -0.2245%
2022-10-14 | -0.9314%
2022-10-18 | -1.352%
2022-11-04 | -0.3717%
2022-12-03 | -0.1814%
2023-03-14 | -0.9949%
2023-04-21 | -1.083%
2023-05-30 | -0.7046%
2023-06-13 | -0.7671%
2023-06-19 | -0.7227%

## Summary_GTR1 (2) (hidden)
GTR1
Average dose (Gy)
| 60 (MeV) | 70 (MeV) | 100 (MeV) | 120 (MeV) | 150 (MeV) | 170 (MeV) | 200 (MeV) | 226 (MeV)
Reference dose (Gy) | 0.95011 | 0.81344 | 0.71141 | 0.68254 | 0.67451 | 0.6748 | 0.67844 | 0.67903
RSD (%) | 0.3716% | 0.2806% | 0.268% | 0.2646% | 0.3521% | 0.316% | 0.3901% | 0.3217%
| Average dose (Gy) | [7 empty] | Difference from Reference value
Date | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226 | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226
(DD/MM/YYYY) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV)
2018-10-16 | 0.95015 | 0.81474 | 0.71295 | 0.68457 | 0.67736 | 0.6774 | 0.68194 | 0.68151 | 0.004674% | 0.1597% | 0.2177% | 0.2976% | 0.4231% | 0.3847% | 0.5159% | 0.3657% | Reference measurements
2018-10-18 | 0.95145 | 0.81307 | 0.71074 | 0.68213 | 0.67293 | 0.67392 | 0.67745 | 0.67811 | 0.1416% | -0.04575% | -0.09327% | -0.06027% | -0.2342% | -0.1302% | -0.1462% | -0.1356%
2018-10-19 | 0.94958 | 0.81094 | 0.70882 | 0.6802 | 0.67178 | 0.67193 | 0.67496 | 0.67587 | -0.05543% | -0.3081% | -0.3637% | -0.3426% | -0.4044% | -0.4255% | -0.5134% | -0.4651%
2018-10-22 | 0.94481 | 0.81184 | 0.71092 | 0.68167 | 0.67393 | 0.67444 | 0.67789 | 0.67918 | -0.5577% | -0.1976% | -0.06789% | -0.1275% | -0.08621% | -0.05377% | -0.08174% | 0.02305%
2018-10-26 | 0.95454 | 0.81663 | 0.71359 | 0.68413 | 0.67654 | 0.67632 | 0.67997 | 0.68047 | 0.4669% | 0.3917% | 0.3071% | 0.2327% | 0.3017% | 0.2248% | 0.2254% | 0.212%
2018-11-05 | 0.95125 | 0.81528 | 0.71067 | 0.68167 | 0.67242 | 0.67358 | 0.67753 | 0.67727 | 0.1208% | 0.2257% | -0.1028% | -0.1277% | -0.3102% | -0.1809% | -0.1336% | -0.2589% | after cyclo PM
2018-11-15 | 0.95002 | 0.81271 | 0.7101 | 0.6799 | 0.6726 | 0.67274 | 0.67561 | 0.67642 | -0.009585% | -0.09059% | -0.1841% | -0.3868% | -0.283% | -0.3054% | -0.4173% | -0.3844% | after cyclo PM
2018-11-27 | 0.94824 | 0.8104 | 0.70777 | 0.67819 | 0.66838 | 0.67018 | 0.67307 | 0.67417 | -0.196% | -0.3739% | -0.5115% | -0.6368% | -0.908% | -0.685% | -0.7919% | -0.7146% | before IC cyclo calib change
2018-12-12 | 0.94855 | 0.81218 | 0.71069 | 0.68135 | 0.67281 | 0.67433 | 0.67719 | 0.67794 | -0.164% | -0.1559% | -0.09992% | -0.1735% | -0.2511% | -0.06899% | -0.1836% | -0.1595% | after IC cyclo calib change
2019-01-08 | 0.94398 | 0.8099 | 0.70739 | 0.67723 | 0.66769 | 0.66908 | 0.67322 | 0.67399 | -0.6452% | -0.4355% | -0.5641% | -0.777% | -1.011% | -0.8475% | -0.7689% | -0.7421%
2019-01-24 | 0.94256 | 0.81162 | 0.71113 | 0.68071 | 0.67407 | 0.6743 | 0.67684 | 0.67755 | -0.7941% | -0.2242% | -0.03877% | -0.2677% | -0.06551% | -0.07435% | -0.2354% | -0.2178%
2019-03-04 | 0.9408 | 0.80904 | 0.70663 | 0.67721 | 0.66901 | 0.6699 | 0.67199 | 0.67303 | -0.98% | -0.5416% | -0.6719% | -0.7804% | -0.8147% | -0.7266% | -0.9509% | -0.8827%
2019-03-17 | 0.95033 | 0.81198 | 0.7096 | 0.67933 | 0.66986 | 0.67195 | 0.6761 | 0.67687 | 0.02298% | -0.1798% | -0.2537% | -0.4694% | -0.6889% | -0.4223% | -0.3446% | -0.3169% |  | v
2019-04-11 | 0.94519 | 0.81168 | 0.70847 | 0.67781 | 0.67088 | 0.67211 | 0.67526 | 0.67687 | -0.5181% | -0.2172% | -0.4129% | -0.6928% | -0.5385% | -0.3985% | -0.4686% | -0.3174%
2019-04-22 | 0.95444 | 0.81453 | 0.71263 | 0.68515 | 0.67572 | 0.67565 | 0.67959 | 0.68164 | 0.4565% | 0.1332% | 0.172% | 0.383% | 0.1796% | 0.1264% | 0.17% | 0.3851%
2019-05-07 | 0.96982 | 0.82848 | 0.72357 | 0.69384 | 0.68501 | 0.68532 | 0.6868 | 0.6869 | 2.075% | 1.848% | 1.71% | 1.656% | 1.558% | 1.56% | 1.232% | 1.16%
2019-05-13 | 0.94181 | 0.81112 | 0.70838 | 0.67735 | 0.67026 | 0.67064 | 0.67366 | 0.67593 | -0.8737% | -0.286% | -0.4251% | -0.7595% | -0.6294% | -0.617% | -0.704% | -0.4555% | new optical solution with new iso
2019-05-13 | 0.94546 | 0.8103 | 0.70827 | 0.67876 | 0.66928 |  | 0.67385 | 0.6743 | -0.4892% | -0.3862% | -0.4409% | -0.553% | -0.7758% |  | -0.6769% | -0.6965%
2019-07-01 | 0.9457 | 0.8095 | 0.7095 | 0.6812 | 0.6709 | 0.6706 | 0.6738 | 0.6765 | -0.4639% | -0.4849% | -0.2678% | -0.1959% | -0.535% | -0.6223% | -0.6839% | -0.372%
2019-07-17 | 0.94537 | 0.80988 | 0.70845 | 0.67942 | 0.67094 | 0.67042 | 0.67429 | 0.67339 | -0.4991% | -0.4378% | -0.4155% | -0.4561% | -0.5286% | -0.6497% | -0.6122% | -0.8296%
2019-08-14 | 0.94759 | 0.81103 | 0.70858 | 0.67976 | 0.67124 | 0.67216 | 0.67452 | 0.67535 | -0.2651% | -0.2967% | -0.3968% | -0.4069% | -0.484% | -0.3912% | -0.5771% | -0.5407%
2019-10-08 | 0.95025 | 0.81534 | 0.71174 | 0.68363 | 0.67552 | 0.67598 | 0.67858 | 0.68085 | 0.01513% | 0.233% | 0.04688% | 0.1606% | 0.1501% | 0.1749% | 0.0205% | 0.2683%
2019-10-14 | 0.95258 | 0.81468 | 0.712 | 0.68435 | 0.67548 | 0.67513 | 0.67829 | 0.67833 | 0.2606% | 0.1518% | 0.08407% | 0.265% | 0.1448% | 0.04883% | -0.02243% | -0.1022%
2019-11-07 | 0.94764 | 0.81389 | 0.71031 | 0.68182 | 0.6729 | 0.67289 | 0.67652 | 0.6774 | -0.2593% | 0.05493% | -0.1544% | -0.1056% | -0.2392% | -0.2834% | -0.2826% | -0.2396%
2020-01-09 | 0.94972 | 0.81442 | 0.7095 | 0.68135 | 0.67371 | 0.67349 | 0.67646 | 0.67777 | -0.04091% | 0.12% | -0.2681% | -0.1741% | -0.1191% | -0.1934% | -0.2919% | -0.1849%
2020-01-18 | 0.94991 | 0.81328 | 0.7107 | 0.68305 | 0.67376 | 0.67402 | 0.67649 | 0.67841 | -0.0208% | -0.02075% | -0.09904% | 0.07582% | -0.1105% | -0.116% | -0.2878% | -0.09099% | new iso position
2020-01-27 | 0.94361 |  | 0.7088 |  | 0.67289 |  | 0.67578 |  | -0.6838% |  | -0.3667% |  | -0.2397% |  | -0.3914%
2020-02-03 | 0.94861 | 0.81352 | 0.71206 | 0.68615 |  | 0.67739 | 0.68102 | 0.68137 | -0.158% | 0.009676% | 0.09226% | 0.529% |  | 0.3835% | 0.3803% | 0.3458% | deflector change
2020-03-06 | 0.94515 | 0.81282 | 0.71013 | 0.68265 | 0.67446 | 0.67449 | 0.67531 | 0.67677 | -0.5218% | -0.07655% | -0.1798% | 0.01641% | -0.007376% | -0.04527% | -0.4617% | -0.3325% | site config update
2020-04-21 | 0.9476 | 0.8125 | 0.7108 | 0.6836 | 0.6753 | 0.6753 | 0.6778 | 0.679 | -0.2639% | -0.1161% | -0.0851% | 0.1557% | 0.1173% | 0.07416% | -0.09432% | -0.003874%
2020-05-22 | 0.94456 | 0.81085 | 0.70953 | 0.68106 | 0.67343 | 0.67317 | 0.67573 | 0.67702 | -0.5837% | -0.3191% | -0.2641% | -0.2169% | -0.1595% | -0.242% | -0.3994% | -0.2962%
2020-06-15 | 0.9481 | 0.8126 | 0.712 | 0.6836 | 0.6752 | 0.6752 | 0.6774 | 0.6782 | -0.2112% | -0.1038% | 0.08358% | 0.1557% | 0.1025% | 0.05934% | -0.1533% | -0.1217%
2020-06-30 | 0.95011 | 0.8121 | 0.71113 | 0.68404 | 0.67592 | 0.67569 | 0.67833 | 0.67979 | 0.0007412% | -0.1658% | -0.03916% | 0.2196% | 0.2089% | 0.1318% | -0.01675% | 0.1129%
2020-07-06 | 0.94077 | 0.80899 | 0.70968 | 0.68379 | 0.67464 | 0.67539 | 0.679 | 0.67963 | -0.9827% | -0.548% | -0.2429% | 0.183% | 0.02006% | 0.08754% | 0.08311% | 0.08877%
2020-07-21 | 0.94905 | 0.81049 | 0.70842 | 0.6815 | 0.67209 | 0.67186 | 0.67493 | 0.67645 | -0.1116% | -0.3629% | -0.4189% | -0.1519% | -0.3588% | -0.4351% | -0.5169% | -0.3787%
2020-07-27 | 0.94737 | 0.81029 | 0.70719 |  | 0.67306 |  | 0.67635 | 0.67543 | -0.2877% | -0.3881% | -0.5932% |  | -0.2149% |  | -0.3084% | -0.5302%
2020-08-18 | 0.9499 | 0.8132 | 0.7121 | 0.6852 | 0.6771 | 0.6767 | 0.6799 | 0.6817 | -0.0218% | -0.03008% | 0.09763% | 0.3901% | 0.3842% | 0.2816% | 0.2152% | 0.3938%
2020-09-08 | 0.94772 | 0.81061 | 0.70845 | 0.68136 | 0.67256 | 0.67247 | 0.67546 | 0.67732 | -0.2513% | -0.3483% | -0.4151% | -0.1719% | -0.2883% | -0.3454% | -0.4395% | -0.2515%
2020-10-12 | 0.94642 |  | 0.70967 |  | 0.67303 | 0.67317 |  | 0.67669 | -0.3881% |  | -0.2445% |  | -0.219% | -0.2419% |  | -0.3434% | cyclo opening
2020-11-03 | 0.95122 | 0.81207 | 0.71174 | 0.68279 | 0.67471 | 0.6745 | 0.67688 | 0.67771 | 0.1174% | -0.1688% | 0.04687% | 0.03657% | 0.03053% | -0.04385% | -0.2307% | -0.1933%
2020-11-18 | 0.94816 | 0.81184 | 0.71084 | 0.68323 | 0.67497 | 0.67584 | 0.67749 | 0.67811 | -0.205% | -0.1979% | -0.07988% | 0.1013% | 0.06876% | 0.1544% | -0.1396% | -0.1349% | Acute deflector change
2021-01-12 | 0.94984 | 0.81269 | 0.70774 | 0.67907 | 0.67071 | 0.67027 | 0.67252 | 0.67373 | -0.02759% | -0.09334% | -0.5149% | -0.5082% | -0.5639% | -0.6719% | -0.8729% | -0.7807%
2021-01-26 | 0.9501 | 0.81344 | 0.70968 | 0.68243 | 0.67312 | 0.67384 | 0.67575 | 0.67677 | -0.001077% | -0.0009431% | -0.2428% | -0.01536% | -0.2059% | -0.1426% | -0.3958% | -0.3317%
2021-02-11 | 0.94566 | 0.81145 | 0.70791 | 0.68024 | 0.67135 | 0.6712 | 0.67376 | 0.67477 | -0.4676% | -0.2454% | -0.491% | -0.3371% | -0.468% | -0.5333% | -0.6904% | -0.6265% | Before cyclo PM
2021-02-14 | 0.95017 | 0.81326 | 0.71093 | 0.6842 | 0.67506 | 0.67553 | 0.67789 | 0.67783 | 0.006916% | -0.02319% | -0.06693% | 0.2439% | 0.0818% | 0.1083% | -0.08077% | -0.1758% | After cyclo PM
2021-02-15 | 0.94532 | 0.81285 | 0.70993 | 0.68155 | 0.67285 | 0.67267 | 0.6756 | 0.67542 | -0.5036% | -0.07346% | -0.208% | -0.1451% | -0.2461% | -0.3158% | -0.4192% | -0.5305%
2021-03-23 | 0.94509 | 0.814 | 0.71052 | 0.68363 | 0.67519 | 0.67553 | 0.67724 | 0.67711 | -0.5279% | 0.06883% | -0.124% | 0.1603% | 0.1016% | 0.1077% | -0.1775% | -0.2828%
2021-04-08 | 0.95559 | 0.81595 | 0.70947 | 0.68003 | 0.67117 | 0.67212 | 0.67381 | 0.67438 | 0.5771% | 0.3085% | -0.2727% | -0.3673% | -0.4946% | -0.3966% | -0.6827% | -0.6845%
2021-05-05 | 0.9432 | 0.80854 | 0.70688 | 0.67789 | 0.67025 | 0.66986 | 0.67196 | 0.67226 | -0.727% | -0.6025% | -0.6361% | -0.6813% | -0.6312% | -0.7315% | -0.9544% | -0.9961%
2021-05-09 | 0.94435 | 0.80915 | 0.70461 | 0.676 | 0.66759 | 0.66778 | 0.66963 | 0.66932 | -0.606% | -0.5276% | -0.9552% | -0.9572% | -1.026% | -1.04% | -1.298% | -1.43% | After IC cyclo replacement
2021-06-17 | 0.94491 | 0.81071 | 0.7061 | 0.67946 | 0.67179 | 0.67201 | 0.67354 | 0.67351 | -0.5475% | -0.3362% | -0.7461% | -0.4504% | -0.4028% | -0.4131% | -0.7227% | -0.8125%
2021-07-13 | 0.94181 | 0.80736 | 0.70331 | 0.67662 | 0.66814 | 0.66833 | 0.67006 | 0.67049 | -0.8735% | -0.7482% | -1.137% | -0.8672% | -0.9439% | -0.9581% | -1.236% | -1.257%
2021-07-14 | 0.94177 | 0.80781 | 0.70408 | 0.67605 | 0.66757 | 0.66793 | 0.66963 | 0.6705 | -0.877% | -0.6921% | -1.03% | -0.9508% | -1.029% | -1.018% | -1.298% | -1.256% | Before IC1 replacement
2021-07-19 | 0.94156 | 0.80992 | 0.70455 | 0.67772 | 0.66903 | 0.66888 | 0.67084 | 0.67141 | -0.9% | -0.4329% | -0.9632% | -0.7061% | -0.8117% | -0.877% | -1.121% | -1.121% | After IC1 replacement
2021-09-14 | 0.94406 | 0.80999 | 0.70556 | 0.67924 | 0.66927 | 0.66852 | 0.67038 | 0.67164 | -0.6366% | -0.4248% | -0.822% | -0.4836% | -0.776% | -0.9302% | -1.187% | -1.087%
2021-10-25 | 0.94608 | 0.8123 | 0.70835 | 0.68174 | 0.67182 | 0.67253 | 0.67426 | 0.67515 | -0.4239% | -0.1412% | -0.4299% | -0.1165% | -0.3983% | -0.3361% | -0.616% | -0.5703% | After BCREU replacement
2022-04-25 | 0.9482 | 0.81362 | 0.7098 | 0.68258 | 0.67334 | 0.67308 | 0.67519 | 0.67571 | -0.2011% | 0.02131% | -0.2262% | 0.006108% | -0.1726% | -0.2554% | -0.4794% | -0.4884%
2022-10-14 | 0.94247 | 0.80838 | 0.70376 | 0.67756 | 0.66879 | 0.66798 | 0.6702 | 0.67124 | -0.8037% | -0.6231% | -1.075% | -0.7299% | -0.8473% | -1.011% | -1.215% | -1.147%
2022-10-18 | 0.93904 | 0.80474 | 0.70122 | 0.6744 | 0.66546 | 0.66553 | 0.66763 | 0.66786 | -1.165% | -1.07% | -1.432% | -1.192% | -1.342% | -1.373% | -1.593% | -1.644%
2022-11-04 | 0.94706 | 0.81177 | 0.70796 | 0.68131 | 0.67271 | 0.67206 | 0.67423 | 0.6757 | -0.3206% | -0.2063% | -0.484% | -0.1798% | -0.2671% | -0.4053% | -0.6207% | -0.4899%
2022-12-03 | 0.94872 | 0.81296 | 0.70907 | 0.6829 | 0.67405 | 0.67342 | 0.67565 | 0.67708 | -0.1458% | -0.05933% | -0.3279% | 0.05282% | -0.06864% | -0.2049% | -0.411% | -0.2865%
2023-03-14 | 0.94274 | 0.8083 | 0.70394 | 0.67621 | 0.66735 | 0.66713 | 0.6698 | 0.67153 | -0.7749% | -0.6325% | -1.05% | -0.9265% | -1.062% | -1.136% | -1.274% | -1.104%

## Summary_GTR2 (2) (hidden)
GTR2
Average dose (Gy)
| 60 (MeV) | 70 (MeV) | 100 (MeV) | 120 (MeV) | 150 (MeV) | 170 (MeV) | 200 (MeV) | 226 (MeV)
Reference dose (Gy) | 0.94552 | 0.8077 | 0.70131 | 0.67749 | 0.66515 | 0.66681 | 0.67741 | 0.68259
RSD (%) | 0.3407% | 0.3113% | 0.3666% | 0.4565% | 0.3987% | 0.326% | 0.3199% | 0.3398%
| Average dose (Gy) | [7 empty] | Difference from Reference value
Date | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226 | 60 | 70 | 100 | 120 | 150 | 170 | 200 | 226
(DD/MM/YYYY) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV) | (MeV)
2018-10-18 | 0.94574 | 0.80654 | 0.70348 | 0.67705 | 0.66265 | 0.66717 | 0.67621 | 0.6805 | 0.02373% | -0.1436% | 0.3093% | -0.06582% | -0.3749% | 0.0538% | -0.1763% | -0.3062% | Reference measurements
2018-10-19 | 0.94284 | 0.80634 | 0.69719 | 0.67475 | 0.66351 | 0.66537 | 0.67478 | 0.68025 | -0.2834% | -0.1687% | -0.5875% | -0.4044% | -0.2466% | -0.2157% | -0.3883% | -0.3427%
2018-10-22 | 0.94461 | 0.80971 | 0.70252 | 0.68037 | 0.66824 | 0.66864 | 0.68 | 0.68471 | -0.09611% | 0.2491% | 0.1718% | 0.4242% | 0.4643% | 0.2747% | 0.3828% | 0.3103%
2018-10-25 | 0.95093 | 0.81095 | 0.70291 | 0.68098 | 0.66781 | 0.66899 | 0.67924 | 0.68525 | 0.5721% | 0.4021% | 0.2286% | 0.5153% | 0.4% | 0.3269% | 0.2704% | 0.3896%
2018-10-26 | 0.94347 | 0.80496 | 0.70045 | 0.67431 | 0.66353 | 0.66387 | 0.67681 | 0.68224 | -0.2163% | -0.3388% | -0.1222% | -0.4693% | -0.2429% | -0.4397% | -0.0885% | -0.05108%
2018-11-05 |  |  | 0.69707 | 0.67179 | 0.66151 | 0.66296 | 0.67582 | 0.68027 |  |  | -0.6051% | -0.8423% | -0.5462% | -0.5775% | -0.2344% | -0.34% | after cyclo PM
2018-11-06 | 0.94178 | 0.80321 | 0.69695 | 0.67253 | 0.66076 | 0.66184 | 0.67397 | 0.67819 | -0.3949% | -0.5565% | -0.622% | -0.7318% | -0.6598% | -0.7447% | -0.5071% | -0.6456% | after cyclo PM
2018-11-27 |  |  | 0.69893 | 0.67258 | 0.6596 | 0.66227 | 0.6744 | 0.6798 |  |  | -0.3393% | -0.7246% | -0.8344% | -0.6803% | -0.4443% | -0.4089% | before IC cyclo calib change
2018-11-27 | 0.94117 | 0.80337 | 0.69749 | 0.67084 | 0.65955 | 0.66176 | 0.67376 | 0.6795 | -0.4599% | -0.5357% | -0.5451% | -0.982% | -0.8408% | -0.7568% | -0.5386% | -0.4526% | after IC cyclo calib change
2019-03-05 | 0.9396 | 0.8015 | 0.69708 | 0.67054 | 0.65823 | 0.66033 | 0.66984 | 0.67633 | -0.6256% | -0.7679% | -0.6031% | -1.027% | -1.04% | -0.9716% | -1.116% | -0.9168%
2019-03-18 | 0.93981 | 0.80287 | 0.69771 | 0.67249 | 0.66563 | 0.66569 | 0.67344 | 0.67994 | -0.6034% | -0.5974% | -0.5141% | -0.7389% | 0.07317% | -0.1678% | -0.5858% | -0.3894%
2019-04-22 | 0.94157 | 0.80273 | 0.69651 | 0.67093 | 0.65978 | 0.65975 | 0.6721 | 0.67666 | -0.4177% | -0.6155% | -0.684% | -0.9693% | -0.8063% | -1.058% | -0.7833% | -0.8688%
2019-04-22 | 0.94015 | 0.80384 | 0.69885 | 0.67278 | 0.66073 | 0.66132 | 0.6727 | 0.67723 | -0.568% | -0.4784% | -0.3515% | -0.6959% | -0.6648% | -0.8233% | -0.6949% | -0.7852%
2019-05-28 | 0.94439 | 0.80629 | 0.70302 | 0.67715 | 0.6647 | 0.66885 | 0.67651 | 0.67892 | -0.1198% | -0.1741% | 0.2431% | -0.05056% | -0.06761% | 0.3067% | -0.132% | -0.5388%
2019-07-24 | 0.94542 | 0.80693 | 0.70211 | 0.67732 | 0.6649 | 0.66777 | 0.67726 | 0.68168 | -0.01033% | -0.09596% | 0.1142% | -0.02544% | -0.03795% | 0.1439% | -0.02203% | -0.1341%
2019-10-01 | 0.9501 | 0.81067 | 0.7041 | 0.67997 | 0.66613 | 0.66941 | 0.67779 | 0.68291 | 0.4845% | 0.3683% | 0.3973% | 0.3662% | 0.1473% | 0.3907% | 0.0569% | 0.04652%
2019-10-14 | 0.94479 | 0.80798 | 0.70129 | 0.67686 | 0.66845 | 0.66762 | 0.67669 | 0.67871 | -0.0769% | 0.03471% | -0.003027% | -0.09319% | 0.497% | 0.1214% | -0.1051% | -0.5691%
2019-11-14 | 0.94396 | 0.80757 | 0.70194 | 0.67547 | 0.66563 | 0.66739 | 0.67639 | 0.681 | -0.1645% | -0.01557% | 0.09009% | -0.2987% | 0.07318% | 0.08802% | -0.1497% | -0.2327%
2019-12-07 | 0.94549 | 0.80639 | 0.699 | 0.67162 | 0.66338 | 0.6656 | 0.67397 | 0.67891 | -0.003457% | -0.1623% | -0.329% | -0.8669% | -0.2661% | -0.1812% | -0.5078% | -0.5391%
2020-01-09 | 0.9434 | 0.8042 | 0.6982 | 0.6753 | 0.6622 | 0.6629 | 0.6732 | 0.6777 | -0.224% | -0.4333% | -0.4437% | -0.3236% | -0.4431% | -0.5858% | -0.6208% | -0.7168%
2020-02-03 | 0.9458 | 0.8063 |  | 0.6818 | 0.6699 |  | 0.6799 | 0.6832 | 0.0298% | -0.1734% |  | 0.6359% | 0.7145% |  | 0.3683% | 0.08896% | deflector change
2020-03-05 | 0.947 | 0.807 | 0.7021 | 0.6797 | 0.6672 | 0.6689 | 0.6775 | 0.6829 | 0.1567% | -0.08669% | 0.1124% | 0.3259% | 0.3086% | 0.314% | 0.01398% | 0.04501%
2020-04-09 | 0.94642 | 0.80726 | 0.70255 | 0.67841 | 0.66638 | 0.66719 | 0.67593 | 0.68068 | 0.09542% | -0.05458% | 0.1767% | 0.1357% | 0.1852% | 0.05757% | -0.2176% | -0.2809%
2020-04-21 | 0.9462 | 0.8087 | 0.7041 | 0.6793 | 0.6676 | 0.6708 | 0.6788 | 0.6836 | 0.0721% | 0.1238% | 0.3976% | 0.2669% | 0.3687% | 0.5989% | 0.2059% | 0.1476%
2020-05-27 | 0.94658 | 0.807 | 0.70298 | 0.67829 | 0.66613 | 0.66945 | 0.67773 | 0.68247 | 0.112% | -0.08679% | 0.2372% | 0.1179% | 0.1471% | 0.397% | 0.04837% | -0.01841%
2020-06-17 | 0.94338 | 0.80608 | 0.70224 | 0.67722 | 0.6649 | 0.6673 | 0.67688 | 0.68051 | -0.2258% | -0.2003% | 0.1317% | -0.03968% | -0.03756% | 0.07473% | -0.07709% | -0.3056%
2020-06-30 | 0.9448 | 0.8066 | 0.703 | 0.6783 | 0.6681 | 0.6693 | 0.679 | 0.683 | -0.07596% | -0.1362% | 0.2407% | 0.1193% | 0.4439% | 0.374% | 0.2354% | 0.05966%
2020-07-06 | 0.9453 | 0.8109 | 0.7066 | 0.6839 | 0.6735 | 0.6757 | 0.6845 | 0.6892 | -0.02308% | 0.3962% | 0.7541% | 0.9458% | 1.256% | 1.334% | 1.047% | 0.968% | Cyclo opening
2020-07-06 | 0.94204 | 0.80526 | 0.70081 | 0.67775 | 0.66542 | 0.6678 | 0.67682 | 0.68069 | -0.3684% | -0.3023% | -0.07208% | 0.03769% | 0.04129% | 0.1496% | -0.08676% | -0.2795% | Repetition in the afternoon
2020-07-27 | 0.9476 | 0.80975 | 0.70618 | 0.6847 | 0.67256 | 0.67649 | 0.68362 | 0.68773 | 0.2197% | 0.2542% | 0.694% | 1.064% | 1.115% | 1.452% | 0.9171% | 0.7521%
2020-08-07 | 0.9468 | 0.80864 | 0.70658 | 0.68363 | 0.67263 | 0.67533 | 0.68246 | 0.68493 | 0.1358% | 0.1167% | 0.7516% | 0.9066% | 1.125% | 1.279% | 0.7457% | 0.343%
2020-08-24 | 0.94571 | 0.80713 | 0.70255 | 0.67905 | 0.6696 | 0.67143 | 0.67981 | 0.68208 | 0.01986% | -0.07059% | 0.1764% | 0.2303% | 0.6698% | 0.6935% | 0.3545% | -0.07492% | New site config
2020-10-06 | 0.94631 | 0.81054 | 0.70534 | 0.68415 | 0.67365 | 0.67537 | 0.68353 | 0.68881 | 0.08389% | 0.3521% | 0.5744% | 0.9833% | 1.278% | 1.285% | 0.9046% | 0.9102%
2020-10-11 | 0.94852 | 0.81122 | 0.70558 | 0.68066 | 0.66994 | 0.67093 | 0.67886 | 0.6845 | 0.3175% | 0.4362% | 0.6085% | 0.4676% | 0.72% | 0.6179% | 0.2141% | 0.2791% | Cyclo opening
2020-11-04 | 0.94763 | 0.80868 | 0.70282 | 0.67716 | 0.66554 | 0.66778 | 0.67566 | 0.6816 | 0.2232% | 0.1215% | 0.2152% | -0.04935% | 0.0585% | 0.1455% | -0.2574% | -0.1453%
2020-11-18 | 0.94318 | 0.80427 | 0.70069 | 0.67755 | 0.66499 | 0.66758 | 0.67579 | 0.6806 | -0.2474% | -0.4243% | -0.08837% | 0.00807% | -0.02388% | 0.1153% | -0.2383% | -0.2913% | Acute deflector change
2021-01-21 | 0.94363 | 0.80661 | 0.70133 | 0.6765 | 0.66585 | 0.66718 | 0.67532 | 0.68046 | -0.1994% | -0.1346% | 0.003305% | -0.1464% | 0.1061% | 0.05629% | -0.3075% | -0.3128%
2021-02-14 | 0.94328 | 0.80593 | 0.7015 | 0.67789 | 0.66586 | 0.66892 | 0.67794 | 0.68209 | -0.2368% | -0.2187% | 0.02621% | 0.05896% | 0.1075% | 0.3171% | 0.07884% | -0.07304% | after cyclo PM
2021-03-23 | 0.94329 | 0.80443 | 0.69887 | 0.67505 | 0.66247 | 0.66549 | 0.67387 | 0.67908 | -0.2359% | -0.4048% | -0.3481% | -0.3612% | -0.402% | -0.1977% | -0.5225% | -0.5149%
2021-04-27 | 0.93803 | 0.80171 | 0.69618 | 0.67016 | 0.65895 | 0.65963 | 0.66878 | 0.67411 | -0.792% | -0.7421% | -0.7324% | -1.082% | -0.9314% | -1.076% | -1.274% | -1.242%
2021-05-05 | 0.93941 | 0.804 | 0.6977 | 0.6744 | 0.66171 | 0.66229 | 0.67169 | 0.67587 | -0.6464% | -0.4575% | -0.5152% | -0.4558% | -0.5171% | -0.6775% | -0.8436% | -0.9842%
2021-05-09 | 0.94002 | 0.80167 | 0.69532 | 0.6691 | 0.65709 | 0.65899 | 0.66752 | 0.67297 | -0.5817% | -0.7464% | -0.8543% | -1.239% | -1.211% | -1.173% | -1.459% | -1.409% | After IC cyclo replacement
2021-06-15 | 0.9398 | 0.8046 | 0.6988 | 0.6737 | 0.661 | 0.6631 | 0.6722 | 0.6774 | -0.6048% | -0.3838% | -0.3582% | -0.5597% | -0.6236% | -0.5558% | -0.7684% | -0.7607%
2021-07-20 | 0.94114 | 0.80499 | 0.69992 | 0.67488 | 0.66358 | 0.66407 | 0.67297 | 0.67981 | -0.4635% | -0.3357% | -0.1991% | -0.386% | -0.235% | -0.4108% | -0.6542% | -0.4071%
2021-10-04 | 0.94358 | 0.80848 | 0.70262 | 0.67846 | 0.6656 | 0.66742 | 0.67418 | 0.68061 | -0.205% | 0.09642% | 0.1871% | 0.143% | 0.06817% | 0.09158% | -0.4767% | -0.2911% | After BCREU replacement
2022-04-25 | 0.94466 | 0.80567 | 0.70287 | 0.6778 | 0.66579 | 0.66746 | 0.67566 | 0.68041 | -0.09031% | -0.2519% | 0.2215% | 0.04549% | 0.09612% | 0.09854% | -0.2579% | -0.3195%
2022-10-14 | 0.94425 | 0.80302 | 0.69982 | 0.67649 | 0.66629 | 0.66831 | 0.6763 | 0.68162 | -0.1342% | -0.5793% | -0.2128% | -0.1485% | 0.1717% | 0.2256% | -0.1634% | -0.1423%
2022-11-18 | 0.94255 | 0.80358 | 0.70026 | 0.67598 | 0.66701 | 0.66791 | 0.67559 | 0.67961 | -0.3137% | -0.5101% | -0.1493% | -0.2229% | 0.2796% | 0.1657% | -0.2677% | -0.4363%